## Slide 1
Takashi Hachiya,
Nara Women’s University & RIKEN BNL Research Center
March 2nd, 2021
Final Design of the INTT Ladder and Production Readiness Review
Ladder Performance
WBS: 3.01

## Slide 2: The INTT ladders
Tracking System: MVTX+INTT+TPC
Ladder
CC
BEX
HDI
SSS
ROC
SSS
HDI
BEX
CC
SSS
SSS
ROC
Silicon module
Silicon module
Data
Data
Si : 320um thick
Channel size : 78um x 16mm(20mm),
Nchannels : 128 x 26 x 2 = 6656
  Ch / chip = 128
  Readout Chip (FPHX) = 26 x 2

## Slide 3: Sensor + Readout Schematics
Performance is studied using:
  Charged particle by beam test, cosmic ray and RI sources
  Test pulse for Readout electronics
Vbias
Test Pulse
Analog
ADC
ROC
+
FEM
Digital data processing
Readout Chip (FPHX)
h+
Notes: Practical
Condition of readout chain

## Slide 4: Beam Test Setup in 2019
Beam test at FNAL 2019
Proton beam w/ 120 GeV
prototype ladder
Beam position is clearly seen
MIP peak
Efficiency
prototype half ladder

## Slide 5: MIP peak in the beam test
MIP peak measurement with 3bit ADC
  Scan with fine ADC threshold
MIP peak clearly seen
  separated from noise
MIP peak puzzle
  The peak is ~700mV from the beam test and cosmic ray.
  Eloss ~ 100keV in 320um Si
    Expected ~ 600mV
  Conversion formula from Eloss to the measured voltage fixed
Conversion formula (Eloss/strip  V)
  Gain value (G) (mV/fC)
  Offset (O) (mV)

## Slide 6: MIP peak by cosmic ray
MIP peak ~ 600mV
  Double hits are ~200mV higher from single hit
    Double hit : V1 + V2 = MIP + offset * 2
    Single hit : V1 = MIP + offset
  MIP peak is nicely reproduced by ToyMC & GEANT w/ Gain + Offset
Raw ADC have ~200mV offset
  Confirmed by test pulse
Single hit
Double hits
Pulse height after clustering
Gain (mV/fC) | 100 (85~100)
Offset (mV) | 198
Single hit
Double hits

## Slide 7: Efficiency in beam test 2019
Analysis cuts
  For L1&L2: single hit & ADC>3
    L0 : multiple hits & No ADC cuts
  Correlated hits in 3 layers
Efficiency
Efficiency (L0)=96.0  0.6 %
  consistent w/ result in 2018 (95.8  0.2 %)
  S/N from MIP suggests higher efficiency
  Investigating why the efficiency drop occurs
    Hypothesis is the timing of signal processing
Exp. vs L0
Exp – L0

## Slide 8: Test pulse response
Scan ADC with changing TP input amplitude
  Bias voltage is applied to Sensor
Results:
  ADC threshold clearly seen
  Linearity between ADC (output) vs amplitude (input)
Amplitude (input)
Outputs (max =20)
Single ch : 20 TP inputs
Threshold
Working as is expected
Amplitude (input)
ADC
All thresholds works
Linear
Amplitude (input)
Channels

## Slide 9: TP response for one ladder
All channels work nicely
TP is useful for health check of the readout chain
Amplitude vs Channels
ADC vs Amplitude

## Slide 10: 90Sr response
90Sr on Chip=3 and 16
Chip3
90Sr
Trigger scintillator is placed behind the sensor
Chip16
e– from 90Sr is also seen.
  Good to use for sanity check during the mass production

## Slide 11: Performance summary
MIP measurement
  Mip peak clearly seen from beam test and cosmic ray
  MIP peak is consistent with the expected value with the gain + offset
Efficiency in beam test
  96.0  0.6% due to the readout timing
  Investigating the time scan in the readout chain using cosmic ray
Test pulse is used for ch-by-ch health check of the readout
  RI source is also useful to check the sensor + RO
Evaluation test will be used for the mass production of the ladder

## Slide 12: Back Up

## Slide 13: TP response for one ladder
All channels work nicely
TP is useful for health check of the readout chain
Amplitude vs Channels
ADC vs Amplitude

## Slide 14: Slowing Effect on Efficiency
Data1
Data0
B-Clock
Slewing effect may cause the efficiency loss
  If Data0 pulse is small, data0 can be latched by one clock later than Data1.
To study this effect, we calculate the efficiency with the different clock timing
  Timing can be changed by ADC
Efficiency
  Different timing : 97.2  1.0 %
  Default : 96.0  0.6 %
  Slewing effect = readout timing is an possible cause of the efficiency loss.
We will study the timing effect at the next beam test.
This effect will be cross-checked by Cosmic ray measurement.

## Slide 15: MIP measurement with FPHX chip
3 bit ADC but these threshold can be set 8 bit DAC width.
MIP Scan
  Took data with changing DAC value to cover the full MIP range
ADC (3bit) | DAC (8bit)
0 | 20
1 | 25
2 | 48
3 | 98
4 | 148
5 | 172
6 | 223
7 | 248
SCAN4
SCAN1
SCAN2
SCAN3
SCAN0
charge

## Slide 16: MIP peak by cosmic ray
Single hit
Double hits
Single hit
Double hits
Toy MC
Double hits are shifted ~200mV higher from single hit
  If no offset, these should be the same
Data is nicely reproduced by ToyMC w/ Gain + Offset
  Single and double hits peaks are nicely consistent
  GEANT model also reproduce the MIP peak
MIP peak puzzle is fixed with new conversion formula
Cosmic ray data
Pulse height after clustering
 | Original | NEW
Gain (mV/fC) | 300 | 100 (85~100)
Offset (mV) | 0 | 198
by Mika

## Slide 17: Result2: noise and data error rate
Noise: check the slope of the threshold
  Width represents noise level
Transfer efficiency: check if all the data comes with high amplitude
These tests should used for the ladder (readout) QA for the mass production
Pulse height
Outputs (max =20)
Single ch : 20 TP input
Threshold
Noise : Width of the slope
Avg: 3.3±0.7
95
96
97
99
98
100
efficiency ~ 100%
Ch

## Slide 18: Readout Check by Test pulse
TP injected to a channel with changing pulse heights
  10 TP for an amplitude (Amplitude : unit of TP pulse height (1-64))
Results
  Data comes out above the threshold for 128 channels
  Linearity of ADC comfirmed
TP response for a channel
TP for a chip (128ch)
ADC vs Amplitude

## Slide 20: The INTT Components
Tracking System: MVTX+INTT+TPC
Ladder
CC
BEX
HDI
SSS
ROC
SSS
HDI
BEX
CC
SSS
SSS
ROC
Silicon module
Silicon module
Data
Data
1.2 m Bus Extender (BEX) Prototype-II
HDI Pre-production
Silicon Pre-production
FPHX Production
Conversion Cable (CC) Pre-Production

## Slide 21: INTT module
Si : 320um thick
Channel size : 78 x 1200um(),
Nchannels : 128 x 26 = 3328
  Ch / chip = 128
  Readout Chip (FPHX) = 26