## Slide 1: Understanding Systematics in the SRS: Timing, Saturation, and Calibration
EIC Tracking Workshop, Temple university
May 9, 2015
Michael Phipps, BNL

## Slide 2: Background for our SRS Studies
SRS used for our Mini-drift detector test beam data
Mini-drift detector features:
  Extended drift gap (16.5 mm) to allow vector reconstruction of track
  Chevron-style readout pads to segment charge and minimize channel cost
SRS not optimal for detectors with large charge/channel, since originally designed for silicon detectors at CMS
To understand the performance of our mini-drift detector we have to be able to ascribe systematics to the readout system and calibrate it for optimal use within our dynamic range

## Slide 3: SRS Architecture: A Quick Look
APV25
FEC
Preamp
Shaper
Multiplexer Gain (MUX Gain)
ADC
FPGA
Multiplexer
Analog Pipeline
Preamp/shaper Gain: 100 mV/mip
(25,000 electrons)
Dynamic Range: ~5-6 mip
Shaping Time Constant: 50 ns
Ring Buffer of 192 cells
Analog gain: 100 m
128:1
10 bit at 40 MHz
System Clock
LHC clock: 40 MHz (25 ns)

## Slide 4: Timing Uncertainty
D(t)
Trigger window: trigger equally anywhere within this window
25 ns time bins
t
Measured Uncertainty with External Pulser
At least 7.2 ns of “unfoldable” uncertainty

## Slide 5: Hybrid Saturation
SRS designed for silicon MIPs of 25k electrons
Shaper has narrow dynamic range with saturation around 175k electrons
Serious problem for wide readout strips (chevrons) or large drift detectors.
Time bin centroid measurements can help minimize by using many unsaturated samples
Clipping

## Slide 6: Nonlinear Response
MUX1 correction curve
Average Response Across All Channels Across all Input Values -> measures response nonlinearity
Nonlinearity as function of MUXGAIN and IMUXIN
Conclusions:
Each MUXGAIN needs its own calibration curve
IMUXIN is just an additive constant -> nonlinearity is in the hybrid and independent of location in FEC range
Correction applied AFTER pedestal subtraction
Nonlinearity around 70k electrons
FEC saturation
Convert from electrons to corrected ADC

## Slide 7: SRS Calibration: Signal/Noise
** For better signal to noise, use Muxgain 3 (which is really Muxgain 4)
Mux 1
Mux 2
Mux 3
Mux 4
Fe55 Spectrum: Peripheral pad from highly collimated xray gun

## Slide 8: SRS Calibration: Dynamic Range
** For better dynamic range, use Muxgain 1
- IMUXIN of 0x01 is safest location for pedestal
Fe55 spectrum across all pads for scaled Muxgains
Peripheral primary peak
- Scale the spectrum for each Muxgain from the mean of this peak
Primary Fe55 peak in center of cluster
- Falls at edge of saturation, so convenient measure of dynamic range

## Slide 9: SUMMARY
At least 7.2 ns timing uncertainty in SRS
Saturation in the hybrid ~175k electrons
Significant nonlinearity beyond 70k electrons
Gain of multiplexer can be optimized to either improve signal/noise or dynamic range

## Slide 10: Backups

## Slide 11: FEC Saturation
Dynamic range between ~400 - 3200 ADC
Since largest MUX5 signal is ~2700 ADC, MUX5 should be avoided with GEMs
MUX3 and MUX4 pedestals have to be carefully aligned within the FECs dynamic range (modify the IMUXIN of the preamp)
Upper limit:
3200 ADC
Lower limit:
400 ADC
Pedestal above threshold
FEC Saturation

## Slide 12: Channel-to-channel calibration
APV Internal Pulser
Internal Pulser
Mux 2
Internal Pulser
Mux 4
Channel dependent non-linearity curves
Signal response per channel

## Slide 13: MUXGAIN: *Resistors for Mux3 and 4 appear to be switched in all APVs (ie Mux 4 is nominal Mux)
MUX GAIN | Approximate Hybrid Range (Pedestal-subtracted ADC) | Multiplexer Gain (mA/mip)
1 | 1250 | 0.6
2 | 1700 | 0.8
3 | 2500 | 1.2
4 | 2100 (~11 bit) | 1.0
5 | 2900 | 1.4
MUXGAIN 1
MUXGAIN 2
MUXGAIN 3
MUXGAIN 4
MUXGAIN 5