## Slide 1: BMX Interferometer: Calibration Studies for a 21cm Intensity Mapping Experiment
P. O’Connor, Anze Slosar, Justine Haupt, Ben Saliwanchik, Paul Stankus, Tom Throwe (BNL)
Laura New burgh, Maile Harris, Emily Kuhn, Annie Polish, Will Tyndall (Yale)
Gregory Troiani ((UMKC)
O'Connor, et al.. (2020, December). The Baryon Mapping eXperiment: a 21cm intensity mapping pathfinder. In SPIE Ground-based and Airborne Telescopes VIII (Vol. 11445, p. 114457C).
arXiv:2011.08695

## Slide 2: Studying Dark Energy using 21cm intensity mapping
Study of Dark Energy and Inflation today relies on observations of:
  CMB temperature and polarization fluctuations;
  Distance-redshift of distant supernovae;
  Large-scale structure of matter in the universe, via galaxy spectroscopic surveys, weak gravitational lensing, etc.
The 21cm radio emission of neutral hydrogen is an alternate tracer of large scale structure of the Universe:
  Hydrogen line provides accurate, unambiguous redshift measure
  Low-resolution telescope can be used to detect aggregate emission all the hydrogen within a voxel (Intensity Mapping, IM)
  Emission is detectable through Earth’s atmosphere to redshift z > 10
Uses precisely calibrated radio interferometers to achieve survey speed is potentially faster and cheaper (per source-equivalent) than an optical survey
  Observe source redshifts in the range 0  z  6 inaccessible to CMB and optical surveys
  Telescope = antenna array designed for wide instantaneous field of view, modest angular resolution, and large collecting area
Takes advantage of rapid commercial development of digital RF and high-performance computing technology.

## Slide 3: Origin and ubiquity of 21cm emission
Spin-flip transition of neutral hydrogen
Spectral line at 21cm rest-frame wavelength is sharp and isolated
 Once detected, provides precise redshift
Nearby galaxy M31 imaged in 21cm
Fernandez
2016
Highest redshift detection (z = 0.376)
21cm and optical surveys detect similar structure to z ~ 0.07
Haynes et al. 2011
Red: SDSS optical
Blue: ALFALFA 21cm
Grey: optical
Purple: 21cm

## Slide 4: Intensity Mapping technique
Several 1st-generation IM dark energy experiments online, under construction, and proposed: CHIME (Canada), HIRAX (S. Africa), BINGO (Brazil), TIANLAI (China)
US participation minimal at this time
Traditional galaxy survey:
individual sources observed, one at a time with spectrograph
Intensity mapping survey:
integrated emission observed as a function of frequency (redshift)
choose Dq, Df to be sensitive to scales of interest to cosmology
Df ~ 1MHz
Dq~ 0.2°
Build up tomographic reconstruction of density field across large volume of space
pointing direction
redshift
(line-of-sight distance)
redshift
(line-of-sight distance)
Scale of interest for cosmology

## Slide 5: A concrete example: Packed Ultrawideband Mapping Array (PUMA)
A next-generation cosmic survey using intensity mapping of the 21-cm emission from neutral hydrogen
Proposal submitted to the ASTRO2020 Decadal Survey and Snowmass LOI call
Interferometric array of 32,000 (5,000) six-meter dishes closely packed
Redshift range 0.3 < z < 6 corresponding to 1100 < n < 200 MHz
Primary science goals:
  Probing physics of dark energy in the pre-acceleration era
  Searching for signatures of inflation
  Probing the transient radio sky (fast radio bursts and pulsars)

## Slide 6: BMX instrument
Parent paraboloid
EM Sim of illumination pattern on dish
(fixed)
Off-axis configuration of single dish + receiver
Front end electronics
Calibrated noise diode

## Slide 7: Instrument design: radio interferometer
One baseline (Ndish = 2)
Interference fringes based on path length difference between 2 stations
D
b
l
Angular resolution l/|b|
Angular field of view l/D “primary beam”
Collecting area pD2Ndish/4
NRAO
Phase of the interference fringes encodes angle to source (q ):
=
B(q) = angular response of dish primary beam
Airy disc and Gaussian approximation
EM sim (not BMX)

## Slide 8: Array baselines

## Slide 9: BMX field of view on sky

## Slide 10: Calibrators
 | DIODE | UAV | GNSS | MW | CygA
For: | gain(n) | beam(n) | beam | beam, frequency | beam, array layout
Range | 0 | 170 m | 2e7 m | 5e20 m | 6e24 m
Heading angle | - | Any | ~NE, ~SW | W | W
Angular rate | - | 1 – 1.1 ◦/s | 0.006 - 0.01 ◦/s | 0.0032 ◦/s (sidereal) | 0.0032 ◦/s (sidereal)
Passes/day | - | up to 10 | 20 – 30 | 1 | 1
Power | 3K | 200,000K adj. | 20,000K max. | 1 -- 40K | 5 K @ 1.42GHz
Polarization | No | Yes | No | No | No

## Slide 11: GNSS satellites: constellations, passes over BMX

## Slide 12: GNSS passes over BMX on 1/3/2020 (13 satellites observed in 4 constellations)
GPS
GPS
GPS
GPS
GAL
GAL
GAL
BEI
BEI
BEI
GPS
GAL
GAL
Assigned transmit bands
Predicted tracks from ephemeris

## Slide 13: Fitting beam parameters to multiple GNSS signals
Integrate measured single-dish power over GNSS band from 1100 to 1378 MHz.
Remove DC offsets
Define 1-hour time windows centered around transits of up to 16 GNSS satellites (one day at a time).
Assemble “stitched” data:
  Data set is predicted [qx(t), qy(t)], signal(t).
Jointly fit 2D Gaussian beam model having parameters amplitude, beam pointing center and width (qx0, qy0, sx, sy).
Concatenated input data set (16 satellites on 1/10/2020)
Joint fit
Method

## Slide 14: Fitted beam parameters over 21-25 Jan. 2020. Data for W dish, Y polarization
beam parameter | mean | st. dev.
qx0 | 0.3639 | 0.1229
qy0 | -1.5392 | 0.1863
sx | 1.5953 | 0.1568
sy | 1.7222 | 0.1030
 | deg | deg
qx0
qy0
sx
sy
degrees (mean=0)
pointing direction
beam width
statistics (mean subtracted)

## Slide 15: 21cm clouds in Milky Way
Milky Way map in RA and frequency for the BMX strip. Top map is BMX data, bottom is rebinned data from the HI4PI survey1. Note, both maps start at 18h RA (left). After 24 hours, the BMX map extends to 22h the following day while the HI4PI map simply wraps around. Obs. Start: 21:45 UTC 20 October 2020.
1Bekhti et al. A&A 2016
1420.12 – 1420.25
1420.25 – 1420.39
1420.39 – 1420.52
1420.52– 1420.65
1420.65– 1420.79
Detail of galactic HI structure. Each of the five sub-panels shows BMX spectrometer data in adjacent 130kHz frequency bins during the MW transit of 1 Dec 2017. The x-axis in each sub-panel covers the same ∼ 100 degrees of RA. After transforming frequencies to a Local Standard of Rest (LSR) which takes into account motions of the Earth and Sun, data is fit to a Gaussian beam model (shape, pointing, frequency offset) with the HI signal predicted by the HI4PI map.
Analysis by C. Sheehy

## Slide 16: UAV calibrations
Vehicle: commercial DJI hexacopter. Flight time ~25 minutes per battery change.
Transmitter: switched, broadband noise source + biconic antenna (polarized). Variable attenuator to set power level.
Position determination: differential GPS, ~1cm accuracy.
Timebase: GPS
Flight plan: ascend to far field (~170m), fly raster pattern over telescope. Drone “yaw” orientation sets polarization direction either aligned with or perpendicular to direction of travel. While flying legs of the raster, typical speed 2 - 4m/s.
Thanks to Yale team of L. Newburgh for use of DJI
Flight patterns
Repeat rasters in N-S, E-W heading, each polarization, each power level
Higher transmit power and larger extent used to probe sidebands

## Slide 17: UAV calibrations
Angular response of the 4 dishes
Beam width vs. frequency
Beam centers vs. frequency
Analysis by G. Troiani

## Slide 18: Cygnus A transit interference fringes
BMX DAQ records cross-correlations of each pair of dishes (baselines) across the full frequency band
Fringes with SNR ~200 are seen during transit of pointlike CygA source
For every baseline and polarization, fit observed fringes to 1-D model having parameters: amplitude, pointing offset from zenith, width of (composite) beam in Gaussian approximation, projected baseline length in E-W direction, and time delay between signal transmission paths. Position of CygA in horizon coordinates known to high precision.
Repeat fit for several frequency bins, avoiding regions likely to be contaminated by nearby GNSS passes.
GNSS interference
Altitude of CygA
E-W baseline b24
Fringes at 1105MHz
Data (blue), fit (red), envelope (green)
Fitted parameters

## Slide 19: Fringes from transit of 17 April 2019
polarization Y
polarization X
polarization Y
polarization X

## Slide 20: Interferometer results
Left: Pointing offsets from CygA visibility fitting. Upper: X-polarizations. Lower: Y-polarizations. Colors: (blue, orange, green, red, magenta, brown) = (N-E, N-S, N-W, E-S, E-W, S-W) dish pairs. Statistics for 81 days of observations in early 2020. Right: fitted baselines.

## Slide 21: SUMMARY
21cm intensity mapping is a new, cost-efficient observational technique that is complementary to optical and CMB surveys.
It opens the largely unexplored redshift range 2.5 < z < 6 where beyond-LCDM physics can be studied - dynamic DE, modified GR, inflationary relic signatures.
Leverages industry advances (wireless, AI) and requires no specialized detector environments (cryo, radiation).
Research needs center on DAQ architectures (with 2030-era electronics), and on calibration methods, including sub-picosecond phase synchronization.
BMX can serve as an early pathfinder to a future large project such as PUMA
  well-matched to HEP expertise and is synergistic with many emerging trends in EF and IF electronics.

## Slide 22: Future plans
BMX “tune up”:
  Improve self-RFI environment at BMX site
  Align beams (adjust dish and horn positions)
EM Sims of beam shape
Improve joint fitting of multiple calibrators (drone, GNSS, CygA)
Replace digitizer + channelizer with RFSoC-based platform
GPS message decoding to improve accuracy of array layout
Acquire stable data for ~months
  detect cosmic 21cm signal at 0 < z < 0.3 in cross-correlation with galaxy survey

## Slide 23
BACKUP MATERIAL

## Slide 24: Current/Upcoming OHEP-sponsored cosmic surveys
Dark Energy Spectroscopic Instrument (DESI)
Dark Energy Survey (DES)
Large Synoptic Survey Telescope (LSST)
Galaxy Imaging, broadband filters
504 Mpix CCD Focal Plane
Operations 2014 - 2019
Galaxy Spectra
5000 Fiber Focal Plane
Operations 2019- 2024
Galaxy Imaging, broadband filters
3.2 Gpix CCD Focal Plane
Operations 2023 - 2033
?
Power of a cosmic survey to measure cosmological parameters is limited by
  redshift range and accuracy
  sensitivity (number of sources)
  scale
Improved statistics has to come from increasing survey speed and/or increasing sensitivity to fainter/redder sources, while preserving redshift accuracyr.

## Slide 25: Obstacles to scaling optical surveys
Detector technology for IR
Atmospheric transparency in IR
Telescope cost scaling
Let’s consider another wavelength range…
z = 0
z = 3.5
Airglow

## Slide 26: Key technological challenges
Since the cosmological 21cm signal is ~104 weaker than astrophysical foregrounds, the major challenge for this method is foreground mitigation.
Interferometers couple bright foregrounds into spatial and frequency modes, making straightforward subtraction difficult.
The key to achieving clean foreground separation is calibrations:
  Antenna element primary beam angular response vs. frequency
  Gain and phase response of the signal chain vs. frequency
  Accurate sky maps of (polarized) galactic and extraglactic sources
Precision timing distribution to ensure coherent recording of GHz signals arriving at thousands of stations separated by km
Power-efficient, real-time processing of network data streams approaching 1Pb/sec
Robust and mass-producible methods of dish and receiver manufacture

## Slide 27: Instrument design: radio interferometer
One baseline (Ndish = 2)
Interference fringes based on path length difference between 2 stations
D
b
l
NRAO
50 marcsec
EHT
Collecting area pD2Ndish/4
Angular resolution l/b
Angular field of view l/D
Large b, small l
Large D, small N

## Slide 28: Instrument design: radio interferometer
One baseline (Ndish = 2)
Interference fringes based on path length difference between 2 stations
D
b
l
Collecting area pD2Ndish/4
Angular resolution l/b
Angular field of view l/D
NRAO
Large b, small l
Large D, small N
Small b, large l
Small D, large N
EHT
CHIME
50 marcsec

## Slide 29: 21cm intensity mapping experiments (0.8 < z < 2.6)
CHIME (Canada)
HIRAX (S. Africa)
TIANLAI (China)
PAON-4 (France)

## Slide 31: PUMA Science
Expansion history
Growth of structure
Fast Radio Bursts (FRBs)
Primordial nonGaussianity
arxiv:1810.09572
arxiv:1907.12559
LCDM
Mod. GR
PUMA

## Slide 32: Galaxy statistics and redshift range vs. optical galaxy surveys
Length of bar = redshift range
Area of bar = effective galaxy number
Speed of LSST, accuracy of DESI

## Slide 33: Instrument design: radio interferometer (large-N)
Many baselines (Ndish = 5000/32000)
Every pair of stations provides a baseline
Nbaselines =
PUMA 5K
PUMA 32K

## Slide 34: Instrument design: radio interferometer (large-N)
Many baselines (Ndish = 5000/32000)
Every pair of stations provides a baseline
Nbaselines =
Each baseline probes a corresponding spatial frequency
PUMA 5K
PUMA 32K

## Slide 35: Instrument design: radio interferometer (polychromatic)
One baseline (Ndish = 2)
Interference fringes based on path length difference between 2 stations
b
l1
Collecting area pD2Ndish/4
Angular resolution l/b
Angular field of view l/D
NRAO
AMP
FFT
ADC
AMP
FFT
ADC
X
X
X
X
 
 
 
 
“F-engine”
“X-engine”
l2

## Slide 36: Full array network
SWITCH
CORRELATOR
Artifact removal
Component selection
Mapmaking
MASTER TIMING
PHASE/FREQ TRANSFER (active)
25-50 Gbps
jitter < 600fs
relative phase drift < 150fs

## Slide 37: Switch and X-engine details
D1
D2
D3
D4
D5
…
DN
dishes
freq. channels
To X-engine
Re-organize inputs, group by frequency channels
Send to X-engine
Perform all auto- and cross-correlations (realtime)
Average redundant baselines
Integrate results

## Slide 38: Calibration challenges
Since the cosmological 21cm signal is ~104 weaker than astrophysical foregrounds, the major challenge for this method is foreground mitigation.
Interferometers couple bright foregrounds into spatial and frequency modes, making straightforward subtraction difficult.
The key to achieving clean foreground separation is calibrations:
  Antenna element primary beam angular response vs. frequency
  Gain and phase response of the signal chain vs. frequency
  Accurate sky maps of (polarized) galactic and extraglactic sources
Additional techniques are needed to cope with man-made RFI.
W. Tyndall 2019
Antenna beam pattern (sim.)
frequency 
Cygnus A E-W transit Fringe Fit

## Slide 39: Front End Electronics Development
BMX RF analog chain
Xilinx ZCU111 RFSoC FPGA digitizer/channelizer
400kHz channelized data from RFSoC digitizers operating at 3.1GS/s
Real-time FFT in firmware (unwindowed)
Software FFT with Blackman window
39
BMX RF analog chain
Measured noise temperature
Xilinx ZCU111 RFSoC FPGA digitizer/channelizer
400kHz channelized data from RFSoC digitizers operating at 3.1GS/s
Real-time FFT in firmware (unwindowed)
Software FFT with Blackman window

## Slide 40: Antenna directional response characterization
Beam calibration using UAV and/or satellite
~20 GNSS satellites transit within 10 degrees of BMX zenith each day (4 constellations)
Drones!
Global navigation satellites
L. Newburgh
fixed-wing (BMX)
quadcopter (Yale/OVRO)

## Slide 41: Low-cost dish construction methods
2.5m
6m
15m
Notional PUMA dish
BMX dish
Composite dish construction
CHORD
HIRAX
SKA

## Slide 42: By the numbers…
 | PUMA-5K | PUMA-32K | 
Raw data rate | 240 | 1500 | Tbit/s
Real-time computation | 15 | 100 | PFLOP
Output data rate | 13 | 82 | GB/s
Data volume | 0.9 | 5.7 | PB/day
Power * | .23 | 1.5 | MW
DATA
DOLLARS
DEVELOPMENT TIMELINE
*no ASIC

## Slide 43: Summary
puma.bnl.gov
21cm intensity mapping is a new, cost-efficient observational technique that is complementary to optical and CMB surveys.
It opens the largely unexplored redshift range 2.5 < z < 6 where beyond-LCDM physics can be studied - dynamic DE, modified GR, inflationary relic signatures.
Leverages industry advances (wireless, AI) and requires no specialized detector environments (cryo, radiation).
Detector elements require exquisite control of spatial and frequency characteristics.
Research needs center on DAQ architectures (with 2030-era electronics), and on calibration methods.
The PUMA project is well-matched to HEP expertise and is synergistic with many emerging trends in EF and IF DAQ.
Opportunities exist for multi-agency and international partnering.
2005 cellphone

## Slide 44
Thank you
puma.bnl.gov