| Tests | Purpose | Method | Criteria |
| --- | --- | --- | --- |
| Number of channel entries | To count the number of dead, hot and abnormal channels | - 1-D histogram of the channel entry distribution from all channels of all ladders - Fit with a Gaussian and define the boundary for good channels. | Good channels within around f 5-sigma of the Gaussian distribution For short conversion cable: 280 < entries < 400 For long conversion cable: 260 < entries < 370 |
| Width of amp distribution from all ADC channels with mean-shift | To quantify the effect of noise. | - Obtained the mean of the amplitude distribution for each ADC channel - Define the "central line" as the position of the ADC0 mean. - Shift the mean of the amplitude distribution of each ADC channel to the "central line" to produce an accumulated distribution. With their means shifted to the common position, i.e. "central line", the width of the cumulated distribution is from the noise. The influence from slope is corrected by the "mean shift". - Fit the width distribution from all channels of all ladders with a Gaussian and define the boundary for good channels. | Good channels within around five sigma of the Gaussian distribution For short conversion cable: 2.2 < width < 4.2 For long conversion cable: 2.5 < width < 4.3 |
| Width of amp distribution vs. ADC | To quantify the variation of noise vs. ADC. | - plot "width vs. ADC", where width is the Gaussian width of the amplitude distribution. - fit the plot with a constant line, i.e. "pol0" in ROOT, and obtain chi2/NDF from the ROOT. - Calculate p-Value, i.e. TMath::Prob(chi, NDF) - Set cut based on the distribution. - e.g. - p-Value = 0.006%--> 5 sigma | To be filled by Chen-wei and Han-sheng when the result is out |
| Linearity of ADC vs. amplitude | To check the linear response of ADC vs amplitude | - plot "mean vs. ADC", where the mean is the Gaussian mean of the amplitude distribution. - fit the plot with a straight line, i.e. "pol1" in ROOT, and obtained chi2/NDF from the ROOT. - Calculate p-Value, i.e. TMath::Prob(chi, NDF) - Set cut based on the distribution. - e.g. - p-Value = 0.006%--> 5 sigma | To be filled by Chen-wei and Han-sheng when the result is out |
| Position of amplitude mean of ADC0 | To check the offset of ADC response. Note: can be corrected via calibration in principle | - make a distribution of the amplitude mean of ADC0 from all channels of all ladders. - set boundary cut to reject bad channels | To be filled by Chen-wei and Han-sheng when the result is out |
| Slope of amplitude mean vs. ADC | To check the gain of ADC. Note: can be corrected via calibration in principle | - plot the distribution of the slope of amp mean vs, ADC from all channels of all ladders. - set boundary accordingly - Note: the slope will vary with channel ID due to, e.g. different signal path-length. | To be filled by Chen-wei and Han-sheng when the result is out |
| Unbounded channel | To quantify the number of channel missing wire-bounding. Note: not needed at BNL according to Rachid’s feedback. | - Run the calibration test without bias voltage. - The noise of an un-bonded channel will be very small as it doesn’t connect with sensor which is the source of noise. | To be filled by Chen-wei |
| Classification | To classify a ladder | - remove double counting on bad channels from all above criteria - plot “number of half ladders” vs. “number of bad channels” - printout the ladder information for each class. | classification based on the percentage of bad channels in a half ladder - 1: 0.5% - 2: 1.0% - 3: 1.5% - 4: 2.0% |

- Cheng-wei’s presentation on classification

- https://indico.bnl.gov/event/12518/contributions/52256/attachments/36063/59166/HanSheng0720.pdf