## Slide 1: sPHENIX TPC Module QA
TK Hemmick

## Slide 2: Standard Operating Procedure Document
sPHENIX SOP for Module Assembly draft was written by Dick Majka and Nikolai Smirnov
  Basically a copy of the ALICE document for the IROC manufacture.
  Adjustments made to radiation load test criterion
Initially Yale was to perform the testing using the ALICE apparatus.
  This agreement fell through.
  SBU built a “philosophical duplicate” of the Yale system.
  Procedures are criterion are basically identical except:
    Different brands of components.
    Different software running the show.
sPHENIX changes to ALICE Procedure:
55Fe test is skipped.
  sPHENIX doesn’t use dE/dx.
  Handling of 55Fe not worth it.
Stability test scaled to 40 nA/cm2
  See next slide.

## Slide 3: Radiation Load Test
sPHENIX specification:
  40 nA/cm2 for 6 hours
  R1 = 12 mA; R2 = 21 mA; R3 = 29 mA
Experience:
  Sounds extreme
  Not really troublesome in experience.
ALICE specification:
  10 nA/cm2 for 6 hours
  IROC = 17 mA
Experience:
  Sounds extreme
  Not really troublesome in experience.

## Slide 4: Nitty Gritty Details: GEM Trimming
Every GEM is barcoded and has a status from the factories (WSU, WIS, Temple, Vanderbilt) and has a status recorded in the database with regard to all factory tests.
Only Green & Yellow (e.g. visually marred not affecting performance) are trimmed.
NOTE: GEMs only handled on clean TABLE inside clean room (better than best clean room)
Database (courtesy Peter Glaesel)
Bar Code

## Slide 5: Nitty Gritty Details: Resistor Potting
ALICE experience revealed a weak point: Sharp exposed solder points.
sPHENIX design eliminates those solder points using “tails”
  Distance to next gap retained.
  Only exposed voltage is CERN-mounted inside “frame pocket”
  Fully potted to completely eliminate all exposed voltage.
Syringe injection at 45o
Cured at 90o
Sealed

## Slide 6: Nitty Gritty Details: Continuity Test
Each GEM is stacked onto the pad plane.
100 Volts is applied to the pad plane:
  Each stripe has 20 MW current limiting resistor.
  Voltmeter has 10 MW input resistance.
  33 Volts expected on each stripe
    Can reveal disconnect at resistor. (3 GEMs out of 240 failed)
    Can reveal break in HV trace to tail. (zero failures so far)
    Can reveal shorts to neighbor. (zero failures so far).
100 Volts
33 Volts
Test Probe Access Holes

## Slide 7: Nitty Gritty Details: Air HV test
Over-stress each individual GEM
  MAKE it fail now, if it every will.
  Started at 525 Volts in air, upgraded to 535 volts in air.
    NOTE: SBU clean room has lots of fans and is HOT.
    Humidity always 35% even when the rest of the lab is muggy.
  If any activity is seen, send to 555 V, leave it there (try to kill it)
Slice away bar code if & only if GEM passes HV

## Slide 8: Nitty Gritty Details: Tail Potting
HV tail exist the detector via small slot in the strongback.
Potting for gas seal.
After cure, final test can be performed for GEM-by-GEM
  This test has been replaced during the Xray box testing…

## Slide 9: Nitty Gritty Details: Xray Gas Enclosure
The Xray test box has two potential planes:
  Al-mylar sheet collects positive ions for IBF mapping.
  Open “mask” extends top GEM potential to smooth the field.
Module bolts to back plate and “window assembly” seals.
  Dove-tail O-ring grooves so O-rings never fall.
IBF Collection
“Mask”
Module “raised” to ease bolt insertion
Box top alignment pins ensure safety
Sealed box suitable for travel to non-clean environment

## Slide 10: Nitty Gritty Details: Xray Box Installation
Xray Gun mounted to X-Y translational stage
Module installation:
  Connect Gas.
  Connect HV
  Connect pad-charge collection boards.
mPORE
High Voltage
Door Interlock(s)
Pad Charge Collection
Module Install
Xray Gun

## Slide 11: Nitty Gritty Details: Xray Box Installation
Xray panel has HARDWARE readout of settings (trust but verify)
Xray panel has human-actuated interlock (allow or crash)
Standard “Zagreb” 12-channel pAmmeter
HV
NIM
pA
Xray
Venerable “Klaus’ Gas System”
  Computer control all flows/
  O2 and H2O readings
  Pressure Reading
  Temperature Reading
Used as pass-through here.
Used recirculating on TPC.

## Slide 12: Nitty Gritty Details: Control Software
Purschke Philosophy:
  All actions are prompt level shell commands
  Stateless GUI (Tcl/Tk) submits shell commands (buttons) and monitors status (labels)
X-ray controls: Zhenia Shulga
Gas Flow controls: Zhenia Shulga
High Voltage controls: Sara Kurdi
Motor Controls: Luke Legnosky
pAmmeter Controls: Zhenia Shulga
Stateless GUI: TKH

## Slide 13: Nitty Gritty Details: In Operation
Just coming alive now.
Quick scan (X-ray aperture too big & too far back)
3 hour “bombardment”…no issues.
Xrays from 32-ganged pads

## Slide 14: Schedule
Routine module testing begins after Thanksgiving.
  System working, but still needs code development, database interface, pass/fail analysis code…
Installation of modules into the TPC will proceed in parallel with the Xray characterization of the modules throughout January 2022.
For the Xray testing, we assume 2 modules per day 5 days per week from a fully dedicated crew (Luke Legnosky, Chris Reynolds, Sara Kurdi)
Schedule lag can be compensated using weekends (working well for Assembly crew)
10 days float for on-time module installation by February 1.
float
Thanksgiving
2 modules/day
3 modules/day
Today
float
Christmas
New Year’s

## Slide 15: Summary
QA at the factory (complete, largest source of failure is single GEMs):
  High voltage stability of individual GEM foils.
QA at Assembly (3 GEMS out of 240 failed…not used):
  Connectivity, stripe-by-stripe
  Overstressed in air.
QA in Xray box (starting up…refining procedure…)
  GEM-by-GEM and gap-by-gap on working gas.
  Gain map.
  IBF Map.
  Stress test (40 nA/cm2)
QA in the TPC (Spring)
  Recirculation flow.
  Sector-by-sector cosmics (ain’t streaming readout grand!)
  Demonstrate KPP (live channels, efficiency, …)
  Deliver July 2022