## Slide 1: The barrel ladder tests G. Nukazuka (RBRC), R. Nouicer(BNL), R. Pisani(BNL), M. Hata(NWU), M. Watanabe(NWU), W. Tang(NCU)

## Slide 2: Tests of the barrel ladders
 | Date | Tested layers | Results
1st test | June/10-June/14 | B0L0 | Good.
2nd test | June/17-July/07 | B0L1, some of B0L0 | Some ladders in B0L1 were noisy. GND condition was investigated.
3rd test | Aug/10- now | B1L0 (some of) B0L1&B0L0 | Ongoing Tested ladders showed good results.
4th test | later… |  | 
Notation: BxLyzz
  x: Barrel ID (0 for inner or 1 for outer)
  y: Layer ID (0 for inner or 1 for outer)
  zz: Ladder ID (from 0 to 15)
B0L0
June/06 : Ladder installation and cabling done
June/10 : Ladder survey done
June/14 : Ladder calibration tests done
B0L1
June/15: Ladder installation and cabling done
June/17: Ladder survey done
now: Ladder calibration tests ongoing
B0L0
B0L1
B1L0
B1L1

## Slide 3: The barrel calibration tests: The setup
3 good ROCs (NE2, NE4, and #29)
  3 ROCs were put on each side in the 3rd test
The FEM/FEM-IB system
The same LV system as the test bench
  upgraded in the 3rd test
The new temperature monitor system
The ladder cooling and the ROC cooling
The FPC conversion cable or the μ-coaxial conversion cable was used. No difference found b/w the 2 types.
A single ladder was operated to test it.
NE2
NE4
#29

## Slide 4: Results: A typical good case (the 1st&2nd tests)
  B0L002S (PB1-L031)
Barrel
(1st test)
Bias current
  type-A: 135 nA
  type-B: 102 nA
Data: 20220613_1546
Chip13
Test bench
Bias current
  type-A: 184 nA
  type-B: 141 nA
Data: 20220513_1217, 111 cm Bus extender + 20 cm FPC conversion cable
Chip13
south side
Generally speaking, many ladders showed results as good as ones in the test bench

## Slide 5: Results: A typical difficult case (the 2nd test)
Many ladders showed results as good as ones in the test bench
Some ladders had more noise… …, for example: B0L103N
Barrel
Bias current
  type-A: 117 nA
  type-B: 83 nA
Data: 20220622_1248 (after modification of GND, good results were obtained)
Chip13
Test bench
Bias current
  type-A: 139 nA
  type-B: 98 nA
Data: 20220420_1927
Chip13
Due to the radiation source↓
north side
  B0L103N (PB1-L0??)
Results in the barrel had more noise than those in the test bench on some B0L1 ladders.

## Slide 6: The barrel calibration tests: The setup
B0L0 (East)
B0L1 (East)
HDI
Bus extender
Conversion cable
ROC power supply cables
Bias cables
to the bias power supply
to the power supplies
Bias
power
supply
ROC
power
supplies
FPHX power supply cables
Stave GND cables
Stave GND cables
Stave
GND
cables
Stave
GND
cables
Aluminum supporting structure
Note: CFC cone omitted
power for the lack

## Slide 7: GND investigations in the 2nd test
Aluminum supporting structure
Note: CFC cone omitted
B0L0 (East)
B0L1 (East)
HDI
Bus extender
Conversion cable
ROC power supply cables
Bias cables
to the bias power supply
to the power supplies
Stave
GND
cables
Stave
GND
cables
Without the stave GNDs connection to the aluminum supporting structure,
  B0L0 <—> the supporting structure: ~0 Ω
  B0L1 <—> the supporting structure: 180 Ω
Stave
GND
cables

## Slide 8: GND investigations: The harvest
6
7
8
9
10
11
11
10
9
8
7
6
B0L111S
South side
North side
Connecting the stave GNDs to the aluminum supporting structure drastically improves the condition.
June/29
July/07
B0L0 ↔ the supporting structure: ~0 Ω
B0L1 ↔ the supporting structure: 180 Ω
B0L0 ↔ the supporting structure: ~0 Ω
B0L1 ↔ the supporting structure: ~0 Ω
All staves must be well connected to GND. In the IR, there is a GND port for this purpose. In the silicon lab, the aluminum supporting plate on the table can be large enough GND port.
Power supplies with less noise are also necessary.

## Slide 9: Upgrades on HV&LV
HV & LV power supplies were upgraded to the same one used in IR, so the environment is much cleaner now.
old HV&LV
system ↓
The new HV&LV system
Maki Wakata
(Rikkyo Univ.)
Wei-Che Tang
(NCU)
Thanks to Maki&Wei-Che, the new power supply system is operational.
See their reports for more details:
HV by Maki
LV by Wei-Che
The 3rd test was started on Aug/10.

## Slide 10: The 3rd test
new ROC power cable
new FPHX power cable
new bias cables

## Slide 11: Results: Good news from 3rd test
  B0L109N (PB1-L023)
Many ladders showed results as good as ones in the test bench
Some ladders had more noise… …, for example: B0L103N
Barrel
the 2nd test
Bias current
  type-A: 117 nA
  type-B: 83 nA
Data: 20220811_1137
Chip13
Test bench
Bias current
  type-A: 204 nA
  type-B: 105 nA
Data: 20220621_1316
Chip26
north side
The noise in the 2nd test is not in the 3rd test. It’s as good as the one on the test bench.
Barrel
the 3rd test
Data: 20220506_1747
Chip26
Chip26
Bias current
  type-A: 143 nA
  type-B: 87 nA
Bias current
  type-A: 262 nA
  type-B: 291 nA

## Slide 12: The 3rd test: Current situation and prospect
3rd test is ongoing now, mainly by M. Hata and M. Watanabe.
For the moment, all tested ladders showed very good results like in the test bench.
In the best cases, testing 1 half-ladder takes ~10 min. If we test all ladders, there are 2×(12 + 12 + 18) = 84 ladders, it takes about 14 hours.
20 min/half-ladder is in average. → 28 hours
Maybe 6 working days are enough.
Once the 2nd Felix operation system is available for these tests, we will migrate from the FEM/FEM-IB system to the new one.
Tested, very good
Tested, noisy
Tested, slightly noisy
North South
Mai Watanabe
(NWU)
Misaki Hata
(NWU)
: OK, but the results are with readout issues

## Slide 13: Summary
3rd test is ongoing now, mainly by M. Hata and M. Watanabe.
For the moment, all tested ladders showed very good results like in the test bench.
In the best cases, testing 1 half-ladder takes ~10 min. If we test all ladders, there are 2×(12 + 12 + 18) = 84 ladders, it takes about 14 hours. 20 min/half-ladder is in average. → 28 hours Maybe 6 working days are enough.
Once the 2nd Felix operation system is available for these tests, we will migrate from the FEM/FEM-IB system to the new one.
 | Date | Tested layers | Results
1st test | June/10-June/14 | B0L0 | Good.
2nd test | June/17-July/07 | B0L1, some of B0L0 | Some ladders in B0L1 were noisy. GND condition was investigated.
3rd test | Aug/10- now | B1L0 (some of) B0L1&B0L0 | Ongoing Tested ladders showed good results.
4th test | later… |  | 
Status of the 3rd test
Tested, very good
Tested, noisy
Tested, slightly noisy
North South

## Slide 14: backup

## Slide 15
Notation: BxLyzz
  x: Barrel ID (0 for inner or 1 for outer)
  y: Layer ID (0 for inner or 1 for outer)
  zz: Ladder ID (from 0 to 15)
Axis
  x-axis: y⃗ × z⃗
  y-axis: Vertically upward direction
  z-axis: The blue beam direction (pointing to the north)
West
B0L008
B0L009
B0L108
B0L109
B1L011
B1L012
B1L110
B1L111
x
(horizontal)
y (vertical)
z
East
B0L000
B0L001
B0L002
B0L003
B0L004
B0L005
B0L006
B0L007
B0L010
B0L011
B0L100
B0L101
B0L102
B0L103
B0L104
B0L105
B0L106
B0L107
B0L110
B0L111
B1L000
B1L001
B1L02
B1L003
B1L004
B1L005
B1L006
B1L007
B1L008
B1L009
B1L010
B1L013
B1L014
B1L015
B1L100
B1L101
B1L102
B1L103
B1L104
B1L105
B1L106
B1L107
B1L108
B1L109
B1L112
B1L113
B1L114
B1L115

## Slide 16: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L000
B0L006
North
South
ROC: NE4, A3
ROC: #29, C3
ROC: NE4, C1
ROC: #29, C3
West
barrel
East
barrel

## Slide 17: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L001
B0L007
North
South
ROC: NE2, C2
ROC: #29, C1
ROC: NE4, C2
ROC: #29, C3
West
barrel
East
barrel

## Slide 18: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L002
B0L008
North
South
ROC: NE2, A3
ROC: NE2, A3
ROC: NE2, C3
ROC: NE2, C2
West
barrel
East
barrel

## Slide 19: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L003
B0L009
North
South
ROC: NE2, C3
ROC: NE2, C3
ROC: NE2, A3
ROC: NE2, A2
West
barrel
East
barrel

## Slide 20: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L004
B0L010
North
South
ROC: #29, C3
ROC: NE4, C2
ROC: #29, C1
ROC: NE4, C2
West
barrel
East
barrel

## Slide 21: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L005
B0L011
North
South
ROC: #29, C3
ROC: #29, C3
ROC: NE4, C3
ROC: NE4, C3
West
barrel
East
barrel

## Slide 22: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L100
B0L106
North
South
ROC: NE2, A1
ROC: #29, D2
ROC: #29, B1
ROC: #29, C2
West
barrel
East
barrel

## Slide 23: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L101
B0L107
North
South
ROC: NE2, C3
ROC: #29, A1
ROC: #29, C3
ROC: #29, C3
West
barrel
East
barrel

## Slide 24: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L102
B0L108
North
South
Chip12 always has only half entries
ROC: NE2, C3
ROC: NE2, A3
ROC: NE2, C3
ROC: NE2, C2
West
barrel
East
barrel

## Slide 25: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L103
B0L109
North
South
ROC: NE2, C3
ROC: NE2, C2
ROC: NE2, C3
ROC: NE2, A2
Chip14 always has only half entries or strange results
West
barrel
East
barrel

## Slide 26: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L104
B0L110
North
South
ROC: NE2, C3
ROC: NE2, C3
ROC: NE2, C3
ROC: NE2, A3
West
barrel
East
barrel

## Slide 27: The barrel calibration tests: Results
0
1
2
3
4
5
5
4
3
2
1
0
6
7
8
9
10
11
11
10
9
8
7
6
B0L105
B0L111
North
South
ROC: NE2, C3
ROC: NE4, C2
ROC: NE4, C1
ROC: NE4, C2
West
barrel
East
barrel
→ a very good improvement obtained today