## Slide 1: Path to an Economical Low-Power Sub-psec Resolution ASIC
D. Braga, F. Fahim, N. J. Pastika, P. M. Rubinov, T. N. Zimmerman
N. Sullivan
H. J. Frisch, M. Heintz, E. Oberla, J. Park, C. Poe, F. Tang
G. S. Varner
B. W. Adams
C. Ertley
E. Angelico
1. Fermi National Accelerator Laboratory 2. Angstrom Research, Inc. 3. Enrico Fermi Institute, the University of Chicago 4. Univ. of Hawaii 5. Quantum Optics Applied Research 6. Southwest Research Institute 7. Stanford University
7
6
5
4
3
2
1
Notes: Hi everyone. I’m Jin, and we will discuss why fast timing electronics is the future and what we have done and doing right now in the field.

## Slide 2: Why Fast Timing?
Notes: First, I want to highlight the importance of fast timing electronics.

## Slide 3: Why Fast Timing? Explained through Time of Flight (ToF) system
W. Badgett et al., Precision Time-of-Flight at the Fermilab Testbeam Facility (2018)
Notes: Look at time of flight system as an example.
When the speed of the particle is close to the speed of light, resolving small difference between v and c determines the momentum of the particle.
You can see what time scale we are on from the bottom plot.
We will look a few exemplary detector systems which can utilize this high resolving power.

## Slide 4: An Electron-Ion Collider Study(2019)
EIC is the local particle collider at BNL with of .
Notes: In the Electron-Ion collider in Brookhaven, the momentum resolution of the detector system will be at best 1% at 50 Giga electron volts. Looking back at our diagram of Pion and Kaons, for example, this resolution indeed corresponds to a region of sub picosecond time resolution. So it is evident that we need a fast timing electronics corresponding to the high performance detector system.

## Slide 5: LHCb Detector Performance(2014)
The LHCb collaboration, https://arxiv.org/pdf/1412.6352.
Notes: LHCb is another example of detector system with high momentum resolution, again at best 1% at 50 Giga electron volts and higher. And as I said, this indicates that we need a sub picosecond fast timing electronics.

## Slide 6: Fermilab Test Beam Facility
E. Angelico, Development of Large-Area MCP-PMT photo-detectors for a Precision Time-of-Flight System at the Fermilab Test Beam Facility(2020)
Further development of the system is now taking place in hands of Joe Pastika.
Notes: Fianally, Fermilab test beam facility is the testbed of our chips and detector systems pushing fast timing boundaries.
We had set up a prototype time of flight system using our fast timing chip, psec4, and micro channel plate photomultiplier tube based detector which we call LAPPDs, the setup is shown in the diagram.

## Slide 7: PSEC4 ASIC Measurement and Calibration
Notes: Now, I will briefly discuss the structure of the fast timing chip we have right now.

## Slide 8: PSEC4 ASIC(2014)
E. Oberla , J.-F. Genata, H. Grabasa, H. Frisch, K. Nishimura, G. Varner, arxiv.org/abs/1309.4397
12-bit depth
6 channel/chip
5ps resolution
130ns IBM 8RF
Notes: This is a brief structure of ASIC we have been using for almost a decade now. We will focus on the components within the green dotted line because they determine the performance of the chip.
The signal goes into the chip from the left. First, its DC component is removed and set to a constant voltage, which can be controlled by an FPGA for example.
Then there is an array of 12 bit registers, the number of registers is equal to the total sample number. I will further explain each components of interest through next couple of slides.

## Slide 9: Sample time calibration
Time between samples are determined by a VCDL.
Each element of the VCDL has different time offset.
Notes: Time difference between samples are determined by a voltage-controlled delay line, which is an array of transistors, usually FETs, which a signal propagates through the array at a desired rate.
However, no chip manufacturing process is perfect, so each element of the VCDL has different time offset.
It is necessary to measure each time offset to correctly sample a signal.

## Slide 10
1. Measure many events of a sine wave.
2. Sum vs. Difference of two adjacent samples, over all events, form an ellipse.
3. Time offset can be determined from the coefficients.
determined by eccentricity and rotation.
10
Chart: 
 | 0.5 | 0.63049 | 0.735 | 0.79272 | 0.79215 | 0.73342 | 0.62821 | 0.49748 | 0.36724 | 0.26344 | 0.20674 | 0.20843 | 0.26817 | 0.37407 | 0.50504
Chart: 
 | 0.79272 | 0.79215 | 0.73342 | 0.62821 | 0.49748 | 0.36724 | 0.26344 | 0.20674 | 0.20843 | 0.26817 | 0.37407 | 0.50504 | 0.63501 | 0.7381 | 0.79378
Chart: 
 | 0.62821 | 0.49748 | 0.36724 | 0.26344 | 0.20674 | 0.20843 | 0.26817 | 0.37407 | 0.50504 | 0.63501 | 0.7381 | 0.79378 | 0.79097 | 0.73022 | 0.62364
Chart: 
 | 0.26344 | 0.20674 | 0.20843 | 0.26817 | 0.37407 | 0.50504 | 0.63501 | 0.7381 | 0.79378 | 0.79097 | 0.73022 | 0.62364 | 0.49243 | 0.36274 | 0.26037
Input Voltage
Time
Event 1
Event 2
Event 3
Event 4
Noise  fainter, more spread curve.
Nonlinearity  distorted curve
Notes: We have developed a method to measure time offset of each sample.
1. Measure many events of a sine wave.
2. Plot Summation vs. Difference of two adjacent samples, over all events, should form an ellipse.
3. Time offset can be determined from the coefficients.
The diagram explains which event corresponds to which dot in the ellipse plot.
When you think about it, the plot does not have to be on two adjacent samples, but samples which are m samples apart. This way the effect of voltage noise of each sample on the time offset is smaller.
In other words, the sampled curve is more like a smooth ellipse as m gets larger.
However, the time difference cannot be larger than one fourth of the period of the sine wave, which in our case corresponded to m=10.
When we look at the right ellipse plot,
Noise: fainter, more spread curve.
Nonlinearity: distorted curve

## Slide 11: Voltage linearization
It is necessary to reconstruct voltage from ADC value. The curve differs for each capacitor.
1. V_ped is set to a voltage of interest, 0.0v-1.2v
2. PSEC4 is left in sampling mode for a few cycles, so that all capacitors are set to V_ped. ()
3. Each register is read. Each curve corresponds to a capacitor.
Linearizable Domain
~0.75V
Notes: What bothered us next is the fact that the capacitors, shown in a green circle in the diagram on the right, lose or gain charges during the short time between signal input and measurement, so the 12-bit number from each ADC and the input voltage appear not linear.
It is necessary to reconstruct input voltage value from 12-bit ADC value. The curve is a differing characteristic for each capacitor in the array.
The way we measured this curve is very straightforward.
1. V_ped, an independent variable which is represented by x axis in the plot, is set to a voltage of interest, 0.0v-1.2v
2. PSEC4 is left in sampling mode for a few cycles, so that all capacitors are set to V_ped. The signal to noise ratio in this case is 0.5 millivolts.
3. Each register is read. Each curve corresponds to a capacitor.

## Slide 12: Performance Measurement
Same signal fed into two different channels of the same readout board.
1. measured for many events.
2. Distribution of acquired.
Depends on the source
  Sin
  Pulse/square
Depends on algorithm determining .
PSEC4
Signal
Ch. 1
Ch. 2
Trigger
Chart: Ch. 1 Voltage[V]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.375 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.5 | 0.4625 | 0.415 | 0.375 | 0.3225 | 0.2775 | 0.2325 | 0.2125 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Chart: Ch. 1 Voltage[V]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.35 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.6375 | 0.78863 | 0.92388 | 0.9 | 0.8875 | 0.8625 | 0.825 | 0.82977 | 0.55901 | 0.5 | 0.36701 | 0.24243 | 0.14643 | 0.13815 | 0.21785 | 0.33682 | 0.28204 | 0.2259 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Chart: Time[ns]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.375 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.5 | 0.4625 | 0.415 | 0.375 | 0.3225 | 0.2775 | 0.2325 | 0.2125 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Chart: Time[ns]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.35 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.6375 | 0.78863 | 0.92388 | 0.9 | 0.8875 | 0.8625 | 0.825 | 0.82977 | 0.55901 | 0.5 | 0.36701 | 0.24243 | 0.14643 | 0.13815 | 0.21785 | 0.33682 | 0.28204 | 0.2259 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
t=0 when the VCDL loop starts
Notes: After the formentioned calibration process,
We finally want to measure the performance of the chip, and one way to do so is to measure *time distribution of the sampled signal when sampling the same waveform many times.*
We set up an experiment where Same signal is fed into two different channels of the same readout board.
Then, delta t is measured for each event the way described figure on the right side.
1. measured for many events.
2. Distribution of acquired.
One should be wary however, because delta t depends on the source and the algorithm determining where the pulse starts.
We want to measure the performance of the chip, and one way to do so is to measure time distribution of the sampled signal when sampling the same waveform many times.
We set up an experiment where Same signal is fed into two different channels of the same readout board.
Then, delta t is measured for each event the way described figure on the right side.
1. Δ𝑡 measured for many events.
2. Distribution of Δ𝑡 acquired.
One should be wary however, because delta t depends on the source and the algorithm determining where the pulse starts.

## Slide 13
[ps]
# of Events
Chart: 
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.375 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.5 | 0.4625 | 0.415 | 0.375 | 0.3225 | 0.2775 | 0.2325 | 0.2125 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Chart: Ch. 1 Voltage[V]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.35 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.6375 | 0.78863 | 0.92388 | 0.9 | 0.8875 | 0.8625 | 0.825 | 0.82977 | 0.55901 | 0.5 | 0.36701 | 0.24243 | 0.14643 | 0.13815 | 0.21785 | 0.33682 | 0.28204 | 0.2259 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Chart: 
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.375 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.5 | 0.4625 | 0.415 | 0.375 | 0.3225 | 0.2775 | 0.2325 | 0.2125 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Chart: Time[ns]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.35 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.6375 | 0.78863 | 0.92388 | 0.9 | 0.8875 | 0.8625 | 0.825 | 0.82977 | 0.55901 | 0.5 | 0.36701 | 0.24243 | 0.14643 | 0.13815 | 0.21785 | 0.33682 | 0.28204 | 0.2259 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
Notes: We have acquired timing resolution below 5ps through this method. This histogram displays ten thousand events and the standard deviation is about 3 picoseconds.
Delta T histogram

## Slide 14: Setup over a distance
Two channels of each chip is used: one for signal, one for synchronization sine wave.
Sync. across far distance is a component of the total timing resolution.
PSEC4
Detector 1
Trigger
PSEC4
Detector 2
Trigger
WR
WR
Phase Sync’d
Sine wave
1ps
Synchronization via optical fiber
Ch. 1
Ch. 2
Ch. 1
Ch. 2
>100m
F. Girela-Lopez, J. Lopez-Jimenez, M. Jimenez-Lopez, R. Rodriguez, E. Ros and J. Díaz,
IEEE 1588 High Accuracy Default Profile: Applications and Challenges in IEEE Access, vol. 8, pp. 45211-45220, 2020
Notes: Same can be done for signal sources a long distance apart, for detector systems like time of flight system.
Two channels of each chip is used: one for signal, one for synchronization sine wave. So in total four channels are used.
The sync. Sine wave is generated by special module, called white rabbit, which outputs phase synchronized sine waves Over long distance via optical fiber. The phase jitter of this sine wave is way below 1ps for our application.
Sync. across far distance is a component of the total timing resolution of detector systems like this.

## Slide 15
t=0 when the sine wave crosses zero.
Chart: Detector 1 Voltage[V]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.375 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.5 | 0.4625 | 0.415 | 0.375 | 0.3225 | 0.2775 | 0.2325 | 0.2125 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
 | 0.5 | 0.63049 | 0.735 | 0.79272 | 0.79215 | 0.73342 | 0.62821 | 0.49748 | 0.36724 | 0.26344 | 0.20674 | 0.20843 | 0.26817 | 0.37407 | 0.50504 | 0.63501 | 0.7381 | 0.79378 | 0.79097 | 0.73022 | 0.62364 | 0.49243 | 0.36274 | 0.26037 | 0.20572 | 0.20966 | 0.2714 | 0.37867 | 0.51009 | 0.6395 | 0.74114 | 0.79476 | 0.7897 | 0.72695 | 0.61902 | 0.48739 | 0.35827 | 0.25737 | 0.20478 | 0.21097 | 0.2747 | 0.3833 | 0.51513 | 0.64394 | 0.7441 | 0.79566 | 0.78835 | 0.72362 | 0.61438 | 0.48235 | 0.35385 | 0.25444 | 0.20393 | 0.21236 | 0.27807 | 0.38796 | 0.52016 | 0.64835 | 0.747 | 0.79647
Chart: Time[ns]
| 0 | 0.03 | 0.06 | 0.09 | 0.12 | 0.15 | 0.18 | 0.21 | 0.24 | 0.27 | 0.3 | 0.33 | 0.36 | 0.39 | 0.42 | 0.45 | 0.48 | 0.51 | 0.54 | 0.56999999999999995 | 0.6 | 0.63 | 0.66 | 0.69 | 0.72 | 0.75 | 0.78 | 0.81 | 0.84 | 0.87 | 0.9 | 0.93 | 0.96 | 0.99 | 1.02 | 1.05 | 1.08 | 1.1100000000000001 | 1.1399999999999999 | 1.17 | 1.2 | 1.23 | 1.26 | 1.29 | 1.32 | 1.35 | 1.38 | 1.41 | 1.44 | 1.47 | 1.5 | 1.53 | 1.56 | 1.59 | 1.62 | 1.65 | 1.68 | 1.71 | 1.74 | 1.77
 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2125 | 0.2325 | 0.2775 | 0.3225 | 0.375 | 0.415 | 0.4625 | 0.5 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.525 | 0.5 | 0.4625 | 0.415 | 0.375 | 0.3225 | 0.2775 | 0.2325 | 0.2125 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2
 | 0.5 | 0.63049 | 0.735 | 0.79272 | 0.79215 | 0.73342 | 0.62821 | 0.49748 | 0.36724 | 0.26344 | 0.20674 | 0.20843 | 0.26817 | 0.37407 | 0.50504 | 0.63501 | 0.7381 | 0.79378 | 0.79097 | 0.73022 | 0.62364 | 0.49243 | 0.36274 | 0.26037 | 0.20572 | 0.20966 | 0.2714 | 0.37867 | 0.51009 | 0.6395 | 0.74114 | 0.79476 | 0.7897 | 0.72695 | 0.61902 | 0.48739 | 0.35827 | 0.25737 | 0.20478 | 0.21097 | 0.2747 | 0.3833 | 0.51513 | 0.64394 | 0.7441 | 0.79566 | 0.78835 | 0.72362 | 0.61438 | 0.48235 | 0.35385 | 0.25444 | 0.20393 | 0.21236 | 0.27807 | 0.38796 | 0.52016 | 0.64835 | 0.747 | 0.79647
Notes: Again, this is how we determine delta t and t=0.
As I said, different algorithms can be used for different waveforms. In this case, the variation of delta t will increase by factor of 2, compared to close distance setup.

## Slide 16: Future <1ps chip
Notes: Moving toward the end of this talk, I want to assert that now is the right time for a 1 picosecond timing chip.

## Slide 17: Time is right for a 1-psec timing chip
  Must optimize three aspects:
    Signal rise time
    voltage noise of samples
    sampling time jitter/uncertainty
Adopted from Stefan Ritt “The role of analog bandwidth and signal-to-noise in timing for waveform digitizing“
Notes: We hope the new chip is better in three aspects we have discussed before in this talk.
    Signal rise time
    voltage noise of samples
    sampling time jitter/uncertainty
    As we can see from the equation at the bottom, these three are the factors contributing to the overall timing resolution.
    And we will see that the improvements can be achieved by using more modern chipmaking processes.

## Slide 18
Great deal of work needs to be done but clear that 65nm CMOS technology superior characteristic to 130nm and older
Smaller process nodes give intrinsically faster transistors
Alessandra Pipino, Design of Analog Circuits in 28nm CMOS Technology for Physics Applications (2016)
Some selected results from literature:
(Romanova and Barzdenas: “A Review of Modern CMOS Transimpedance Amplifiers for OTDR Applications”, 2019)
Improved signal rise time
Improved or equal voltage noise of samples
Notes: Great deal of work needs to be done but
One thing sure is that smaller process nodes give faster transistors.
This implies potentially
Improved signal rise time
Improved or equal voltage noise of samples

## Slide 19
Timing uncertainty reduced by an optimized LC-oscillator based PLL
J. Prinzie, J. Christiansen, P. Moreira, M. Steyaert and P. Leroux, "Comparison of a 65 nm CMOS Ring- and LC-Oscillator Based PLL in Terms of TID and SEU Sensitivity," in IEEE Transactions on Nuclear Science,
vol. 64, no. 1, pp. 245-252, Jan. 2017.
From abstract: “Both independent PLLs have identical loop dynamics to allow a fair comparison … Furthermore these circuits consume the same amount of power. The PLLs were processed in a commercial 65 nm CMOS technology.”
Ring: 5.6ps jitter
LC: 0.325ps jitter
Notes: As I said before, we are currently using ring osc for the phase locked loop for the Voltage controlled delay line
This can be replaced to LC oscillator based phased locked loop to achieve better timing uncertainty.

## Slide 20: Improved packaging also important
Chi-Tsung Chiu, Sung-Mao Wu and Chi-Ping Hung, "High speed electrical performance comparison between bump with RDL and wire bond technologies," Proceedings of the 4th International Symposium on Electronic Materials and Packaging, 2002., 2002
 | (ps) observed for 35ps risetime input pulse | Degradation (%)
Case A (wirebonded) | 65ps | 86%
Case B (flip-chip + RDL) | 42ps | 20%
 |  | Degradation (%)
Case A (wirebonded) | 65ps | 86%
Case B (flip-chip + RDL) | 42ps | 20%
Notes: Also, we can improve the rising time of the pulse through better packaging, since shorter trace length results in lower degradation of rising time.

## Slide 21: Time is right for a 1-psec timing chip!
Notes: These three combined, here is the final table of specifications we expect for the new chip.
One final assertion of our team is that
1-psec timing chip is possible and it is the right time for it.

## Slide 22: Nonlinearity

## Slide 23: Citations(All from the doc lib)
Preliminary Specifications for a 1-Picosecond Resolution ASIC for the FTBF TOF System
A Modular Data Acquisition System using the 10 GSa/s PSEC4 Waveform Recording Chip
DEVELOPMENT OF LARGE-AREA MCP-PMT PHOTO-DETECTORS FOR A PRECISION TIME-OF-FLIGHT SYSTEM AT THE FERMILAB TEST BEAM FACILITY
Measurement Noise and Time-Base Calibration in ACDC Data