## Slide 1: Pixel-based accelerated aging of Large Area Picosecond Photodetectors (LAPPDTM)
UTA Picosecond Test Lab
V. A. Chirayath & A. Brandt
LAPPD Virtual Workshop
October 26, 2022

## Slide 2: Lifetime of Microchannel Plate Photomultiplier Tubes (MCP-PMT)
A fraction of incoming photons are converted to photo-electron (PE) by the photocathode (Pk) according to its quantum efficiency.
The charge is multiplied by the collision of the incoming PE with the walls of the MCP. Amplified charge is collected at Anode.
Collision can sometime lead to ionization of the residual gas/desorption of positive ions from the MCP surface: Ion feedback
Ions can react with or even sputter the Pk material – Degradation of Quantum efficiency
MCP Gain can also be affected by the desorption of ions from MCP surfaces.

## Slide 3: Lifetime of Microchannel Plate Photomultiplier Tubes (MCP-PMT)
Suppress the release of positive ions through Atomic Layer Deposition (ALD) of additional layers of Al2O3 or MgO in the inner surfaces of MCP micropores. Main challenge is optimizing the coating process to get the desired lifetime.
Blocking the ions from reaching the photocathode once they are created (active or passive ion barriers).

## Slide 4: Investigation of Aging of MCPs with Enhanced Lifetime
Pioneering work in simultaneously comparing the aging behavior of different kinds of MCPs by Lehmann et al [1]
Diffuse 460 nm LED to homogeneously illuminate the whole MCP at single PE rate. The LED pulsed at 1 MHz.
MCP-PMT with ALD coated MCP's by Arradiance (Arradiance-UTA SBIR) is first tube with >10 C/cm2 lifetime
Photonis MCP PMT with resistive and emissive ALD coating showed lifetime exceeding 26 C/cm2.
Extended period of irradiation to achieve integrated anode charge providing clear feedback on absolute detector lifetime.
1. Lehmann et al., Nuclear Inst. and Methods in Physics Research, A 958 (2020) 162357.
Fig. 1 . Results from the Lifetime testing conducted by Lehmann et al., NIMA (2020) [1]. The Quantum efficiency as a function of integrated anode charge for ALD coated MCP-PMTs. The inset in top panel for non-ALD MCP-PMTS.
10 C/cm2
26 C/cm2

## Slide 5: Pixel-based Lifetime testing
The pixel-based lifetime time testing method: Exposes a pixel of the MCP-PMT to high photon rate.
Damage is primarily a local phenomenon.
Key advantages:
The MCP-PMT can be reused after testing.
Multiple lifetime tests can be performed on a single device.
Figure: PE scan of a 5 cm x 5cm Photonis MCP-PMT. Comparison of the estimated photo electrons (PE) before (Pre-LT) and after (Post-LT) the lifetime test for same incoming light intensity.
5 cm
Notes: Reduce the number of words!!

## Slide 6: Accelerated Aging
Figure: Measured increase (blue) in output current with increasing LASER frequency (in expressed in kHZ) compared to expected increase (orange) when employing (a) two MCPs at maximum gain and (b) employing only the top MCP. The input light intensity is the same in both cases. In second case the second MCP is used as the anode.
(a)
(b)
Laser Rate (kHz)
Laser Rate (kHz)

## Slide 7: Entry MCP plays critical role in after-pulse generation
Figure: (a) Pulse height distribution as a function of LASER repetition rate (demonstrates saturation) (b) After Pulse distribution measured under similar conditions as a function of repetition rates. Even though the detector is saturated, the after-pulse rate remains same.

## Slide 8: 20 cm x 20 cm LAPPDTM from Incom Inc.
Feature | Parameter
Seal Date | 02/04/2020/LAPPD-64
Photodetector Material | Borosilicate glass lower tile assembly Interior stripline anode
Photocathode Material | Alkali metal
Wavelength – Maximum Sensitivity (nm) | ≤ 365 nm
Anode Data Strip Configuration | 28 silver strips, Width = 5.2 mm, Nominal 50 ‎Ω Impedance
MCP Substrate | Incom C14/C5 Glass
Resistive and Emissive Coatings | Chem 1, Applied via Atomic Layer Deposition (ALD)
Secondary Emission (SEE) Layer Material | MgO
MCP resistance (Entry/Exit) | 13.1 / 4.1 MΩ at 925 V A

## Slide 9: Pre-Lifetime Characterization : Dark box with Pulsed LASER
Device characterization using pulsed LASER (405 nm). The exposure arrangement is equivalent to that in Lifetime test box with LED.
Main Tests
Single PE pulse height distribution
Transit Time Spread
Behavior under high-rate exposure
After pulse characterization
Lifetime Testing of the MCP-PMT is performed in a dedicated dark box with LED
LASER
ND Filters
Mirrors
Fiber Holders
MCP-PMT

## Slide 10: Gain Distribution and Transit Time Spread
Figure: (a) Gain and (b) Transit time spread as a function of rate. Performance measurements were carried using fiber arrangement which leads to ~ 4.6 mm beam spot. Saturation happens ~ at 500 kHz for a beam spot of diameter 4.6 mm. Measurements performed with 925 V across the MCPs, 100 V on Photocathode.
(a)
(b)

## Slide 11: Lifetime Testing Rate
LAPPD-64 can achieve a gain of ~ 3 × 106 at a rate of 500kHz & with a beam spot of diameter ~ 4.6 mm.
Output Current that can be extracted from LAPPD 64:
Iout per unit area = Gain × No. of PE × Charge × Rate per unit area
=
i.e. 0.124 C/cm2 per day or ~ 1 C/cm2 in 8 days

## Slide 12: Pixel based Lifetime Testing setup
LED
ND filter
Inline ND filter
Si Photodiode
Tube under test
LED is operated in continuous mode (CW) for irradiation and pulsed mode (PW) for evaluation.
LED illuminates 4 fiber bundle
Allows pixelized exposure of up to 4 devices with inline neutral density (ND) filters to control the light intensity to each device.
A Si Photodiode for independent monitoring of light intensity
Fiber bunch

## Slide 13: Pixel-Based Lifetime Testing Algorithm
Measure the pulse Height distribution at single PE level in the LED box at non saturation level (200 Hz) before the test.
Irradiate a small region of 4.6 mm diameter with MCP-PMT close to saturation.
Measure the pulse height distribution at single PE level in the LED box at 200 Hz at regular intervals by stopping the irradiation.

## Slide 14: Pulse Height distribution as a function of collected charge
Figure: Pulse height distribution as a function of the charge collected at the anode.

## Slide 15: MCP-PMT performance as a function of collected charge
Figure: Average pulse height calculated for pulses above threshold (15 mV). Sensitive to Gain variations.
Figure: Fraction of pulses above threshold (15 mV). Sensitive to QE variations.

## Slide 16: Comparison of pre and post test performance using pulsed LASER

## Slide 17
Comparison of pre and post test performance QE Scan @ Incom Inc.
QE Scan : Before shipping to UTA. Blue square shows the irradiation spot. Spot size is to scale.
QE Scan : After receiving the tile from UTA

## Slide 18
PC
MCP1
MCP2
Anode
100 V
925
925
200 V
LAPPD Electrode Arrangement
Spectra of After pulses
when only one after pulse occurs in coincidence with main pulse and LASER.
LASER Pulse
Ch1: One end of strip
Ch4: Other end of strip
Ch3: Unirradiated strip
Main Pulse
After Pulses
After Pulse measurement settings
Pulse Height Distribution of Main Pulse
At a gain of 1x107, the tube has an after-pulse rate of 5%

## Slide 19: Conclusions and Ongoing work
LAPPD-64 is a high gain MCP-PMT with good timing resolution and the lifetime tests has shown little degradation in performance till 5.6 C/cm2
The accelerated aging is a promising complementary method to the pioneering lifetime testing works - provide feedback in realistic time scales - appropriate changes to the design and chemistry.
An MCP-PMT with higher after pulse rate is in testing to generate a correlation between after pulse rate and lifetime.
Testing the performance of an MCP-PMT with modified photocathode chemistry that it is suitable for detecting Cherenkov spectrum from Aerogel (like in EIC) (interesting talks later by Cole and Alexey).

## Slide 20: Thank You
DOE Grant DE-SC0011686