## Slide 1: Magnetic Field Testing at Argonne National Laboratory
M. Popecki, Junqi Xie & Jack McKisson
February, 2023

## Slide 2: Two LAPPDs & One HRPPD at a Solenoid Magnet
Two capacitively coupled LAPPDs and One HRPPD
  144, 20 um, glass
  139, 20 um, ceramic
  One directly-coupled HRPPD: 6
Magnetic field strength: 0.02 T to 2.0 T
Dark box
  Aluminum case
  Laser input fixed in the center near the bottom – on the centerline of the solenoid when the LAPPD is vertical.
  Waveforms collected for 12 pixels (LAPPD) and 16 pixels (HRPPD D/C)
Rotation in the magnetic field:
  LAPPD tips into or out of the region of stronger magnetic field
  Move the LAPPD in or out at each angle to compensate for the change in field strength
Data products
  Gain
  Position
  Position resolution
  Transit Time Spread
  Afterpulse rate
  Detection efficiency estimate
  Dark rates

## Slide 3: Dark Rates vs. Field Magnitude
Dark rates measured using:
  Two pulse counters – Jorway counter and an Ortec 778
  BK 2559 oscilloscope with trigger crossing counter
Dark counts decline with increasing magnetic field strength
A persistent count rate peak was observed at one field strength.
The response was isolated to a narrow range of field strength, much like a resonance.
LAPPD144 02-15-2023.xlsx, see AF78

## Slide 4: Dark Rates vs. Photocathode Voltage
Dark rates increased with photocathode voltage.
Rates are lower at 1.69 T than at the Earth’s field.
Is the magnetic field reducing the emission from the photocathode more than reducing the gain?
-11 mV threshold
1.69T
-4 mV threshold
Earth’s B field