## Slide 1: DAC Scan
RIKEN/RBRC
Itaru Nakagawa
July 6, 2023

## Slide 2: DAC Scan Status
1st attempt on June 4~5th Run#9301 ~ 9335.
Condition: 6 servers timed in (except for intt0, 1), n_collision=4.
Yuka’s Analysis
2 hits cluster
https://indico2.riken.jp/event/4503/contributions/21289/attachments/12076/17540/20230705_sugiyama.pdf
2021 Tohoku Beam test
MIP peak isn’t too clear yet. Z-vertex cut may be mandatory.
1 hit cluster
4 hits cluster
3 hits cluster
Large background?

## Slide 3: DAC Scan
Scan | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
DAC0 | 8 | 28 | 48 | 68 | 88 | 108 | 128 | 148 | 168 | 188 | 212 | 236
1 | 12 | 32 | 52 | 72 | 92 | 112 | 132 | 152 | 172 | 192 | 216 | 240
2 | 16 | 36 | 56 | 76 | 96 | 116 | 136 | 156 | 176 | 196 | 220 | 244
3 | 20 | 40 | 60 | 80 | 100 | 120 | 140 | 160 | 180 | 200 | 224 | 248
4 | 24 | 44 | 64 | 84 | 104 | 124 | 144 | 164 | 184 | 204 | 228 | 252
5 | 28 | 48 | 68 | 88 | 108 | 128 | 148 | 168 | 188 | 208 | 232 | 255
6 | 32 | 52 | 72 | 92 | 112 | 132 | 152 | 172 | 192 | 212 | 236 | 255
7 | 36 | 56 | 76 | 96 | 116 | 136 | 156 | 176 | 196 | 216 | 240 | 255
BigPartition together with MBD (Must) no need to be a dedicated run
Can be done with n_collision=127 (w/o waiting for asynchronous timing issue btwn intt servers.
12 settings
> 1M events at ~400Hz
~ 12 hours total
If the series of data are interrupted by the beam dump, repeat the same setting as the last run at the last store.
Extend to max range

## Slide 4: Instruction for a DAQ operator
DAC setting is to be changed by a DAQ operator by running following script: ~/INTT/sphenix_inttpy/run_scripts/INTT_dac_scan.sh [scan number]
Example:
phnxrc@operator1$ bash INTT_dac_scan.sh 1
The script changes DAC setting “scan 1” and loads parameters to FPHX (execute run.py)
Make an entry in INTT e-log for scan number and corresponding run number.
cat fphx_parameters_DAC_scan_01.txt
Vref 1
DAC0 8
DAC1 12
DAC2 16
DAC3 20
DAC4 24
DAC5 28
DAC6 32
DAC7 36
N1sel 6
N2sel 4
FB1sel 4
Leaksel 0
P3sel 0
P2sel 4
Gsel 2
BWsel 8
P1sel 5
Injsel 0
LVDS 63
To be checked the script still works, since the last use was a month ago.
Jaein