## Slide 1: Results from STA/ITL Fully Depleted CCDs for LSST
Michael Lesser & David Ouellette
University of Arizona
Imaging Technology Laboratory
Steward Observatory
lesser@itl.arizona.edu

## Slide 2: Overview
Sensor Description
  STA3800 Science Sensors
  STA4400 Wavefront Sensors
Electro-Optical Characterization
  Measurement Processes
  Results
Metrology Characterization
  Measurement Processes
  Results
Sensors developed for LSST during multiple project phases
Test specifications defined by LSST project

## Slide 3: STA3800 Science Sensors
4072 x 4000 10 um pixels
16 video channels (8x2)
Custom package for LSST
  sensor hybridized to AlN ceramic with connectors, RTD, 1-wire LAN
  hybrid mounted to CE5 Al-Si allow
  alignment pins and studs installed in CE5
STA3800 sensors in shipping containers
packaged sensor exploded view

## Slide 4: STA3800 Science Sensors
“Lithoblack” added to wafer frontside to reduce IR “print-thru” reflections off ceramic metal films
Select masks rewritten using the highest resolution mask writer available to reduce fixed patterns
without/with “Lithoblack” - 1000 nm
without/with high-res masks

## Slide 5: STA4400 Wavefront Sensors
4072 x 2000 10 um pixels
8 video channels (8x1)
Custom package for LSST
High resolution masks used
Lithoblack applied to wafers
packaging jig
STA4400 packaged sensor on test tower

## Slide 6: Electro-Optical Characterization
Three LSST production systems and one test system
  All operate STA3800 or STA4400 sensors
  LN2 for rapid cooling
  Internal heaters for rapid heating
  Internal Fe55 x-ray sources for calibration
  Receives LSST test tower without dismounting sensor
Dedicated lab for LSST
  Main ITL characterization lab also available for tests/development
  Arc lamp, filters, fiber optic light pipe, integrating sphere with reference diode, baffled tube, CryoCon temperature controllers, dewar dock
  Fully automated acquisition using Python scripts

## Slide 7: Electro-Optical Characterization
LN2 dewars on carts
STA Archon controllers
ARC Gen3 controller
turbo pumps

## Slide 8: Electro-Optical Characterization
LSST Testing Lab
Light source + filters
Integrating sphere + baffled tube

## Slide 9: Electro-Optical Characterization
Three STA Archon controllers
  Internal -100 V backside bias control
  530 kHz operation
  Code file delivered with each sensor (typically not changed)
One Gen III ARC controller
  Used for development, testing, calibrations
Acquisition script (automated)
  Sets bias levels
  Calibrates flux at each test wavelength
  Waits for final temperature
  Acquires data
  Emails
  Begins warm-up process
~1 TB of data analyzed for following slides…

## Slide 10: Electro Optical Results
Five 10 minute darks per sensor
Dominate issue is bright column due to defects after etching
Very small number of sensors have “glows” at serial register sides
Very small number of sensors have “edge glows” related to backside bias
Dark Signal
Dark Signal (-95 C) : 54 sensors | 
Fraction < 0.05 e-/pix/sec | 
95.6 | 0.6
LSST spec: 95% pixels ≤ 0.2 e-/pix/sec

## Slide 11: Electro Optical Results
Data acquired with 10 nm band-pass filters
Wavelengths: 330, 350, 370, 450, 500, 620, 750, 870, and 1000 nm
QE reported for bands: u, g, r, i, z, y
Quantum Efficiency
 | Quantum Efficiency – 48 sensors @ -95C |  |  |  |  | 
Band | u | g | r | i | z | y
Mean QE | 0.639 | 0.886 | 0.904 | 0.970 | 0.920 | 0.328
 | 0.070 | 0.040 | 0.016 | 0.000 | 0.013 | 0.011
500 nm

## Slide 12: Electro Optical Results
500 nm

## Slide 13: Electro Optical Results

## Slide 14: Electro Optical Results
60 second Fe-55 images
Each amplifier section analyzed
Some software confusion from defects
Backside bias -35 to -50 V
  increased bias may increase defects
Charge Diffusion
Charge Diffusion “sigma” : 56 sensors | 
Mean sigma (µm) | 
4.73 | 0.4
LSST spec ≤ 5 µm

## Slide 15: Electro Optical Results

## Slide 16: Electro Optical Results
Measured with Extended Pixel Edge Response (EPER) method
High and low level flats (combined)
  5 x 50k e- and 50 x 1k e- images
No VCTE issues seem
Some HCTE issues…
Charge Transfer Efficiency

## Slide 17: Electro Optical Results
Some batches have shown poor HCTE (0.9999x)
Suspect process issue by a vendor
Poor HCTE channels also have “sloped” lead-in bias level
single channel fail
serious HCTE fail

## Slide 18: Electro Optical Results
Measured twice per sensor
  PTC @ ~1600 e-
  Fe55 overscan noise using Fe55-derived system gain
Read Noise
noise PASS
LSST spec < 8 e- rms

## Slide 19: Electro Optical Results
Sensors with poor HCTE usually have higher noise (~15 e-)
Read Noise
noise FAIL

## Slide 20: Electro Optical Results
fringing seen only @ 973 nm ± < 5 nm (internal lithoblack/epoxy)
Backside
column defects
Frontside
low bias: -5V bias
No phosphorescence or radioactivity found in AlN but vacuum gauge glow detected (no lithoblack)
gauge on
gauge off
Miscellaneous

## Slide 21: Electro Optical Results
“Tree Rings”
Backside bias voltage:
0 V to -70 V

## Slide 22: Metrology Characterization
VIEW Summit 600 CMM
Machined-glass mounting jig
Mounting torque of sensor important
  Added correction for equipment/lab temperature drift by measuring four standards on each “line” scan throughout sensor measurement process

## Slide 23: Metrology Characterization
Machined glass mounting jig

## Slide 24: Metrology Characterization
data products

## Slide 25: Metrology Characterization
Zygo NextVIEW 8000 3D Optical Surface Profiler
  surface studies & calibrations
  < 1 nm – 20 mm height
Analysis of sensors during processing
image area ±1 µm
cold surface flatness < 1 µm

## Slide 26: Metrology – Room Temp
STA3800 – means of a sample of 59 sensors
Mean (mm) | Median (mm) | Z-95 (mm) |  (µm)
12.9923 | 12.9924 | 12.9918 | 1.2
Flatness
F-95 (µm) | F-Range (µm) |  (µm)
1.80 | 6.28 | 0.8
Height
F-95 is flatness half-band (spec ± 5 µm)
Z-95 is height half-band (spec ± 9 µm)
F-Range is full range (max-min)

## Slide 27: Metrology – Room Temp
95% data (boxes) and 100% data (lines) – 32 sensors
95% limits

## Slide 28: Metrology – Room Temp
STA4400 – means of a sample of 14 sensors
Mean (mm) | Median (mm) | Z-95 (mm) |  (µm)
9.0036 | 9.0037 | 9.0034 | 1.3
Flatness
F-95 (µm) | F-Range (µm) |  (µm)
1.69 | 4.77 | 0.8
Height
F-95 is flatness half-band (spec ± 5 µm)
Z-95 is height half-band (spec ± 9 µm)
F-Range is full range (max-min)

## Slide 29: Metrology - Cold
Nanovea chromatic confocal optical pens mounted on large area X-Y scanner
  0.9 µm and 3.0 µm accuracy
LN2 cooled dewar
Sample spacing 1 mm
  Edge distance issue due to larger beam size and large working distance
Lab air/equipment temperature compensation included
ITL “CryoScanner”

## Slide 30: Metrology – Cold
When cooled from +23 to -95 C surface bows < 5 µm (center high)
STA3800

## Slide 31: Summary
Characterization systems allows testing two sensors per day
  Major issue is water vapor removal with short duty cycle
  Active heaters for warm-up are very helpful
STA3800 (16 channel) and STA4400 (8 channel) provide essentially identical performance – STA Archon controllers
  ~5 electrons @ 530 kHz
  Full well > 175,000 electrons (10 micron pixels)
  QE close to model
Metrology specification has been consistently met
  Precision packaging jig development and automated measurement critical
  Cold metrology shows <5 micron change @ -95 C
Processing difficulties
  Bright columns from defects after thinning
  Excessive noise/poor HCTE which was batch-specific
  Occasional damage (ESD?) by vendor processes such as backside grinding and the multiple “lithoblack” applications steps

## Slide 32: Acknowledgements
Thanks to entire staff of UA Imaging Technology Lab
Special thanks to Semiconductor Technology Associates, Inc.
  R. Bredthauer, K. Bredthauer, K. Boggs
Funding provided by SLAC/DOE/BNL/UofA