## Slide 1: Intrapixel Effects of CCD and CMOS Detectors
Hu Zhan, Li Cao, & Xin Zhang
National Astronomical Observatories
Chinese Academy of Sciences
PACCD, Dec 1-2, 2016, BNL

## Slide 2
The non-uniformity of CCDs and the effects of spatial undersampling
SPIE 1994, 2198, 836
BSI

## Slide 3: Intrapixel Effects and Undersampling
Zhou, X., et al., Modeling the CCD Undersampling Effect in the BATC Photometric System, PASP 2005, 117, 86
Mag. residuals of bright stars, unguided, one night
Ford Aerospace 2k×2k CCD FSI
pixel size: 15μm square
pixel scale: 1.7″/pix

## Slide 4: New Context: Precision Astronomy
Precision goal
  Photometry 0.005 mag
  Astrometry 0.5 mas
Ivezic et al., ARAA 2012, 50, 251

## Slide 5: New Context: Precision Astronomy
Weak lensing shear
  Induced ellipticity: 10-3 -10-2
  Mult. shear error: 10-3
  Add. shear error: 3×10-4

## Slide 6: New Dev’t: Thick CCDs & BSI CMOS-APS
DES CCD 250m
LSST CCD 100m
TAOS II BSI CMOS
GSENSE400 BSI CMOS
Precision astronomy places demanding requirements on detectors. Previously negligible effects need closer examination.

## Slide 7: Measuring Intrapixel Effects
Device | Model
LED light source | Thorlabs M365L2、M850L3
X-Y precision stages | Thorlabs LTS150、LTS300
Lens | Schneider-KREUZNACH XENON-E 2.2/50 AF Nikkor 50mm f/1.8D
Source on X-Y stages
Notes: 可以把像斑放大的照片放在这里，我没找到。

## Slide 8: Measuring Intrapixel Effects
LED |  | 
model | M365L2 | M850L3
Peak Wavelength | 365nm | 860nm
Bandwidth (FWHM) | 7.5nm | 30nm
Precision Stage | 
Travel Range | 150 mm，300 mm
Minimum step size | 0.1 µm
Bidirectional Repeatability | 2 µm
Backlashe | 2 µm
Lens f=50mm
Cameras/detectors tested:
GSENSE400 BSI CMOS
Andor DV435 (CCD)
Andor Neo sCMOS
Andor DU934 (CCD)
LED Diffuser Pinhole
Notes: 可以把像斑放大的照片放在这里，我没找到。

## Slide 9: Results for Neo sCMOS & DU934
Flux variation
Ellipticity variation
 | Andor Neo sCMOS | Andor DU934
 | Front illuminated CMOS | e2v CCD 47-10
FWHM | 2.0 pix | 1.7 pix
Flux variation PV | 28% | 2.2%
Ellipticity var PV | 0.084 | 0.017
Notes: “CCD与CMOS20140403詹.docx”中以下描述是什么一丝？光阑在什么位置？像斑大小？CCD的流量变化达到2.2%和3% rms
使用sCMOS时，光阑大小设置为3μm（像元大小是光阑直径的2.2倍），平台沿sCMOS的y方向移动。像斑x方向的变化在一个像元内（0.12~0.24像元内），y方向移动了约6个像元。
使用CCD进行实验时，光阑大小设置为5μm（像元大小是光阑直径的2.6倍），平台分别沿x、y方向移动进行测试。

## Slide 10: GSENSE400BSI & DV435
Camera | GSENSE400 Evaluation Board (G400) | Andor DV435
Sensor | GSENSE400BSI Back illuminated CMOS | e2v CCD 47-20 Back illuminated CCD
Shutter model | Rolling | Mechanical shutter
Active pixels | 2048×2048 | 1024×1024
Pixel Size | 11μm×11μm | 13μm×13μm
Well depth | 32,000e- | 80,007e-
Readout Noise | 1.73 e-/pix | 3.02 e-/pix
Temperature | Room temperature | -40℃
FWHM ≈ 2.6 pixels
Step size: 1/8~1/10 pix
Length of scan: 4~5 pixels
CMOS spot image
16 frames combined
CCD spot image
10 frames combined
Notes: PSF size the radius enclosing 80% flux ≈2pixels——defocus
Step：1/8~1/10 pix
distance: 4~5 pixels

## Slide 11: Flux Variations for G400 & DV435
RMS：0.89%
RMS：0.20%
RMS：0.80%
RMS：0.22%
GSENSE400BSI
DV435: CCD47-20
365nm
850nm
Notes: 张鑫：CMOS 850nm shot noise=1.2‰（16帧合并）， CCD 850nm shot noise=0.8‰(10帧合并)
右侧：CCD，365nm上，850nm下

## Slide 12: Ellipticity Variations for G400 & DV435
mean：0.020
RMS：0.0047
mean：0.010
RMS：0.0045
mean：0.038
RMS：0.0129
mean：0.014
RMS：0.0048
GSENSE400BSI
DV435: CCD47-20
365nm
850nm

## Slide 13: Centroid Variations for G400 & DV435
RMS：0.012pix
RMS：0.009pix
RMS：0.022pix
RMS：0.009pix
GSENSE400BSI
DV435: CCD47-20
365nm
850nm
Notes: 左侧：CMOS，365nm上，850nm下
右侧：CCD，365nm上，850nm下

## Slide 14: Recap for G400 & DV435
 |  | Unit | GSENSE400BSI |  | DV435: CCD47-20 | 
 |  |  | 365nm | 850nm | 365nm | 850nm
Flux variation (RMS) |  | % | 0.89 | 0.80 | 0.22 | 0.22
Ellipticity | Mean |  | 0.020 | 0.038 | 0.014 | 0.010
 | RMS |  | 0.0047 | 0.0129 | 0.0048 | 0.0045
Centroid variation (RMS) |  | pixel | 0.012 | 0.022 | 0.009 | 0.009
  Photometry 0.005 mag
  Astrometry 0.5 mas
  Shear 1% signal to 0.1%

## Slide 15: Other Issues with CMOS
CMOS detectors sometimes combine two 12bit A/D channels, one high gain and the other low gain, to achieve a 16bit dynamic range. The jump of the readout noise at the end of the high-gain channel is clearly seen. How does it affect astronomical measurements?
GSENSE400BSI: lower readout noise but wide distribution
DV435: higher readout noise and narrow distribution

## Slide 16: Summary
The level of CCD and CMOS intrapixel effects on spot (PSF) flux, centroid, and ellipticity are relevant to precision astronomy.
The CMOS detectors tested exhibit worse intrapixel effects than CCDs do.
Intrapixel effects depend very much on the detailed pixel structure and the pixel scale (PSF sampling). Therefore, one has to quantify the effects for each project.
Need for intrapixel effect correction!