## Slide 1: Germanium CCDs for Large-Format SWIR and X-Ray Imaging
C. Leitz, S. Rabe, I. Prigozhin, B. Burke, K. Ryu, M. Zhu, M. Cooper, K. Johnson, W. Hu, B. Felton, M. Cook, C. Stull, V. Suntharalingam
1 December 2016
This work was sponsored by the Assistant Secretary of Defense (ASD) for Research and Engineering (R&E) under Air Force Contract No. FA8702-15-D-0001. Opinions, interpretations, conclusions, and recommendations are those of the authors, and do not necessarily represent the view of the United States Government.
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 2: MIT-LL Silicon Charge-Coupled Devices (CCDs)
Exquisite sensitivity and dynamic range
Single electron read noise
Noiseless, efficient charge transfer enables time delay integration (TDI) operation, pixel binning
High spatial uniformity enables large-format imagers
PS1-2 (2007, 2015)
SST (2011, 2015)
Chandra (1999)
TESS (2015)
Can we realize these advantages in other bands?

## Slide 3: Advanced Imager Technology and MIT-LL Microelectronics Laboratory
Microelectronics Laboratory
Class-10 clean room
Production-class sub-90-nm toolset on 200-mm-diameter wafers
Lithography capabilities including large fields (50  50 mm field), fine features (e-beam)
Advanced Imager Technology
Suite of imaging technology spanning both Si and compound semiconductors
  Si CCDs, Geiger-mode APD arrays, and APS
  InGaAs linear- and Geiger-mode APDs
  Digital focal planes
In-house capabilities from design through sub-system fabrication

## Slide 4: Unique Positioning of MIT-LL Microelectronics Laboratory (ML)
Microelectronics Laboratory, with its unique capabilities, is ideally suited for development of germanium imaging devices.
Extreme-Low-Power CMOS
90 nm
Innovation
Process Maturity
Superconducting Electronics
1 mm
2 mm
Mature Device Understanding
Experimental Devices
Single Devices
100M’s Devices
Foundry Processes
University Research
Range of ML
Fabrication Processes
Large-Format CCDs
& APD Arrays
20 mm
Bond line
3D-Integrated Devices
Notes: VIPER ROIC = 500M transistors!

## Slide 5: Advantages of Germanium for Imagers
Visible + SWIR Quantum Efficiency Comparison [1-3]
X-Ray Sensitivity Comparison
Strong absorption from UV through short-wave infrared (SWIR)
X-ray absorption to 30 keV and beyond for high-energy imaging
[1] T. Martin & P. Dixon, Laser Focus World (11/2004); [2] MIT-LL measured; [3] average from Nakano et al., J. Non-Cryst. Sol. 358 (2012) 2249 & N. Posthuma et al., Proc. 3rd World Conf. Photo. Ener. Conv. (2003), scaled to 45 µm.

## Slide 6: Advances in Germanium Process Technology
Build on recent advances in complementary fields
  Next-generation electronics
  Concentrator photovoltaics and thermophotovoltaics
  Material quality – 200-mm bulk wafers commercially available
Enables fabrication of high-quality devices not possible in recent past
High-Quality Gate Stacks [1]
GeO2
Ge (100)
10 nm
2 nm
Low-Temperature Surface Passivation [2]
[1] T. Sasada, et al. J. Appl. Phys. 106 (2009) 073716[2] J. van der Heide, et al. Sol. Energ. Mat. Sol. Cells 93 (2009) 1810; http://www2.imec.be/be_en/press/imec-news/tpvcell.html.

## Slide 7: Comparison to Compound Semiconductor Imagers
The Ge CCD represents a potentially significant improvement over compound semiconductor imagers in terms of format, pixel size, and capability
  More advanced processes to enable larger formats, small pixel sizes
  No bump bonding required
One disadvantage: Germanium must be cooled due to high intrinsic carrier concentration
Wafer Size Comparison
Sensor Size Comparison
200 mm
Germanium
InGaAs (SWIR)
CZT (X-ray)
InGaAs (1.3 Mpix)
5 cm
Germanium
(> 20 Mpix expected)
NuSTAR
(1 kpix)

## Slide 8: Path to a Germanium CCD
Started with isolated components for quick-turn insights into CCD performance
CCD Viewed Along Charge-Transfer Direction
Vlow
Vhigh
Vlow
Vlow
Vlow
Vhigh
VOG
Output
1
2,3
4
100 nm
Germanium
GeO2
Metal
Al2O3
 | Device Component | Status
1 | Metal-Insulator-Semiconductor (MIS) capacitor | Surface state density rivaling that of silicon electronics, low gate leakage
2 | P-N junction diodes | Low dark current
3 | MOSFET | Functional output circuits
4 | Buried channel | Imagers being analyzed

## Slide 9: Metal-Insulator-Semiconductor (MIS) Capacitor Analysis
Excellent quality MIS capacitors realized
Room-temperature measurements on MIS capacitors confirm basic properties of Ge suitable for CCDs
Low surface state density
Gate leakage < 10 pA/cm2 for +/- 10V gate swing
Capacitance-Voltage
Gate Leakage
Surface State Density

## Slide 10: Diode Dark Current Analysis
Estimate of Dark Current Sources
Dark current sources most relevant to CCD are small
Dark current sources in a diode [1]
Depletion
Front surface
Diffusion
Back surface
Peripheral
Sources can be isolated by varying relevant dimensions
Dominant in a CCD
Contact Metal
n-Ge
GeO2/Al2O3
p-Ge
3
4
5
1
2
Diodes yield insights into CCD dark current
[1] D.K. Schroder, Sol.-State Electron. 27 (247) 1984.

## Slide 11: Diode Dark Current Analysis
Dark current extrapolates to
1 pA/cm2 at 150K, validating material and device quality
Dark Current vs. Temperature
Diodes yield insights into CCD dark current
Dark current sources in a diode [1]
Depletion
Front surface
Diffusion
Back surface
Peripheral
Temperature dependence indicates mechanism
Dominant in a CCD
Contact Metal
n-Ge
GeO2/Al2O3
p-Ge
3
4
5
1
2
[1] D.K. Schroder, Sol.-State Electron. 27 (247) 1984.
Fitting Equation:
Jdark = AT 3exp(-Eg/kT)

## Slide 12: Estimating Dark Current in a Germanium CCD
Key variables that govern device performance are nearly equivalent to silicon
Dark current target of < 10 pA/cm2 at 150K
(40 e-/pixel/s for an 8-µm pixel)
Used measurements on building blocks to estimate dark current noise
Germanium CCDs require cooling to achieve low dark current because of high intrinsic carrier concentration

## Slide 13: Germanium CCD Fabrication Challenges
Process Challenge | Mitigation Plan | Other Alternatives
Cannot use overlapping poly process to define phases | Single-level CCD | Low-temperature fabrication processes
Charge transfer inefficiency due to large gap between phases | 248-nm lithography to define sub-250 nm gap | 193-nm or e-beam lithography
Low sensitivity in front illumination due to metal shading | Back illumination | Engineered substrate for rapid prototyping
Silicon CCD
Germanium CCD
Poly 1
Poly 2
Poly 3
Gap Between Phases
Silicon Dioxide
Gap Between Phases
Metal Gate Electrode
Gate Dielectric

## Slide 14: Recent Fabrication Accomplishments
Transitioned to standard 200-mm wafer processing tools, first light this fall
150-mm Wafer
Discrete Devices
200-mm Wafer
Discrete Devices
200-mm Wafer
Imagers
Metal Gate Detail

## Slide 15: Output Transistor Comparison Buried Channel MOSFETs
Germanium PMOSFET
W = 10 µm, L = 2 µm
Silicon NMOSFET
W = 20 µm, L = 2 µm
Both plots normalized by gate width and plotted over same range of |Vg-VT|
Slightly higher off-state current in Ge due to nonoptimized peripheral isolation
Higher current in germanium PMOSFET due to combination of higher carrier mobility and higher gate capacitance
Potential for lower read noise in Ge CCDs

## Slide 16: Germanium Imager Designs
Reticle Layout
  1k × 2k × 8.1 µm frame transfer CCD
  128 × 128 x 8 µm OTCCD
  32 × 32 × 8.1 µm full-frame CCD
  Shift registers
  Shift registers
30 µm
OTCCD Pixel Detail
Single pixel

## Slide 17: Summary
Germanium CCDs offer numerous advantages for X-ray and SWIR imaging
  Wide response
  Able to fabricate large-format CCDs on this material
Our work to date has validated the potential of germanium as a low-noise CCD material
Our first (front-illuminated) imaging devices are in fabrication
  Includes 128 × 128 OTCCD, 1 Mpixel frame transfer CCD