## Slide 1: Flattening your flats or: How I learned to stop worrying and love the CBP
Merlin Fisher-Levine - Princeton

## Slide 3: Why do we flat-field / what does it do?
Corrects for amplifier-to-amplifier gain variation & pixel-to-pixel QE
  i.e. deals the conversion between ADU and incident photons
Quick history of flat-fielding:
  Twilight flats
  Nighttime sky flats
  Conventional dome flats
What comes next?
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Wide field camera → large gradients.
And what about the spectrum?
Time variable night-to-night
Time variable night-to-night
Night sky spectrum
Alt/az telescope + stray/scattered/ghosted light
→ flat-fields depend on rotator position, >5% difference
Typically incorrect input spectrum, + lamp variability

## Slide 4: Why do we flat-field / what does it do?
Corrects for amplifier-to-amplifier gain variation & pixel-to-pixel QE
  i.e. deals the conversion between ADU and incident photons
Quick history of flat-fielding:
  Twilight flats
  Nighttime sky flats
  Conventional dome flats
What comes next?
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Wide field camera → large gradients.
And what about the spectrum?
Time variable night-to-night
Time variable night-to-night
Night sky spectrum
Alt/az telescope + stray/scattered/ghosted light
→ flat-fields depend on rotator position, >5% difference
Typically incorrect input spectrum, + lamp variability
Flat-fields with monochromatic illumination + a CBP

## Slide 5: What is a CBP
CBP stands for Collimated Beam Projector
Essentially, it’s a telescope “run backwards”
A telescope is a device which maps angles to positions
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
A CBP is a telescope which takes a mask at what would have been its focal plane, and projects light through it.
Please note that despite the ponytail
and inept use of a telescope,
that is not Robert!
(You can tell from the shoes)
So reversing this…

## Slide 6: What is a CBP
CBP stands for Collimated Beam Projector
Essentially, it’s a telescope “run backwards”
A telescope is a device which maps angles to positions
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
This results in multiple “pencil-beams” of light coming out of the CBP
The angles at which these beams exit the CBP are defined by the position of the holes in the mask
So reversing this…
(see Giles the cheese hedgehog)

## Slide 7: What is a CBP
CBP stands for Collimated Beam Projector
Essentially, it’s a telescope “run backwards”
A telescope is a device which maps angles to positions
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
The mask pattern is therefore re-imaged by the observing telescope onto its focal plane
Meaning we can place spots of our choosing at will
So reversing this…

## Slide 8: But why?
Because it mimics a (monochromatic) source at infinity
And because the diameter of these output beams == the diameter of the CBP telescope (~30cm)
  it’s finite (unlike the planar wavefront from stars)
  and therefore only samples a portion of the optics of the main telescope (primary mirror, filters etc.)
These properties mean it can be used for all sorts of fun things…
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016

## Slide 9: Like what?
A Collimated Beam Projector allows us to:
    Make ‘photometric flat-fields’
And also, as an aside:
Monitor the filter transmission profiles, including as a function of position on the filter
  Although we are assured that the LSST filter edges won’t move over time (unlike every filter ever made previously), it is probably prudent to confirm/measure this.
Provides a convenient alternative method to differentiate between between electronic crosstalk and optical ghosting, i.e. without relying on cosmics (low signal) or bleed trails (messy)
Measure absolute throughput (?!?!)
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016

## Slide 10
CBP
CBP

## Slide 11: Variable definitions
SQE = System response due to quantum efficiency
Soptical = System response from optical distortions
            (Different pixels subtend different areas on the sky)
Spixel = (static) pixel size variations: intrinsic variation, tree-rings
i = non-uniformity in the flat-field screen illumination
A = ghosting/scattered light from the flat-field screen
Other effects:
  Vignetting - flat in λ, but degenerate with QE, opposite in sign to Soptical
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Additive
Additive
Multiplicative
Multiplicative
F = ( 1 + i + A ) Ssys
Flat-field observation
Ssys = SQE Spixel Soptical
System response function

## Slide 12: Flat-field Illumination (i)
1D
2D
i = 0.05x2
i = 0.05x2 + 0.05y2
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 13: Ghosting (A)
1D
2D
A = -0.07 cos (2πx)
A = - (0.04 cos (πx) + 0.04 cos (πy))
|range| = 0.08
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 14: Optical Distortion (Soptical)
1D
2D
Soptical = 1 + 0.1x4
Soptical = 1 + 0.1x4 + 0.1y4
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 15: Quantum Efficiency (SQE)
2D
SQE = piecewise linear
SQE = piecewise bi-linear
range = 74-97%
(no chromaticity)
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Introduces discontinuities
Random gradients in x and y
Random nominal QE
Figure credit: R. Lupton

## Slide 16: Pixel Sizes (Spixel)
1D
2D
Spixel = 1 + 0.01 sin(50πx)
Spixel = 1 + 0.01 sin(νi πx + ϕi)
ν = [59, 61, 67, 71, 73, 79, 83, 89, 97]
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Origins off-chip by random
amount → random phase
Incommensurate periods
Random chip orientation
Figure credit: R. Lupton

## Slide 17: Pixel Sizes (Spixel)
1D
2D
Spixel = 1 + 0.01 sin(50πx)
Spixel = 1 + 0.01 sin(νi πx + ϕi)
ν = [59, 61, 67, 71, 73, 79, 83, 89, 97]
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Origins off-chip by random
amount → random phase
Incommensurate periods
Random chip orientation
Figure credit: R. Lupton

## Slide 18: Putting these together…
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016

## Slide 19
1D
2D
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Ssys = SQE Spixel Soptical
System response function
Figure credit: R. Lupton

## Slide 20
1D
2D
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
F = ( 1 + i + A ) Ssys
Flat-field observation
Figure credit: R. Lupton

## Slide 21
Reconstruction from CBP data (1)
1) Measure flat-field (green)
2) Take CBP samples (blue)
      NB these see Ssys
      but not additive parts i + A
3) Spline interpolate
Figure credit: R. Lupton

## Slide 22
Reconstruction from CBP data (1)
1) Measure flat-field (green)
2) Take CBP samples (blue)
      NB these see Ssys
      but not additive parts i + A
3) Spline interpolate
Figure credit: R. Lupton

## Slide 23
Reconstruction from CBP data (2)
4) Divide flat-field by
the spline
→ this gives an
estimate of the
illumination
Figure credit: R. Lupton

## Slide 24
Reconstruction from CBP data (3)
5) Smooth this with a
spline
Figure credit: R. Lupton

## Slide 25
Reconstruction from CBP data (4)
6) Divide the flat-field by the
smoothed illumination correction
Figure credit: R. Lupton

## Slide 26
Residuals
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Max = 1.00023
Figure credit: R. Lupton

## Slide 27
1D
2D
Reconstruction from CBP data (1)
Place sample points in 2D
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 28
1D
2D
Reconstruction from CBP data (1)
Place sample points in 2D
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 29
1D
2D
Reconstruction from CBP data (2)
Take data at
sample points & interpolate
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 30
1D
2D
Reconstruction from CBP data (3)
Divide flat-field by
the spline surface
and smooth
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 31
1D
2D
Reconstruction from CBP data (4)
Apply derived
illumination correction
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton

## Slide 32
1D
2D
Reconstruction from CBP data (5)
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Figure credit: R. Lupton
Arrive at
system response,
plot residuals…

## Slide 33
1D
2D
Residuals
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Max = 1.00023
Max = 1.0137
Figure credit: R. Lupton

## Slide 34
1D
2D
Residuals
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Max = 1.000048
Max = 1.000021
Figure credit: R. Lupton

## Slide 35: It’s real!
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Photo credit: M. Coughlin

## Slide 36: It works! (At CTIO/DECam)
DECam spot images
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
1st prototype
Focus was poor due
to rough transit
For a first try it worked wonderfully!

## Slide 37
Filter transmission measurement at Pan-STARRS
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
It works! (At Pan-STARRS)
Figure credit: M. Coughlin, N. Mondrik, J. Tonry, C. Stubbs
CBP v2:
New and improved!
Now with
smaller spots!
made with
CBP v2 !!
made with
CBP v2 !!
made with
CBP v2 !!

## Slide 38: Conclusion
The CBP allows us to:
  Calculate the system response function, Ssys, and correct for the ghosting and non-uniform illumination in dome-flats
  Do this as a function of wavelength by using monochromatic light for both dome-flats and CBP illumination
  Thereby accurately flat-field our images to unprecedented precision, and do so correctly for any given SED
  Sample only a portion of the pupil at a time, allowing characterisation of the filter transmission as a function of position
As a bonus, we also get help with crosstalk measurements and monitoring the evolution of the filter bandpasses over time.

## Slide 39: Backup slides

## Slide 41
Residuals vs. CBP samples
Merlin Fisher-Levine - PACCD - BNL - Dec. 1st 2016
Please ignore me, I am the product
of numerical instability…
How does the error change
with number of number of CBP
samples taken per chip?