## Slide 1: ePIC Clock (Control) Distribution Test
Clock Source
Si5344H
“DAM”
KCU105
(xcku040)
“RDO”
XEM8320
(xcau25p)
Clock jitter cleaner
Si5394A
OUT
ClkRefIn
IN
RxRecClk
GTY_Rx
GTH_Tx
IN#3
OUT
OUT#3
Tektronix
MSO64
8 GHz, 25 GSps
Si5344H
XEM8320
KCU105
Si5394A

## Slide 2: 2 SiliconLab Clock Evaluation boards setting
Si5344H:
XTAL: 48 MHz
Output0: 98.5 MHz
Output1: 197 MHz, or 98.5 MHz
Output2: 250 MHz
Output3: 100 MHz
Si5394A:
XTAL: 48 MHz
Input0: 98.5 MHz
Input1: 197 MHz
Input2: 156.25 MHz
Input3: FeedBack
Output0: 98.5 MHz
Output1: 197 MHz
Output2: 250 MHz
Output3: ZDM_98.5MHz
Clock generator
Clock jitter cleaner, Clock generator for FEB
All the scope traces are triggered by the Si5344H output#1, while the output#0 is used as (KCU105) ClkRef for the MGT
The scope’s measurement functions, Time Interval Error (TIE), Clock frequency, Clock period are used.

## Slide 3
3 Rx Recovered clock (ClkRef_out), with regular 8b/10b coding
Example scope trace:
  Six bytes K28.5 (COMMA) and Two bytes data per (98.5 MHz) clock cycle
Downlink (KCU105XEM8320, to simulate DAM  RDO) Encoding:
Eight Bytes (8b/10b) payload per Clock (98.5 MHz)
7.88 Gbps line rate
Equivalent (downlink) control: 6.304 Gbps
DownLink scope trace: Unit Interval (UI): ~ 127 ps
Connections between KCU105 and XEM8320:
SMA co-ax copper cables for 3.1, 3.2, 3.3
SFP optic transceivers and LC fiber for 3.4
5.076ns, half clock cycle

## Slide 4
3.1 Rx Recovered clock (ClkRef_out), with regular 8b/10b coding
A), All eight bytes are K28.5 (COMMA)
RxRecClk: TIE σ = 2.5 ps , phase stable @ 40 µs
Period
TIE
Frequency
Frequency
Period
TIE
B), Six bytes are K28.5 (COMMA), Two bytes data
RxRecClk: TIE σ = 2.6 ps , phase stable @ 40 µs

## Slide 5
3.2 Rx Recovered clock (ClkRef_out), with regular 8b/10b coding
C), Four bytes are K28.5 (COMMA), four bytes are data
RxRecClk: TIE σ = 2.65 ps , phase stable @ 40 µs
Period
TIE
Frequency
Frequency
Period
TIE
D), Two bytes are K28.5 (COMMA), Six bytes are data
RxRecClk: TIE σ = 2.8 ps , phase stable @ 40 µs

## Slide 6
3.3 Rx Recovered clock (ClkRef_out), with regular 8b/10b coding
E), All eight bytes are regular data
RxRecClk: σ = 2.87 ps, phase stable @ 40 µs
Period
TIE
Frequency
F), (= D), Two bytes are K28.5 (COMMA), Six bytes are data
RxRecClk: σ = 2.77 ps , phase stable @ 40 µs
Period
TIE
Frequency

## Slide 7: 3.4 Rx Recovered clock (ClkRef_out), with regular 8b/10b coding, SFP
All eight bytes are regular data
RxRecClk: TIE σ = 3.2 ps, Phase stable @ 40 µs
Period
TIE
Frequency
Through Avago AFBR-710SMZ optic transceivers and a short multimode LC fiber (850nm)

## Slide 8: 4.1 Jitter cleaning of the Rx Recovered clock
Jitter cleaned (by Si5394A) clock:
TIE: σ = 1.3 ps
But, the phase is (slowly) drifting. (Si5394A: Loop BW 88Hz)
Filtered Clock
drift
“scope traces captured several seconds later”
Period
TIE
Frequency

## Slide 9: 4.2 Jitter cleaning of the Rx Recovered clock (Loop BW at 2.8 KHz)
Jitter cleaned (by Si5394A) clock:
TIE: σ = 1.54 ps @ 40 µs
Si5394A, Loop BW 2.8 KHz
Stable phase relative to the original 98.5 MHz clock
Scope trace (lower right) at ~ 36 µs after (original 98.5 MHz clock) trigger
No drift
Period
TIE
Frequency
40 ps

## Slide 10: 5 Test Summary (so far)
The GTY recovered clock works
  The TIE (σ) is below 4 ps
  The phase is stable. The recovered clock reproduces the original (GTH_Ref) Clock.
The Clock jitter cleaner can further improve the recovered clock
  The cleaned clock TIE (σ) is below 2 ps,
  The jitter cleaner Loop Bandwidth needs be ~ KHz, not ~Hz, to match with the GTY.

## Slide 11: 6.1 To do in the near future
Implement the link coding in the new FELIX card (VERSAL chip);
Recovered clock phase stability (on power cycle, reset etc)
    CERN Tclink technique: ‘fix’ the TX Phase Interpolator, ‘fix’ the Rx equalizer;
To explore the possibility of FELIX clock input (or ePIC clock distribution from GTM):
Dedicated differential Clock (98.5 MHz)
985 Mbps (one byte 8b/10b per 98.5 MHz clock)
to simulate the Fast control from GTM
FELIX
versal
985 Mbps (8b/10b at 98.5 MHz)
To simulate the Status of the FELIX
FPGA IOB, not MGT

## Slide 12: 6.2 To do in the farther future
RDO prototype: for clock recovery design verification in the ‘real world’, long-term recovered clock stability test
Optic transceiver
RX
TX
1:2 fanout
O1
O2
FPGA
xcau25p
GTY
RX
TX
Si53x5
O1
O0
O2
IN
RxRecClk
GENIO
FEB
Phase difference detection
Fiber
FEB
optional emergency circuit I2C
RDO
N* 98.5 MHz
197 or
98.5 MHz
control
ClkRef
pseudoClock
98.5 MHz
OUT
IN3
Zero Delay FB

## Slide 13
Thank you

## Slide 14
Backup slides

## Slide 15
Encoding: 7.88 Gbps line rate (UI = ~127 ps)
80 bits per clock (98.5 MHz) period.
1111111111,1x1x11x1x1,1x1x11x1x1,1x1x111111,000000x0x0,0x0x00x0x0,0x0x00x0x0,0000000000: similar like the scope trace
1111111111,1111111111,1111111111,1111111111,0000000000,0000000000,0000000000,0000000000: ‘pseudo 98.5 MHz clock’
6.2 To do in the farther future, Why the phase Difference detection is needed
XEM8320, UltraScale+ ARTIX, GTY RxRecClk
Frequency: 197 MHz
RxRecClock: σ < 3 ps
But, the phase is jumping at n*(UI = 127 ps) on reset/power.

## Slide 16: 3 Low Pass filter for Clock recovery, Clock feature in TX
39541.64ns
39541.68ns
40 ps
40 ps
Encoding: 7.88 Gbps line rate (UI = ~127 ps)
One 8b/10b payload at Clock (98.5 MHz) high,
One 8b/10b payload at Clock low,
Equivalent control: 1576 Mbps. (can be: 788 Mbps global, 788 Mbps sub-system)
The ‘Clock’ edge P_P < 20 ps
Zoom In on the edges
Downlink, 80 bits per cycle
98.5 MHz clock

## Slide 17: 3.1 Low Pass filter for Clock recovery, Clock feature in TX, BW = 500 MHz
LowPass Filter:
500 MHz,
TIE measurement at 40 µs
(~ 8 km fiber delay)
GTH Ref_Clock: σ < 1 ps
Rising Edge: σ = 18 ps
Falling Edge: σ = 60 ps
Falling Edges, TIE
Rising edges, TIE
Ref_Clock, TIE

## Slide 18: 3.2 Low Pass filter for Clock recovery, Clock feature in TX, BW = 1 GHz
LowPass Filter:
1 GHz,
TIE measurement at 40 µs
(~ 8 km fiber delay)
GTH Ref_Clock: σ < 1 ps
Rising Edge: σ = 11 ps
Falling Edge: σ = 21 ps
Falling Edges, TIE
Rising edges, TIE
Ref_Clock, TIE
Need jitter cleaner

## Slide 19: 3.3 Low Pass filter for Clock recovery, Clock feature in TX, BW = 2 GHz
LowPass Filter:
2 GHz,
TIE measurement at 40 µs
(~ 8 km fiber delay)
GTH Ref_Clock: σ < 1 ps
Rising Edge: σ = 4 ps
Falling Edge: σ = 13 ps
Falling Edges, TIE
Rising edges, TIE
Ref_Clock, TIE
The measurement is on the MATH, Ch#2 – 0.8*Ch#3
Ch#2: TX+, BW 2 GHz
Ch#3: TX-, BW 500 MHz

## Slide 20: 3.4 Low Pass filter for Clock recovery, Clock feature in TX, BW = 3 GHz
LowPass Filter:
3 GHz,
TIE measurement at 40 µs
(~ 8 km fiber delay)
GTH Ref_Clock: σ < 1 ps
Rising Edge: σ = 3 ps
Falling Edge: σ = 6 ps
Falling Edges, TIE
Rising edges, TIE
Ref_Clock, Period
The measurement is on the MATH, Ch#2 – 0.6*Ch#3
Ch#2: TX+, BW 3 GHz
Ch#3: TX-, BW 500 MHz
Hard to implement
Clock
Ch#3
Ch#2
MATH

## Slide 21: 3.5 Rx Recovered clock (ClkRef_out), with Clock feature in TX
XEM8320, UltraScale+ ARTIX, GTY RxRecClk
Frequency: 197 MHz
RxRecClock: σ = 2.0 ps @ TIE measurement at 400ns range (80 meter fiber equivalent);
RxRecClock: σ = 3.0 ps @ TIE measurement at 40 µs (20 µs before and after trigger) range;