## Slide 1: High-Field Combined Function Permanent Magnet
December 20th, 2023
Stephen Brooks

## Slide 2: Magnet Design
Bought 24 permanent magnet wedges from AllStar Magnets
  Material grade N42EH, Br=1.30T
  45mm length
B(0) = -0.9512T, 55.54T/m
±10.5mm good field region
  Bmax,gfr = -1.536T
±7.5mm vertical aperture
6mm total minimum gap
Open midplane design with >1.5T in good field region, 10-3 field accuracy, relevant for CEBAF upgrade
cm grid
10cm

## Slide 3: Assembled Magnet
Outer ½” thick aluminium frame for strength

## Slide 4: Assembled Magnet
Outer ½” thick aluminium frame for strength

## Slide 5: Endcaps for Safety (no shims here)

## Slide 6: High fields exist in aperture

## Slide 7: Fieldmap Measurement
Used Senis 3MH6 teslameter with 3-axis Hall probe
  Accurate to ±0.01%
Two (horizontal X,Z) linear movement stages
  Few microns
±10.5mm × 200mm
  0.5mm, 2.5mm steps
8 repeated scans

## Slide 8: Fieldmap Measurement

## Slide 9: Graphical Fieldmap Plot
Left: linear colour scale
Right: log colour map,
field inversion (flux return) far from magnet is visible
Z
X

## Slide 10: Magnet Field Mapping Run Sequence
Run 1: bare magnet but bad wedge alignment
Run 2: added brass shims between wedges (better alignment)
Run 3: added 26× 35mil Ø iron tuning rods but got X-axis direction definition wrong!
Run 4: tried first tuning iteration again with corrected X axis
Run 5: verify first tuning iteration with different rod holder but field changed
Run 6: investigate bare magnet without tuning rods again, field had changed, wedge alignment worse perhaps
Run 7: tuning rods first iteration based on run 6
Run 8: 2nd tuning iteration from run 6
Run 9: 3rd tuning iteration from run 6
Three consecutive field tuning iterations

## Slide 11: Aperture after Brass Shims (Run 2)

## Slide 12: Magnet with Iron Tuning Rods

## Slide 13: Magnet with Iron Tuning Rods

## Slide 14: Maximum Field Error History
Original
Brass shims
Iter 1, rev. X axis
Iter 1, correct X axis
Iter 1, verify, block slip?
No tuning rods
Iter 1
Iter 2
Iter 3

## Slide 15: Integrated Fields (Run 9)
Chart: 
| 1.2E-2 | 1.15E-2 | 1.0999999999999999E-2 | 1.0500000000000001E-2 | 0.01 | 9.4999999999999998E-3 | 8.9999999999999993E-3 | 8.5000000000000006E-3 | 8.0000000000000002E-3 | 7.4999999999999997E-3 | 7.0000000000000001E-3 | 6.4999999999999997E-3 | 6.0000000000000001E-3 | 5.4999999999999997E-3 | 5.0000000000000001E-3 | 4.4999999999999997E-3 | 4.0000000000000001E-3 | 3.5000000000000001E-3 | 3.0000000000000001E-3 | 2.5000000000000001E-3 | 2E-3 | 1.5E-3 | 1E-3 | 5.0000000000000001E-4 | 1.04083409E-17 | -5.0000000000000001E-4 | -1E-3 | -1.5E-3 | -2E-3 | -2.5000000000000001E-3 | -3.0000000000000001E-3 | -3.5000000000000001E-3 | -4.0000000000000001E-3 | -4.4999999999999997E-3 | -5.0000000000000001E-3 | -5.4999999999999997E-3 | -6.0000000000000001E-3 | -6.4999999999999997E-3 | -7.0000000000000001E-3 | -7.4999999999999997E-3 | -8.0000000000000002E-3 | -8.5000000000000006E-3 | -8.9999999999999993E-3 | -9.4999999999999998E-3 | -0.01 | -1.0500000000000001E-2 | -1.0999999999999999E-2 | -1.15E-2 | -1.2E-2
avg iBx/L (T) | -0.013374 | -0.012294 | -0.011372 | -0.010372 | -0.0094553 | -0.0085976 | -0.0077619 | -0.0068806 | -0.0060013 | -0.0052113 | -0.0044557 | -0.003637 | -0.002821 | -0.0021183 | -0.0013481 | -0.00058805 | 0.00020409 | 0.00093702 | 0.0017421 | 0.0024916 | 0.0032711 | 0.0039927 | 0.0047231 | 0.0054922 | 0.0062953 | 0.0070439 | 0.0077767 | 0.0085771 | 0.0094515 | 0.0103 | 0.011166 | 0.012052 | 0.012933 | 0.013793 | 0.014618 | 0.015505 | 0.016394 | 0.017164 | 0.017949 | 0.018682 | 0.019292 | 0.019754 | 0.020134 | 0.020555 | 0.020874 | 0.021168 | 0.021561 | 0.022101 | 0.022766
avg iBy/L (T) | -0.30835 | -0.33717 | -0.36563 | -0.39365 | -0.42149 | -0.44933 | -0.4772 | -0.50495 | -0.53263 | -0.56042 | -0.58823 | -0.61602 | -0.64372 | -0.67155 | -0.69942 | -0.72721 | -0.75505 | -0.78285 | -0.81074 | -0.83849 | -0.86625 | -0.89408 | -0.92192 | -0.94966 | -0.97743 | -1.0053 | -1.0332 | -1.061 | -1.0887 | -1.1166 | -1.1444 | -1.1722 | -1.1999 | -1.2277 | -1.2553 | -1.283 | -1.3105 | -1.3382 | -1.3659 | -1.3935 | -1.4213 | -1.4493 | -1.4775 | -1.5057 | -1.534 | -1.5621 | -1.59 | -1.617 | -1.6433
avg iBz/L (T) | 0.0058618 | 0.0064108 | 0.0069546 | 0.0074885 | 0.0080165 | 0.0085483 | 0.0090775 | 0.0096132 | 0.01014 | 0.010666 | 0.011208 | 0.011743 | 0.012265 | 0.012801 | 0.013326 | 0.013866 | 0.014391 | 0.014931 | 0.01548 | 0.016003 | 0.016545 | 0.017093 | 0.017627 | 0.018142 | 0.018676 | 0.019207 | 0.019744 | 0.020295 | 0.020821 | 0.021366 | 0.021887 | 0.02243 | 0.022974 | 0.023521 | 0.024041 | 0.024584 | 0.025115 | 0.025679 | 0.026233 | 0.026796 | 0.027409 | 0.027997 | 0.028643 | 0.029265 | 0.029902 | 0.030489 | 0.031108 | 0.031606 | 0.031977
Bx error (T) | -0.0017571 | -0.0014279 | -0.0012573 | -0.0010086 | -0.00084289 | -0.00073638 | -0.00065198 | -0.00052189 | -0.00039377 | -0.00035503 | -0.00035065 | -0.00028318 | -0.00021846 | -0.00026698 | -0.00024798 | -0.00023919 | -0.00019828 | -0.00021658 | -0.00016274 | -0.0001645 | -0.00013615 | -0.0001658 | -0.0001867 | -0.00016881 | -0.00011694 | -0.00011954 | -0.00013795 | -8.882e-05 | 3.4318e-05 | 0.00013152 | 0.00024652 | 0.00038087 | 0.00051083 | 0.00061948 | 0.00069303 | 0.00082923 | 0.00096719 | 0.00098539 | 0.0010196 | 0.0010009 | 0.00086028 | 0.00057072 | 0.00019974 | -0.00013067 | -0.00056298 | -0.0010196 | -0.001378 | -0.0015891 | -0.0016753
By error (T) | 0.0011543 | 0.00015984 | -0.00048753 | -0.00068894 | -0.00071155 | -0.00073974 | -0.00079524 | -0.00072143 | -0.00059132 | -0.0005639 | -0.0005518 | -0.0005307 | -0.00040906 | -0.00042445 | -0.00047562 | -0.00045508 | -0.00047801 | -0.00045468 | -0.00053296 | -0.00046246 | -0.00040502 | -0.00042144 | -0.00044412 | -0.00037452 | -0.00032457 | -0.00037849 | -0.0004542 | -0.0004358 | -0.00035801 | -0.00038995 | -0.00041271 | -0.00039065 | -0.00023339 | -0.00019404 | -6.3521e-05 | 0.00013545 | 0.00039576 | 0.00049923 | 0.00068097 | 0.00084064 | 0.00082549 | 0.00066666 | 0.00027754 | -6.4686e-05 | -0.00052799 | -0.00084064 | -0.00089554 | -0.00015688 | 0.0013563
Good field region ±10.5mm
RHS
RHS
Main Field (T)
Field Error (T)
X (m)
Note: integrated fields are divided by magnet nominal length of 45mm to give Telsa

## Slide 16: Integrated Field Error History
Gradual improvement over iterations 0-3 (runs 6-9)
Short magnets often require more iterations because of end field effects
By Field Error (T)
Bx Field Error (T)
X (m)
X (m)
8mT
2.5mT

## Slide 17: Conclusion and Future Work
A short section of magnet similar to CEBAF upgrade FFA has been tuned to <10-3 accuracy
  No glue/epoxy was used (might reuse wedges), so some issues with blocks staying in place
  A technique with epoxy was also attempted (in backup slides)
Temperature variation and material thermal coefficient appears to be the main source of measurement error now
  Thermal coefficient is approximately 1.1×10-3/K
  Temperature-controlled enclosure is a future option

## Slide 18: Tooling for Assembling a Permanent Magnet

## Slide 19: Tooling Design
Full mould
all pieces and central plug, printed double wall thickness for strength
“Spider”
Channels for half the pieces
Can be flipped over to get other half
Corner holes for ¼”-20 threaded rods throughout, provide loose alignment, constraint

## Slide 20: Can 3D print all parts in place
0.3mm spacing between pieces minimum possible with our printers

## Slide 21: Point field measurement tool for Hall probe, also a “spider”
Single-axis Hall probe inserted in X and Y directions in cross-shaped holes, depth is controlled the centre of magnet length
Used this to confirm field of all 24 wedges at 7 locations  rough magnetisation vector measurement

## Slide 22: Assembly Sandwich

## Slide 23: Assembly Process
1.
2.
3.
4.
5.

## Slide 24: Assembly Process Notes
Insert one half of the blocks and then flip the spiders over and insert the other half
Make sure each block is pushed all the way in because they can repel longitudinally
  Actually a mixture of attraction and repulsion
The later blocks will be tighter fit but it was OK with the 0.3mm gaps (0.15mm larger mould)
  Small gaps between some magnets were visible after central plug removed (0.1-0.2mm), inserted brass shims

## Slide 25: Permanent Magnet Assembly With Epoxy

## Slide 26: Double Tape-Wrapped Spacers
Spacer must not be glued in place!

## Slide 27: Tape-Wrapped Base
Spacer must not be glued in place!
Epoxy should not leak out of base

## Slide 28: Side Tape and Epoxy “Bucket”
RTV sealant used around seams, plus extended spacers

## Slide 29: Epoxy Pour in Vacuum Chamber

## Slide 30: Solid Epoxy after Spacers Removed