## Slide 1: MPGD readout electronics
I. Mandjavidze
Irfu, CEA Saclay
Incremental Design and Safety Review of the EIC Tracking Detectors March 20-21, 2024

## Slide 2: Outlook
Requirements
Readout architecture
  Salsa ASIC
  FEB
  System scale
Organization
  Production strategy
  Planning
  QA and risk mitigation
Summary
References
  Signal - https://indico.bnl.gov/event/20965/contributions/82420/attachments/50649/86604/231026_IM_CyMBal_ExpectedSignal.pdf
  Data collection - https://indico.bnl.gov/event/18118/contributions/72179/attachments/45781/77366/221221_MpgdDataCol_IM.pdf
  Calibration - https://indico.bnl.gov/event/16040/contributions/64090/attachments/41290/69185/220520_MpgdTrack_CalibRates_IM.pdf
  Salsa - https://indico.bnl.gov/event/22053/contributions/86152/attachments/52272/89395/SALSA_EPIC_electronics_DAQ_20240125.pdf
  MPGD LV - https://indico.bnl.gov/event/22316/contributions/87363/attachments/52727/90159/240215_IM_MpgdPower.pdf
  FEB options - https://indico.bnl.gov/event/21104/contributions/83856/attachments/51197/87574/231127_IM_Mpgd_VtrxPlus.pdf
  Detailed study of FEB organization options and their powering schemes – provided in support material
  Run control state machine and proposal of synchronous commands – provided in support material
  Salsa and Prisme PLL IP – proposals, provided in support material

## Slide 3: Charge questions addressed
Are the technical performance requirements appropriately defined and complete for this stage of the project?
Are the plans for achieving detector performance and construction sufficiently developed and documented for the present phase of the project?
Are the current designs and plans for detector, electronics readout, and services sufficiently developed to achieve the performance requirements?
Are plans in place to mitigate risk of cost increases, schedule delays, and technical problems?
Are the fabrication and assembly plans for the various tracking detector systems consistent with the overall project and detector schedule?
Are the plans for detector integration in the EIC detector appropriately developed for the present phase of the project?
Have ES&H and QA considerations been adequately incorporated into the designs at their present stage?

## Slide 4: Reminder on MPGD sub-systems and channel counts
Cylindrical Micromegas Barrel Layer : CyMBaL : ~30k channels
  32 tiles of 1024 channels each
µRWELL Barrel Outer Tracker : µRWell-BOT : ~100k channels
  24 modules of 4 096 U-V strips each
µRWell End Cap Tracker : µRWell-ECT : ~30k channels
  8 half-disks of 4 000 X-Y strips each
~160k-channel heterogeneous system
  Micromegas, µRWell, barrel, endcap, curved, planar, circular
Common approach to acquire data from different types of ePIC MPGDs
  Use same frontend ASIC
  Share frontend design between groups
    Adapt form factor if needed
45 x 60
170 x 35
R 50

## Slide 5: Requirements
Typical signal 1-1.5 keV resulting to 15-30 fC per channel
  Cluster sizes, detector gains, charge collection
Aimed dynamic range of 10-bits
  Signal / noise of ~60
    S/Th =~10 and Th/N = ~6
  Max / signal of ~10
Timing precision of few ns
  Low contribution to the aimed overall time measurement accuracy of ~O(10 ns)
Channel occupancies of ~10 kHz
  Including factor of safety margin
Streaming readout
  With support of in-situ calibration and of on-demand readout
~1.8 T magnetic field
Mild radiation environment
  TID and neutron fluence after 10 years: 10 krad and 1011 neq / cm2
  20 MeV proton flux: 100 particle / cm2 / s
Stringent space for detector readout and services
Chart: Energy deposit in CyMBaL tracker
R(cm)
Z(cm)
B (T)
FEB
FEB
Magnetic field map
CyMBaL
Charge 1

## Slide 6: Readout organization of ePIC detector
FEB – frontend board with readout ASICs
  Sub-detector specific
RDO – readout module – first stage of FEB data aggregation, last stage to dispatch clock & control
  Mostly common design framework between sub-detectors, different form factor
DAM – data aggregation module – interface with computing and global timing and control unit (GTU)
  Common design for all sub-detectors
Downstream towards detector : clock, control, monitoring
Upstream towards storage : physics, calibration, monitoring data
FPGA
E/O
RDO
Electrical or optical
FPGA
O/E
DAM
FEB
FE ASIC
Optical fibers
Patch panel
No restriction on length
On-detector
Within detector or cavern
Friendly area
Distance to be determined
Place to be identified
On-line computing Storage
Clock
Control
GTU
Charges 2, 3
This talk

## Slide 7: Readout strategy: signal shape sampling
Signal is continuously sampled with an ADC
Signal samples above threshold are retained
Nominal (physics data) readout : signal amplitude and timing is derived
  Time of max (as on example) or time of arrival (fitting samples on rising edge)
On-demand readout : signal shapes or raw non-ZS data are provided
  Calibration, detector studies
Guarantees best noise immunity and thus best S/N ratio
  Allows on-line common mode noise (CMN) subtraction before ZS
ADC clock
Signal
ZS threshold
TMax
Charges 2, 4

## Slide 8: Salsa : a 64-channel versatile MPGD readout ASIC
Currently under development by a collaboration of Irfu, CEA Saclay and University of Sao-Paolo
  Support from Project R&D program for Salsa prototyping including clock management IP Prisme
CSA
Analog Shaper
ADC
CMN
Trigger
seeds
Channel 1
Channel 2
Channel 64
DSP
FIFO
Ampli
Timing
Pedestal Baseline
Digi Shaper
ZS
B
U
F
F
E
R
High speed serial IO
FIFO
Tx
Tx
Tx
Clock management
Prisme IP
Bandgap
Bias
Debug
Clock, trigger, synch commands
I2C
Slow control, monitoring & registry
Analog inputs
Test
1 Gb/s serial data links
Envisaged
Baseline
Trigger prim
Data path
Trigger path
1 Gb/s Trigger prim
Salsa
Trigger
seeds
FIFO
Rx
CDR
M U X
MUX
1 Gb/s serial link
Charges 2, 3
ADC

## Slide 9: 64-channel Salsa : brief recap of specifications
Channel features ePIC MPGD needs
  4 dynamic ranges : 50 fC, 250 fC, 500 fC, 5 pC 250 fC
  10 peaking times : from 50 ns to 500 ns 100-200 ns
  Support for high input capacitances up to 1 nF and beyond ~200 pF
  Both signal polarities Negative
  Rate per channel : up to 100 kHz <10 kHz
  Sampling rate : programmable, up to at least 50 MSPS 50 MSPS
  12-bit ADC with >10-bit ENOB >9.6 bits
Digital stage programmable features
  Pedestal equalization, common mode noise subtraction, zero suppression All
  Baseline tracking
  Signal amplitude and timing extraction Main readout option
Clock management with Prisme IP :
  Wide range jitter cleaner PLL, 4 clock frequency synthesizer, phase adjustment 100 MHz
Streaming and triggered readout
  Four 1 Gbit/s serial links Single Gbit/s link
  Non-ZS, signal shape or time-amplitude readout All
Backend
  Traditional interface with separated clock, sync command and control ports
  Innovative unified interface over 1 Gbit/s input link Baseline
Implementation
  65 nm TSMC
  10-15 mW/ch @ 1.2V
  Radiation hardened : SEU, > 300 Mrad, > 1013 neq / cm2 10 krad, 1011 neq / cm2
Charges 2, 3

## Slide 10: 256-channel FEB with optical interface: aimed baseline
FEB
FEB
  ASICs directly connected to 4-lane bidirectional parallel optic FireFly transceivers from Samtec
    Single 1 Gbit/s Rx line encoding clock, sync run-control and asynchronous slow control and monitoring commands
    Single 1 Gbit/s Tx line for physics, calibration, control and monitoring data
  Low active component count
    Easier to adapt to challenging on-detector environment
    Samtec FireFly : reported to stand TID of 50-100 krad and neutron fluence of at least 5x1011 neq / cm2
RDO
  Based on common ePIC design with minimal adaptation for MPGDs
  Can be placed anywhere in experimental hall with no particular environmental restrictions
Optimal tradeoff between complexity of the on-detector electronics and its power consumption
FPGA
E/O
RDO
ASIC 4
ASIC 3
ASIC 2
To / from DAM
On-detector
ASIC 1
Embedded clock, synch, control
O F p i t r i e c F
a l
l y
O F p i t r i e c F
a l
l y
4-Tx and 4-Rx optical FireFly
Optical fibers with no restriction on length
Power distribution
Low restriction area
1.2V
3.3V
Vin=12V
Data
Charges 2, 3
Adapt

## Slide 11: Example of CyMBaL configuration under study
32K channels
128 256-channel FEBs
  4 Salsa ASICs per FEB
32 1024-channel RDOs
  4 FEBs per RDO
256-channel FEBs
32-channel µcoaxial cables
Charges 2, 3
16 Salsa-s, 4 FEBs, 1 RD0 = 1024 channels

## Slide 12: FEBs, Salsa-s and RDOs for MPGDs
Estimates of operational quantities for MPGDs assuming 256-channel FEBs
Production quantities
    Including prototyping, test-bench and quality assurance needs
  750 FEBs
  4 000 Salsa-s
  200 RDOs
Work in progress to understand impact of mechanical constraints on FEB design
  Possibility to use the same form factor and ASIC integration level for all MPGDs
 |  | CyMBaL |  |  | µRWell-BOT |  |  | µRWell-ECT |  |  | Total
 |  | Tile | Sub-detector |  | Module | Sub-detector |  | ½ disk | Sub-detector |  | MPGDs
FEB |  | 4 | 128 |  | 16 | 384 |  | 16 | 128 |  | 640
Salsa |  | 16 | 512 |  | 64 | 1 536 |  | 64 | 512 |  | 2 536
RDO |  | 1 | 32 |  | 4 | 96 |  | 4 | 32 |  | 160
Charge 5

## Slide 13: Feasibility studies
Salsa and clock management Prisme IP prototypes developed
  Analog channel and clock synthesizer validated
Several alternative FEB options have been evaluated
  Passive electrical RDO interface, VTRX+ optical RDO interface, merged FEB/RDO
Data collection and calibration protocols elaborated
  Physics and calibration data throughput estimated
    On-line periodic calibration is possible
  FEB – RDO link occupancy is ~30 % : comfortable safety margin
LV power requirements estimated and power distribution schemes compared
  e.g. Baseline FEB with FireFly optical interface requires : ~9W or 35 mW / channel
Detector – FEB connectivity studies on-going
  Compact and low-cost micro-coaxial cable assemblies from KEL under evaluation
R&D on Samtec FireFly optical transceivers
  In close contact with Samtec optical division
    Bilateral NDA between Samtec and Irfu, CEA Saclay renewed
Charges 2, 3, 4
More info in links on Outlook page and in backup slides
USLS00-40-A
  USLS20-40-x-C

## Slide 14: Design and production strategy
Salsa will be produced and tested in quantities that will cover the needs of all ePIC MPGDs
The FEB design will be shared with all MPGD groups
  If needed, FEB design adaptation for a particular geometry will be under sub-detector responsibility
Common effort to adapt RDO-s and DAM-s to MPGD readout
  Firmware and Software
  And partially RDO hardware
Common effort on services
  LV and HV distribution, interlock, slow control and cooling
    Common concern of several sub-detector systems in magnetic field
FEB and RDO production and qualification tests will be shared between several sites
  Faster completion track
    Especially in case of different FEB form factors
  Lower risk associated with procurement and task execution delays
Charges 4, 5

## Slide 15: Planning
Tight schedule mainly driven by Salsa development phases
  With ~6-9-month contingency
Q2
Q3
Q4
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
2024
2025
2026
2027
2028
Salsa
FEB
Prod review
Prod review
RDO
Valid. w/ DAM
FEB interface / adaptation hw / fw / sw
Prod.
Services
Det interface / comp / form factor
Proto w/ Salsa 2
Pre-series w/ Salsa
Tender
Prod.
Salsa 1: 4-ch.
Salsa 2: 32-ch., final package
Salsa
Tender
Production
System
Det/FEB proto/RDO/DAM
Shipment
 Ready for deployment
Keep-in zones, HV-LV dist., interlocks, cooling
Validation
Prod.
Present in each MPGD group
Charges 4, 5

## Slide 16: Quality assurance
Expertise at Saclay in large scale production of ASICs and frontend boards
  In-house : automated ASIC tester robots and FE production test benches
  In industry : providing turn-key test benches
    Ex: 40 000 Rafael and 80 000 Catia ASICs for CMS Ph2 upgrade
    Ex: 700 Alice Solar, 170 LTDB Atlas, 150 Clas12 FEU boards
Rich set of equipment for Salsa and FEB test benches
  High-end LeCroy and Tectronix oscilloscopes
  High performance phase noise analyzer
  Low jitter precision clock sources
  Climate chamber
  Bonding machines
Expertise in system-level design, production, commissioning, maintenance
  Detector – readout electronics – acquisition software – analysis
    Clas12 MVT, T2K TPCs, Asacusa tracker, …
  Respecting ES&H regulations of host labs (BNL, CERN, JLab, J-PARC)
Expertise in radiation hard ASIC design and validation techniques
  Access to CERN, French and European radiation source facilities
Access to high magnetic field facilities
  At CERN and at Saclay
Charge 7

## Slide 17: Identified risks and their mitigation
Unexpected Salsa behavior or performances issues
  Diagnostic-correction-production-validation cycle may require 1 year
Delays in mass-production of FEBs and RDOs
  At least two production and validation sites
Component procurement delays
  Usual practice of anticipated acquisition of long lead time components of choice
  Establishing privileged partnership with manufacturers
    e.g. Irfu CEA Saclay has renewed bilateral NDA with Samtec optical division
Radiation impact on COTS components
  Use of components from the list of components proven for radiation tolerance
  Validation of components in radiation facilities for ePIC-like ionization doses and particle fluxes
  SEU : implementing error detection and recovery mechanisms adopted by ePIC collaboration
Noise impact on data volume
  On-line coherent noise subtraction
  Enough bandwidth to readout signal shapes and apply more sophisticated discrimination algorithms on-line
Charge 4

## Slide 18: Summary
Envisaged solutions for the ePIC MPGD readout are viable and cover performance requirements
Engaged groups have necessary experience for large scale system design, production and commissioning
R&Ds are on-going
  Detector – FEB connectivity
  Integrated ASIC interface
  COTS components validation for ePIC environment
Design of Salsa ASIC drives current planning
Closely following and contributing to collaboration-wide efforts
  Precise identification of keep-in-out zones between the sub-detectors
  Efficient low form factor magnetic-field tolerant powering means
  Defining of run-control state machine and set of rules that detector partitions need to obey
  Error detection and recovery strategy
Pursuing discussions within the MPGD DSC to identify second site for FEB design and production
  In addition to Irfu, CEA Saclay
Planning is compatible with the overall ePIC installation schedule

## Slide 19: Backup

## Slide 20: Readout strategy: signal shape sampling
ADC clock is derived from bunch crossing clock and is its (sub)multiple (e.g. ~50 MHz)
  Known frequency and phase relationship exists between the two clocks
Signal above threshold is tagged by a timestamp relative to revolution tick
  Max (as on example) or time of arrival (fitting samples on rising edge)
  Timing association with data from other sources possible
BX counter
O+4
O+5
O+3
O+2
O+1
Origin
BX clock
R
R+1
R-1
ADC clock
Signal
ZS threshold
O+6
O+7
O+8
O+9
O+10
O+11
O+12
O+13
O+14
TMax
R+2
Rev counter
Phase

## Slide 21: Backup option: 256-channel FEB with electrical interface
FEB
FEB
  ASICs directly connected to 4-lane bidirectional parallel optic FireFly transceivers from Samtec
    Single 1 Gbit/s Rx line encoding clock, sync run-control and asynchronous slow control and monitoring commands
    Single 1 Gbit/s Tx line for physics, calibration, control and monitoring data
  Low active component count
    Easier to adapt to challenging on-detector environment
    Samtec electrical FireFly : drop-in compatibility with optical counterpart
RDO
  Based on common ePIC design with minimal adaptation for MPGDs
  Possibly can be placed in a low restriction area
Attention must be payed for ground loops and noise pickup over long distance
FPGA
E/O
RDO
ASIC 4
ASIC 3
ASIC 2
To / from DAM
On-detector
ASIC 1
Embedded clock, synch, control
P F a i s r s e i F
v l
e y
P F a i s r s e i F
v l
e y
8-lane passive electrical FireFly
Shielded twinax copper cables
Power distribution
Low restriction area
1.2V
2.5V
Vin=12V
Data
Buffer
driver
Buffer
driver
Adaptation

## Slide 22: Frontend data volume estimates
Physics: support two zero suppression modes
  Nominal : peak finding readout
    12-bit amplitude, 12-bit time of max, 8-bit ToT
  On-demand : full signal shape readout
    All samples (12-bit) above threshold (typically 15-25)
On-line calibration : on demand readout
  Programmable number of non-ZS samples
  Signal shape for calibration pulse
FEB – RDO link occupancy : ~30 %
  Comfortable safety margin, even for on-demand signal shape readout
Overall physics frontend data of MPGDs
  Accurate estimates still to be done
    NB: in nominal operation, the bulk of the physics frontend data will be further processed before being archived
      e.g. clustering
Channel rate kHz |  | Peak finding Mbit/s |  | Signal shape Mbit/s
10 (5 x safety) |  | 40 |  | 265
Estimated calibration data bandwidth per ASIC ~6 Mbit/s
Estimated CyMBaL physics data bandwidth per Salsa ASIC
 |  | CyMBaL |  | µRWell-BOT |  | µRWell-ECT |  | Total
Data (Gbit/s) |  | 35 |  | 100 |  | 35 |  | 170

## Slide 23: Excerpt from data collection protocol proposal
Combine FEB (ASIC) data belonging to the same revolution
  If needed, heartbeat (RevTick) acknowledgement is sent after the last data belonging to the same revolution
  Revolution-level granularity might be handy for fragment building in RDO and in DAM
    A compromise between memory requirement and aggregation
    If too long, data fragments can be sent in a succession of packets
DAM performs revolution record building based on revolution numbers embedded in data
  Heartbeat packet can be used as indication that no more data is expected from RDO or FEB for this revolution
In other words, consider the RevTick as a 78.195 kHz constant rate trigger
  Do classical event building in RDO and DAM with a readout window of ~12.7886 µs
DH
Payload: R
DAM
RDO
n
FEB k
FE ASIC
RDO
m
FEB
j
FE ASIC
DH
Payload: R
HB
HB
R
R+1
R-1
Revolution R-2
R
R+1
Rev counter: R-1
Broadcast “RevTick” sync commands

## Slide 24: Calibration example : single “Calib” request – window readout
A “Calib” request results in a window readout
  N consecutive samples
  Full readout of non-ZS data
    All channels of all ASICs in concerned FEBs
  Window size programmable
“Calib” command distributed in barrel shifter mode
  Only to a predefined FEBs to avoid data congestion
A special “TestPulse” command to fire on FEB pulser
  Read ASIC response to know charge
Sk+11
Sk+12
Sk+13
Sk+7
Sk+8
Sk+9
Sk+10
Sk+4
Sk+5
Sk+6
Sk
Sk+1
Sk+2
Sk+3
Sk+14
Sk+15
Sk+16
“Calib”
Window of N consecutive samples
12 in this example
Pedestal variation
R
R+1
Broadcast “RevTick” sync commands
Multicast “Calib” sync command
Rev counter: R-1
Gap
DAM
RDO
n
FEB k
FE ASIC
RDO
m
FEB
j
FE ASIC
HB
HB
DH
Physics: R
DH
Physics: R
CH
Calibration: R
CH
Calibration: R

## Slide 25: LV power distribution for MPGD frontends
Powering fronted electronics in ~2 T magnetic field is a challenge task
  Efficient scheme requires on-board magnetic field tolerant DC/DC converters
    Coupled with ultra low drop-out (LDO) linear regulators
    Avoids significant power dissipation in LV cables
256-channel FEB with a bidirectional 4-lane FireFly optical interface : 9 W or 35 mW / channel
  LV wire cross-section : 1 mm2
    NB : In absence of DC-DC converters, LDO-based power distribution will require 5.5 mm2 wires
Estimates of MPGD sub-detector power needs
Engaging common intra- and extra-MPGD efforts to devise power distribution and cooling infrastructure
  Profit from expertise available within the Project in small form factor magnetic field tolerant DC-DC converters
1.5V
Remote power supply
12V
<12.5V
80% DC-DC
Ultra LDO
ASIC
1.2V
20 m
 |  | CyMBaL |  |  | µRWell-BOT |  |  | µRWell-ECT |  |  | Total
 |  | Tile | Sub-detector |  | Module | Sub-detector |  | ½ disk | Sub-detector |  | MPGDs
Power(W) |  | 36 | 1.2k |  | 144 | 3.5k |  | 144 | 1.2k |  | 6k
FEB (illustration only)

## Slide 26: Salsa0Analog: very frontend prototype
Prototype ASIC with 4 channels
  4 dynamic ranges : 50 fC, 250 fC, 500 fC, 5 pC
  10 peaking times : from 50 ns to 500 ns
  Support for high input capacitances
  Both signal polarities
Power board
Test bench
SalsaAnalog
500ns
Programmable Peaking time
50 ns to 500 ns
Typical 120 pF input capacitor
1200 e- at 500 ns
2000 e- at 50 ns
250 fC range
Equivalent Noise Charge in 250 fC range

## Slide 27: Prisme: clock management jitter cleaner PLL IP
Prisme prototype to test a new 65 nm hybrid PLL IP for Salsa and high fidelity fan-out ASIC
  Large input frequency range
  4 clock outputs with programmable frequencies and phase
Test boards arrived end of February
  Test bench completed clock generator, high-end 80 GS/S scope and phase noise analyzer
Preliminary results
  Power consomption nominal
  I2C working to read and write registers
  Low voltage differential (CLPS) receivers and transmitters operational
  Digital branch of PLL operational
    Allows to find the right frequency range for the analog branch
  Internal oscillator reaches nominal 3.2 GHz
  Locks occur for 40 MHz and 100 MHz input
  All for outputs tested with programmable frequencies
  Ongoing tests on jitter characterization and optimization
Power board
Prisme test bench
Prisme

## Slide 28: Planning
Salsa
  2024 : Salsa1
    Several fully instrumented channels including analog very fronted and 12-bit MSPS ADC
    Clock management circuitry based on Prisme IP
  2025 : Salsa2
    A 32-channel prototype with main DSP functionalities and serial link interfaces targeting final packaging
    Prototyping of unified backend interface based on a serial link
  2026 : Salsa pre-series
    A 64-channel fully functional prototype
  2027 : Mass production
FEB
  2024-2025 : small prototype developments
    Assessment of COTS components – e.g. power and monitoring solutions for radiation and magnetic field environment
    R&D on detector – FEB interface connectors
    FEB form factor and Salsa integration studies for different MPGDs
    Validation of unified interface with ASICs
  2026 : advanced prototype development
    Based on Salsa2
  2027 : Pre-series production
    Based on pre-series Salsa
  2028 : Mass production

## Slide 29: Planning
RDO
  2024 – 2025 : FEB – RDO interface specification and validation, plus adaptation hardware design
  2026 : FEB-RDO communication validation based on Salsa2 prototype systems
  2027 : RDO production
Services
  2024 : Identify clearly keep-in-out zones
  2025 : Validate LV distribution & come out with cooling scheme, plus interlocks
    Profit from rich expertise within the ePIC groups on magnetic field tolerant DC/DC regulators
  2026-2027 : production
System
  2027 : Fully instrumented slice for every group
    including FEB prototypes, at least two RDOs and a DAM
  2028 : Full system chain validation
    Including slow control and monitoring