## Slide 1
HRPPD Progress Report – March 21, 2024
Arshak Asaturyan, Beni Zihlmann, Carl Zorn (presenter) - JLAB
Alexander Kiselev (BNL)
March 21, 2024
HRPPD #15 – EIC #1

## Slide 2
JLAB purchased 5 initial samples of current generation of HRPPD
@ JLAB  Show that the device is “alive” so that it can “approved” and payment made to Incom
Collect data on all 1024 channels (if possible) and whatever other data can be collected in reasonable time frame
Then ship to BNL for full test program
Eventually JLAB will build full test station (Arshak Asaturyan)
Yale is also preparing full test station
Initial JLAB Test Program

## Slide 3
Two setups
Current – ARC L215 - Detector Group Lab
In preparation for future full test program

## Slide 4
Current Test Setup
HRPPD
HV distribution
32 of 1024 channels readout
Photek MCP reference
Pulsed 405 nm laser
3 filter wheels + diffuser

## Slide 5
Much of HRPPD readout equipment is on loan from BNL – HRPPD enclosure, 3D-printed parts, interposers, readout board + connectors, MMCX adapters, MCX cables – thank you Alexander
We have VME crate with QDC and new V1742 digitizer (both 32 channels) – but need JLAB DAQ personnel to finish the JLAB software/hardware setup
  Have several oscilloscopes on hand
Initial data for now based on oscilloscope measurements (transit time variation) and dark rate studies with NIM electronics
Measurement test program is still evolving
Note – every HRPPD sample comes with test report generated by Incom
We have weekly Wednesday meetings with Incom to discuss status
For those with JLAB computer access https://logbooks.jlab.org/book/hrppd
Status of JLAB setup

## Slide 6
Setup in Darkbox

## Slide 7
GRND
PC
EXBOT
EXTOP
ENB
ENT
GRND
PC
EXBOT
EXTOP
ENB
ENT

## Slide 8
Samtec interposers
Install interposers between anode array and readout board

## Slide 9
GROUND
EXBOT
EXTOP
ENB
ENT
PC
Incom HV scheme

## Slide 10
 | HRPPD 15 (1) | HRPPD 16 (2) | HRPPD 17 (3) | HRPPD 18 (4)
PC | -2,150 | -2,100 | -1,950 | -2,200
ENT | -1,950 | -1,700 | -1,650 | -2,100
ENB | -1,175 | -1,050 | -1,025 | -1,250
EXTOP | -975 | -850 | -825 | -1,050
EXBOT | -200 | -200 | -200 | -200
PC MAX | -2,350 | -2,100 | -2,050 | -2,500
Incom recommended HV values

## Slide 11
HRPPD 15
HRPPD 16

## Slide 12
HRPPD 17
HRPPD 18

## Slide 13
2-56 screws attached to ceramic

## Slide 14
when tightening small nylon 2-56 nuts – found that several of the screws would simply pop off with very small torque
In addition – some HV power pins would also come off when installing or removing HV connections despite the care taken
First two (#15 & #16) shipped back for repairs
Third (#17) was operated and (noisy) signal seen despite ground pin broken – afterwards sent back to Incom for repair
Fourth (#18) is first to operated with minimum problem despite loss of some screws – and excessive glue made it difficult to attach readout board - it is also being returned to Incom (5 screws popped off)
Repaired #15 and #16 are back onsite - #15 operating successfully
Some Notes for first four HRPPD samples

## Slide 15
Missing
screws
Missing
screws
HRPPD 18

## Slide 16
Example pulses
Laser trigger

## Slide 17
Dark Rate Measurements
Incom uses oscilloscope directly – measures the Trigger rate on dark pulses > 4 mV

## Slide 20
Transit time variation (TTV) measured with scope with SPE levels of light - laser intensity attenuated to 4x10-5 (neutral density filters) plus diffusion filter + black mask with 3 mm hole
Laser trigger rate = 3 kHz
TTV = 125 ps for HRPPD #17 and 102 ps for HRPPD #15
Laser Timing Parameters

## Slide 21
Transit Time Variation
HRPPD 18
125 ps
Incom
82 ps

## Slide 22
HRPPD 15
102 ps
Transit Time Variation
Incom
71 ps

## Slide 23
Lots of time lost in dealing with initial mechanical problems – spending time on repair cycle - these issues needs to plateau at some point
Need to get full DAQ up and running for efficient test system
Need criteria for Pass/Fail of HRPPD for QA purpose  pay for module before shipment to full test center (BNL) – BNL is very eager to start their program
  Test signals from all 1,024 channels?
Summary

## Slide 24
Appendix

## Slide 25
Photek LPG-405 pulsed laser
405 nm - pulse width = 75 ps
Freq range = single shot to 300 kHz
Class 1(!!!) – inherently eye safe

## Slide 27
HRPPD 15 - repaired
Visual inspection
✅

## Slide 28
HRPPD 16 - repaired
Visual inspection
✅

## Slide 29
Occasional Noise

## Slide 30
A9
A7
A6
A5
A4
A2
A1
A20
A19
A17
A16
A15
A14
A12
A11
A10
B9
B7
B6
B5
B4
B2
B1
B20
B19
B17
B16
B15
B14
B2
B11
B10
C9
C7
C6
C5
C4
C2
C1
C20
C19
C17
C16
C15
C14
C12
C11
C10
D9
D7
D6
D5
D4
D2
D1
D20
D19
D17
D16
D15
D14
D12
D11
D10
A8
A3
A18
A13
B8
B3
B18
B13
C8
C3
C18
C13
D8
D3
D18
D13
GROUND
A
B
C
D
1
5
10
15
20
GRND
PC
EXBOT
EXTOP
ENB
ENT
Readout Map