## Slide 1: SiPM Prototype Board

## Slide 4: 6010/6015 Boards
Unfortunately I have not had success getting a good IV curve from these boards.
Still debugging issues
  1 board is shorted between Vbias and ground
  3 boards do not have Vbias shorted, but I don’t see any change in current going up to Vop. I suspect output is shorted to something or not connected.

## Slide 8
Chart: Chart 3010PS IV curve
| 37 | 37.25 | 37.5 | 37.75 | 38 | 38.25 | 38.5 | 38.75 | 39 | 39.25 | 39.5 | 39.75 | 40 | 40.25 | 40.5 | 40.75 | 41 | 41.25 | 41.5 | 41.75 | 42 | 42.25 | 42.5 | 42.75 | 43 | 43.25 | 43.5 | 43.75 | 44
SiPM 5 | 0.007759 | 0.007941 | 0.008108 | 0.008333 | 0.008539 | 0.008763 | 0.009004 | 0.009289 | 0.009686 | 0.010174 | 0.01077 | 0.011448 | 0.012241 | 0.01311 | 0.014086 | 0.015182 | 0.016363 | 0.017767 | 0.019145 | 0.020738 | 0.022536 | 0.024456 | 0.026579 | 0.028912 | 0.031295 | 0.034148 | 0.036984 | 0.040303 | 0.043776
SiPM 6 | 0.008168 | 0.008337 | 0.008531 | 0.00873 | 0.008942 | 0.009167 | 0.009411 | 0.009709 | 0.010131 | 0.010667 | 0.011317 | 0.012047 | 0.012928 | 0.013856 | 0.014987 | 0.016187 | 0.01751 | 0.018963 | 0.020522 | 0.022287 | 0.024321 | 0.026374 | 0.02868 | 0.031159 | 0.033825 | 0.036917 | 0.039997 | 0.043625 | 0.047647
SiPM 7 | 0.008077 | 0.008249 | 0.008424 | 0.008615 | 0.008819 | 0.009037 | 0.009277 | 0.009569 | 0.009975 | 0.010503 | 0.011109 | 0.011847 | 0.012687 | 0.013623 | 0.014709 | 0.015877 | 0.017232 | 0.018632 | 0.02021 | 0.021962 | 0.023874 | 0.025936 | 0.028225 | 0.030692 | 0.033352 | 0.036419 | 0.039655 | 0.043044 | 0.04702
SiPM 8 | 0.008263 | 0.008424 | 0.00861 | 0.008811 | 0.00902 | 0.009237 | 0.009498 | 0.009789 | 0.01022 | 0.010766 | 0.011431 | 0.012194 | 0.013062 | 0.014033 | 0.0151 | 0.016325 | 0.01761 | 0.019137 | 0.020697 | 0.022485 | 0.024397 | 0.026533 | 0.028823 | 0.031326 | 0.034057 | 0.036976 | 0.04025 | 0.043823 | 0.047835
SiPM 9 | 0.008457 | 0.008622 | 0.008805 | 0.009002 | 0.009206 | 0.00943 | 0.009683 | 0.00999 | 0.01043 | 0.011001 | 0.011657 | 0.012452 | 0.013304 | 0.014321 | 0.015407 | 0.016612 | 0.01799 | 0.019495 | 0.021031 | 0.022718 | 0.024705 | 0.026879 | 0.029087 | 0.031543 | 0.034357 | 0.037272 | 0.040608 | 0.044227 | 0.048023
SiPM 10 | 0.008515 | 0.008678 | 0.008824 | 0.009012 | 0.009209 | 0.009422 | 0.009656 | 0.009931 | 0.010325 | 0.010846 | 0.011452 | 0.012167 | 0.012942 | 0.013873 | 0.014908 | 0.01602 | 0.017307 | 0.018601 | 0.020097 | 0.021699 | 0.023483 | 0.025399 | 0.027546 | 0.029875 | 0.032307 | 0.035232 | 0.038084 | 0.041481 | 0.044854
SiPM 11 | 0.00865 | 0.008801 | 0.008975 | 0.009159 | 0.009344 | 0.009566 | 0.009793 | 0.010063 | 0.010458 | 0.010963 | 0.011579 | 0.012297 | 0.013118 | 0.014036 | 0.015059 | 0.016225 | 0.017513 | 0.018936 | 0.020472 | 0.022167 | 0.023972 | 0.025962 | 0.028208 | 0.030511 | 0.03313 | 0.036055 | 0.039056 | 0.042275 | 0.045986
SiPM 12 | 0.008185 | 0.008322 | 0.008496 | 0.008662 | 0.008853 | 0.009053 | 0.009271 | 0.009553 | 0.00996 | 0.010486 | 0.011133 | 0.011878 | 0.012769 | 0.013772 | 0.014913 | 0.016133 | 0.017596 | 0.019055 | 0.020801 | 0.022541 | 0.024507 | 0.026705 | 0.029127 | 0.031629 | 0.034473 | 0.037504 | 0.040729 | 0.044401 | 0.048239

## Slide 9
Chart: 3015 IV Curve
| 37 | 37.25 | 37.5 | 37.75 | 38 | 38.25 | 38.5 | 38.75 | 39 | 39.25 | 39.5 | 39.75 | 40 | 40.25 | 40.5 | 40.75 | 41 | 41.25 | 41.5 | 41.75 | 42 | 42.25 | 42.5 | 42.75 | 43
SiPM 45 | 0.005777 | 0.005895 | 0.006033 | 0.00616 | 0.006298 | 0.006449 | 0.006624 | 0.006985 | 0.00775 | 0.008831 | 0.010166 | 0.011781 | 0.013647 | 0.01578 | 0.018083 | 0.020702 | 0.023561 | 0.026504 | 0.029953 | 0.033668 | 0.037582 | 0.041775 | 0.046506 | 0.051079 | 0.056449
SiPM 46 | 0.0061 | 0.006207 | 0.006343 | 0.006478 | 0.006622 | 0.006779 | 0.006957 | 0.007301 | 0.008064 | 0.009153 | 0.010579 | 0.012205 | 0.014146 | 0.016305 | 0.018761 | 0.021388 | 0.024506 | 0.027726 | 0.030951 | 0.034869 | 0.039023 | 0.043469 | 0.04828 | 0.053279 | 0.058845
SiPM 47 | 0.006019 | 0.008129 | 0.006256 | 0.00639 | 0.006535 | 0.006692 | 0.006874 | 0.007217 | 0.007956 | 0.009007 | 0.010327 | 0.011856 | 0.013609 | 0.015617 | 0.018012 | 0.020379 | 0.023025 | 0.025988 | 0.029003 | 0.03268 | 0.036556 | 0.040728 | 0.045032 | 0.049737 | 0.054982
SiPM 48 | 0.006418 | 0.006487 | 0.006588 | 0.006696 | 0.006816 | 0.006941 | 0.007094 | 0.007373 | 0.008059 | 0.009147 | 0.010503 | 0.01215 | 0.014065 | 0.016285 | 0.018733 | 0.021529 | 0.024422 | 0.027698 | 0.031244 | 0.035056 | 0.039319 | 0.043866 | 0.048721 | 0.054017 | 0.059613
SiPM 49 | 0.006228 | 0.006337 | 0.006453 | 0.006592 | 0.006734 | 0.006888 | 0.007062 | 0.007382 | 0.008129 | 0.009222 | 0.010646 | 0.012285 | 0.014242 | 0.016463 | 0.018794 | 0.0216 | 0.024706 | 0.027903 | 0.031362 | 0.035239 | 0.039258 | 0.043646 | 0.048506 | 0.053844 | 0.059254
SiPM 50 | 0.006166 | 0.006277 | 0.006399 | 0.006532 | 0.006669 | 0.006816 | 0.006992 | 0.007357 | 0.00813 | 0.009228 | 0.010606 | 0.012257 | 0.014083 | 0.016312 | 0.018746 | 0.021451 | 0.024271 | 0.027454 | 0.030781 | 0.034755 | 0.038848 | 0.042889 | 0.047606 | 0.052574 | 0.057787
SiPM 51 | 0.006229 | 0.006342 | 0.006458 | 0.006585 | 0.006722 | 0.006872 | 0.007038 | 0.007328 | 0.008057 | 0.009186 | 0.010642 | 0.012475 | 0.014475 | 0.016786 | 0.019445 | 0.022297 | 0.025596 | 0.029022 | 0.03263 | 0.036688 | 0.041206 | 0.045877 | 0.051076 | 0.056366 | 0.062193
SiPM 52 | 0.006439 | 0.006507 | 0.006601 | 0.006704 | 0.006819 | 0.006947 | 0.007101 | 0.00742 | 0.008181 | 0.009253 | 0.010671 | 0.012277 | 0.014172 | 0.0164 | 0.01876 | 0.021405 | 0.024383 | 0.027636 | 0.031196 | 0.034919 | 0.038911 | 0.043552 | 0.048376 | 0.053306 | 0.058726

## Slide 10
Chart: Full 20x 3015 SiPM board IV curve
| 35 | 35.25 | 35.5 | 35.75 | 36 | 36.25 | 36.5 | 36.75 | 37 | 37.25 | 37.5 | 37.75 | 38 | 38.25 | 38.5 | 38.75 | 39 | 39.25 | 39.5 | 39.75 | 40 | 40.25 | 40.5 | 40.75 | 41 | 41.25 | 41.5 | 41.75 | 42 | 42.25 | 42.5 | 42.75 | 43
Trial 4 with pre-amp 1 | 0.11803 | 0.002314 | 0.12441 | 0.12585 | 0.13008 | 0.10148 | 0.24223 | 0.076081 | 0.25013 | 0.052907 | 0.1473 | 0.25035 | 0.15406 | 0.15707 | 0.15962 | 0.16569 | 0.181 | 0.20803 | 0.17414 | 0.26185 | 0.33445 | 0.41347 | 0.45379 | 0.58107 | 0.60073 | 0.70651 | 0.77278 | 0.86954 | 0.9727 | 1.0844 | 1.207 | 1.3333 | 1.4702
Trial 5 with pre-amp 1 | 0.21931 | 0.11996 | 0.12347 | 0.12586 | 0.12784 | 0.20322 | 0.04426 | 0.25266 | 0.049509 | 0.14248 | 0.14622 | 0.14906 | 0.15212 | 0.15521 | 0.15472 | 0.16719 | 0.16986 | 0.21544 | 0.22402 | 0.23809 | 0.33237 | 0.38855 | 0.50651 | 0.46636 | 0.56639 | 0.70571 | 0.76833 | 0.86318 | 0.96507 | 1.0763 | 1.1969 | 1.3241 | 1.4599
20x SiPM 52 | 0.12878 | 0.13014 | 0.13202 | 0.13408 | 0.13638 | 0.13894 | 0.14202 | 0.1484 | 0.16362 | 0.18506 | 0.21342 | 0.24554 | 0.28344 | 0.328 | 0.3752 | 0.4281 | 0.48766 | 0.55272 | 0.62392 | 0.69838 | 0.77822 | 0.87104 | 0.96752 | 1.0661 | 1.1745
Trial 4 and 5 are separated by about 5 minutes

## Slide 12
10% to 90%
Fall time: 9.5 ns Rise time: 98.5 ns

## Slide 14: 8x8x40 mm LYSO crystal array
Created a 2x2x2 array of LYSO crystals each crystal is 4x4x20 mm in size. I placed a ~0.1µCi Co60 source on top of the crystal array.
Lu176 in LYSO creates many gammas and electrons mostly with energy below 300keV
I see the excess for the 2 Co60 gammas but didn’t take enough data to tell if we can resolve them.

## Slide 15: LYSO crystal decays