## Slide 1
RRPL Operations
RRPL - Week of 09/02:
  Production
    Ac-225 Pre-Processing
    Researchers
      Separations and Labeling R&D
      F&O
        Strobic, Exhaust Fan maintenance, Remove line voltage thermostat and install blank cover over box; F&O installing new exhaust Fan & HEPA filter system for 2-66C- roof top work
    Coming weeks:
      LANL Th target processing campaigns early and late Sept.

## Slide 2
BLIP Operation – OFF Season Mode
  Chillers require 24-hr for crankcase heater warmup prior to operation. Therefore, cannot immediately switch to backup chiller until installation of electrical panel is performed in the next phase of cooling system upgrade.
  Off-season maintenance/testing.
  Scheduling chillers 1 + 2 Preventative Maintenance for late September.
  Subsystems and instrumentation restarted after power shut down last Saturday

## Slide 3
  Commissioning Path Forward: Commissioning has restarted.
    RF breaker trip issue
  - USI for RF bypass and testing: RF is operational – needs fine tuning from manufacturer (Sep. 8th).
  USI in development for new access controls ‘testing/tuning mode’ – USI drafted/incorporated ACG technical details; working with C-AD ESSHQ Div. Head to refine document.
  Glycol pump replacement - Coordinating with Facility project manager/ AC group to top up glycol in cooling system.
  New robot - Robot testing/programming in progress in vault room.
  USI for ‘additional actions/mitigations under RF Bypass USI’ – BHSO approved RSI on 8/12 to conduct setup and testing of Non-Radiofrequency system components at the Cyclotron. Testing Ion source injection system. Tested filament power supply – failed, requested quotes for replacements - 8/26 Configuring temporary filament power supply borrowed from power supply group- 9/2.
  Vault room Differential pressure. New Seals installed around doors into MER. Dp levels fluctuating slightly from setpoint. Continuing to monitor.
  Hitachi Air compressor in MER - Ordering replacement scroll head.
Cyclotron Activities