## Slide 1
RRPL Operations
RRPL - Week of 09/16:
    IVR – Implementation Verification Review this week- reviewing proper controls implemented to cover TSR's.
  Production
    LANL target processing has been completed
    Dispensing and Th waste processing this week.
    F&O: monthly generator test to simulate power failure completed
    Coming weeks:
      Another LANL Th target processing campaign late Sept.

## Slide 2
BLIP Operation – OFF Season Mode
  Chillers require 24-hr for crankcase heater warmup prior to operation. Therefore, cannot immediately switch to backup chiller until installation of electrical panel is performed in the next phase of cooling system upgrade.
  Off-season maintenance/testing:
      Scheduling chillers 1 + 2 Preventative Maintenance for the week of September 22.
      This week weekly alarms tests performed. No issues
      This week the annual alarms tests will be performed

## Slide 3
  Commissioning Path Forward: Commissioning has restarted.
    Ion Source Testing – Replacement filament power supply installed and tested 9/5. Gate valve replaced, ion source successfully tested. Arcing is an issue, cause yet to be identified.
    East Side Gate Valve XV01 not functioning, replacement being procured (16-week lead time)
    RF breaker trip issue
  - USI in place for RF bypass and testing: RF is operational. RF Control box working temporarily but needs calibration from manufacturer
  USI for new access controls ‘testing/tuning modes’ – scheduled to be reviewed by committees
  New robot - Robot testing/programming in progress in vault room.
  USI for ‘additional actions/mitigations under RF Bypass USI’ – BHSO approved RSI on 8/12 to conduct setup and testing of Non-Radiofrequency system components at the Cyclotron.
  Hitachi Air compressor in MER - Ordering replacement scroll head.
Cyclotron Activities