## Slide 1
RRPL Operations
RRPL - Week of 10/20:
  Production activities demonstrations for CCR
    Coming weeks:
    Ni/Co Hot Cell set-up in TPL; preparation for BLIP irradiated foil processing
    Be-7 processing in TPL
    F&O:
      Generator Test, Generator Maintenance, Fuel tank Inspection
      Exhaust Fan Maintenance EF1, EF2, FE7 & EF8
      AC1 & AC2 Winterization PM

## Slide 2
BLIP Operation – OFF Season Mode
  Chillers require 24-hr for crankcase heater warmup prior to operation. Therefore, cannot immediately switch to backup chiller until installation of electrical panel is performed in the next phase of cooling system upgrade.
  Off-season maintenance/testing
      This week weekly alarms testing
      Rad monitors not communicating with C-AD Main Control room - still on-going
      dP indication on target water filter not reading correctly. SME from water group discussed purchasing new calibrated gauges and calibrating transducer – still on-going
      Preventive manipulator maintenance - on-going

## Slide 3
  Commissioning Path Forward: Commissioning has restarted.
    Cyclotron is working under module 3 of the commissioning plan.
    USI for new access controls ‘testing/tuning modes’ – LESHC meeting took place 10/9, 1 week to review (in progress)
  Cobot (robotic arm for target movement) testing paused, will resume when access to vault is available.
  Next step in commissioning plan is Ni target irradiation. Target ready, canning record approved, need RWP to cover target removal operation (in progress)
Cyclotron Activities