## Slide 1
RRPL Operations
RRPL - Week of 12/22:
    Special Operating Restriction: BNL has issued a pause in all operations associated with radioactive materials in RRPL and BLIP hot cells/boxes covered under the current SAD/ASE until further notice
    LANL Th target waste neutralized, solidification after special operation restrictions are resolved.
    Coming weeks (pending continuation of operations):
    Ni/Co Hot Cell set-up in TPL; preparation for BLIP irradiated foil processing
    Be-7 processing in TPL

## Slide 2
BLIP Operation – OFF Season Mode
  Special Operation Restriction: BNL has issued a pause in all operations associated with radioactive materials in RRPL and BLIP hot cells/boxes covered under the current SAD/ASE until further notice.
  Issue found with rad wastewater tank alarm – persistent since last test of the alarm
  Off-season maintenance/testing
      This week weekly alarms testing
      dP indication on target water filter is now reading correctly after some adjustments but still working with SME from water group to purchase new calibrated gauges and calibrating transducer
      Request to water group to regenerate/replace the resin in the makeup water deionizer.

## Slide 3
  Commissioning Path Forward: Commissioning in progress.
  Cyclotron is working under module 3 of the commissioning plan.
    Next step in commissioning plan is Ni target irradiation. Target ready, canning record approved, need RWP to cover target removal operation (in progress)
  Cobot (robotic arm for target movement) configuring software & hardware for remote operation from control room. Parts ordered, new remote pendant nearing completion
  Intermittent RF trips attributed to blower (air cooling) flow switch – has been recalibrated
  WO submitted for filter and fluid change of hydraulic pump.
  RF/dummy load interrupt flow switch installed and tested.
Cyclotron Activities