## Slide 1
Data courtesy C. Naylor
Chart: C-AD ACTIVITY DECEMBER 2025
| FY26-week 10: | FY26-week 11: | FY26-week 12: | FY26-week 13:
Physics | 98.43 | 94.05 | 104 | 0
Machine Development | 0 | 2.02 | 0 | 0
Beam Studies | 14.7 | 0 | 12.83 | 0
Experimental setup | 0.17 | 22.57 | 0 | 0
Setup | 13.97 | 33.97 | 22.79 | 0
Scheduled Maintenance | 8.75 | 0 | 0 | 0
Scheduled Shutdown | 0 | 0 | 0 | 0
Unscheduled shutdown | 0 | 0 | 0 | 0
Machine failures | 31.98 | 15.39 | 28.38 | 0

## Slide 2
Data courtesy C. Naylor
Chart: C-AD ACTIVITY NOVEMBER 2025
| FY26-week 06: | FY26-week 07: | FY26-week 08: | FY26-week 09:
Physics | 118.23 | 121.42 | 124.4 | 130.6
Machine Development | 0 | 0 | 0 | 0
Beam Studies | 14.65 | 0 | 9.58 | 0
Experimental setup | 0 | 0 | 0 | 0
Setup | 15.93 | 16.52 | 15.6 | 17.08
Scheduled Maintenance | 0 | 10.62 | 0 | 9.67
Scheduled Shutdown | 0 | 0 | 0 | 0
Unscheduled shutdown | 0 | 0 | 0 | 0
Machine failures | 19.19 | 19.44 | 18.42 | 10.65

## Slide 3
Data courtesy C. Naylor
Chart: C-AD ACTIVITY October 2025
| FY26-week 01: | FY26-week 02: | FY26-week 03: | FY26-week 04: | FY26-week 05:
Physics | 15.37 | 112.82 | 95.1 | 113.1 | 117.75
Machine Development | 0 | 0 | 0 | 0 | 0
Beam Studies | 0 | 0 | 15.12 | 10.85 | 0
Experimenter setup | 0 | 0.1 | 0 | 0.67 | 0
Setup | 5.75 | 15.52 | 17.43 | 14.1 | 19.32
Scheduled Maintenance | 130.85 | 0 | 0 | 0 | 8.92
Scheduled Shutdown | 0 | 0 | 0 | 0 | 0
Unscheduled shutdown | 0 | 0 | 0 | 0 | 0
Machine failures | 16.03 | 39.56 | 40.35 | 29.28 | 23.01

## Slide 4
Data courtesy C. Naylor
Chart: C-AD ACTIVITY SEPTEMBER 2025
| FY25-week 49: | FY25-week 50: | FY25-week 51: | FY25-week 52:
Physics | 0 | 111.13 | 110.65 | 41.07
Machine Development | 0 | 0 | 0 | 0
Beam Studies | 0 | 0 | 11.68 | 0
Experimental setup | 0 | 0 | 0 | 0
Setup | 1.88 | 18.72 | 15.23 | 5.02
Scheduled Maintenance | 153.55 | 0 | 2.67 | 81.87
Scheduled Shutdown | 0 | 0 | 0 | 0
Unscheduled shutdown | 0 | 0 | 0 | 0
Machine failures | 12.57 | 38.15 | 27.77 | 40.04

## Slide 5
Data courtesy C. Naylor
Chart: C-AD ACTIVITY AUGUST 2025
| FY25-week 44: | FY25-week 45: | FY25-week 46: | FY25-week 47: | FY25-week 48:
Physics | 109.68 | 122.68 | 126.78 | 110.3 | 0
Machine Development | 0 | 0 | 0 | 0 | 0
Beam Studies | 0 | 13.93 | 0 | 12.48 | 0
Experimental setup | 0 | 0.48 | 0.2 | 0 | 0
Setup | 19.07 | 17.67 | 17.65 | 15.57 | 0
Scheduled Maintenance | 10.72 | 0 | 9.98 | 0 | 164.3
Scheduled Shutdown | 0 | 0 | 0 | 0 | 1.5
Unscheduled shutdown | 0 | 0 | 0 | 0 | 0
Machine failures | 28.53 | 13.24 | 13.39 | 29.65 | 2.2

## Slide 6
Data courtesy C. Naylor
Chart: C-AD ACTIVITY July 2025
| FY25-week 40: | FY25-week 41: | FY25-week 42: | FY25-week 43:
Physics | 94.82 | 96.75 | 105.75 | 113.3
Machine Development | 8.62 | 5.52 | 0 | 0
Beam Studies | 0 | 7.62 | 0 | 8.15
Experimenter Setup | 0.8 | 0.08 | 0 | 0.23
Machine Setup | 18.58 | 18.57 | 22.9 | 18.42
Scheduled Maintenance | 11.5 | 2.33 | 8 | 0
Scheduled Shutdown | 0 | 0 | 0 | 0
Unscheduled Shutdown | 0 | 0 | 0 | 0
Machine Failures | 33.68 | 37.13 | 31.35 | 27.9

## Slide 7
Data courtesy C. Naylor
Chart: C-AD ACTIVITY JUNE 2025
| FY25-week 36: | FY25-week 37: | FY25-week 38: | FY25-week 39:
Physics | 0 | 51.73 | 100.12 | 91.68
Machine Development | 0 | 2.52 | 3.45 | 0
Beam Studies | 0 | 0 | 0 | 11.02
Experimental setup | 19.97 | 2.7 | 0.42 | 0.1
Setup | 48.97 | 17.98 | 25.03 | 19.62
Scheduled Maintenance | 63.27 | 64.9 | 0 | 0
Scheduled Shutdown | 0 | 0 | 0 | 0
Unscheduled shutdown | 0 | 0 | 0 | 0
Machine failures | 35.79 | 28.17 | 38.98 | 45.58

## Slide 8
Data courtesy C. Naylor
Chart: C-AD ACTIVITY MAY 2025
| FY25-week 31: | FY25-week 32: | FY25-week 33: | FY25-week 34: | FY25-week 35:
Physics | 0 | 0 | 0 | 0 | 0
Machine Development | 0 | 0 | 0 | 0 | 0
Beam Studies | 0 | 0 | 0 | 0 | 0
Experimental setup | 0 | 0 | 0 | 0 | 1.22
Setup | 0 | 0 | 0 | 0 | 31.97
Scheduled Maintenance | 168 | 168 | 168 | 168 | 134.52
Scheduled Shutdown | 0 | 0 | 0 | 0 | 0
Unscheduled shutdown | 0 | 0 | 0 | 0 | 0
Machine failures | 0 | 0 | 0 | 0 | 0.3

## Slide 9
Chart: C-AD ACTIVITY April 2025
| FY25-week 27: | FY25-week 28: | FY25-week 29: | FY25-week 30:
Physics | 0 | 0 | 0 | 0
Machine Development | 0 | 0 | 0 | 0
Beam Studies | 0 | 0 | 0 | 0
Experimenter Setup | 0 | 0 | 0 | 0
Machine Setup | 0 | 0 | 0 | 0
Scheduled Maintenance | 168 | 168 | 168 | 168
Scheduled Shutdown | 0 | 0 | 0 | 0
Unscheduled Shutdown | 0 | 0 | 0 | 0
Machine Failures | 0 | 0 | 0 | 0
Data courtesy C. Naylor

## Slide 10
Chart: C-AD ACTIVITY MARCH 2025
| FY25-week 22: | FY25-week 23: | FY25-week 24: | FY25-week 25: | FY25-week 26:
Physics | 0 | 0 | 0 | 0 | 0
Machine Development | 0 | 0 | 0 | 0 | 0
Beam Studies | 0 | 0 | 0 | 0 | 0
Experimental setup | 0 | 0 | 0 | 0 | 0
Setup | 0 | 0 | 0 | 0 | 8.17
Scheduled Maintenance | 0 | 0 | 0 | 0 | 24
Scheduled Shutdown | 168 | 167 | 168 | 168 | 41.15
Unscheduled shutdown | 0 | 0 | 0 | 0 | 0
Machine failures | 0 | 0 | 0 | 0 | 94.68

## Slide 11
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- MARCH 2025
| 02/25/25/2 to 03/04/25/1 | 03/04/25/2 to 03/11/25/1 | 03/11/25/2 to 03/18/25/1 | 03/18/25/2 to 03/25/25/1 | 03/25/25/2 to 04/01/25/1
PPS_AGS | 0.32154
RfBooster | 0.21877
ESHQ | 0.089256
PPS_Booster | 0.003403
PS_RHIC | 0.28128
sum<1hr | 0.0063199

## Slide 12
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- April 2025
| 04/01/25/2 to 04/08/25/1 | 04/08/25/2 to 04/15/25/1 | 04/15/25/2 to 04/22/25/1 | 04/22/25/2 to 04/29/25/1
.
.
.
.
.
.
.
.
.
.
.
.
Data courtesy C. Naylor

## Slide 13
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- MAY 2025
| 04/29/25/2 to 05/06/25/1 | 05/06/25/2 to 05/13/25/1 | 05/13/25/2 to 05/20/254/1 | 05/20/25/2 to 05/27/25/1 | 05/27/25/2 to 06/03/25/1
.
.
.
.
.
.
.
.
.
.
.
.

## Slide 14
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM - JUNE 2025
| 06/03/25/2 to 06/10/25/1 | 06/10/25/2 to 06/17/25/1 | 06/17/25/2 to 06/24/25/1 | 06/24/25/2 to 07/01/25/1
PS_RHIC | 0.03422 | 0.015097 | 0.0057003 | 0.023243
PS_Bstr/BtA/LtB | 0.0054061 | 0.015814
PPS_Booster | 0.0027214
CntrlsHdRHIC | 0.0039534 | 0.0053877
FoilChange | 0.0032179
RfRHIC | 0.014251 | 0.0032915 | 0.005498
WaterRhicRf | 0.025412
CoolACRHIC | 0.0069875
PPS_RHIC | 0.014766
CntrlsSftwr | 0.0018388
CryoRHIC | 0.0030892
InstRHIC | 0.0037512 | 0.0030156

## Slide 15
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- July 2025
| 07/01/25/2 to 07/08/25/1 | 07/08/25/2 to 07/15/25/1 | 07/15/25/2 to 07/22/25/1 | 07/22/25/2 to 07/29/25/1
PS_RHIC | 0.017319 | 0.0098897 | 0.010243 | 0.0041066
RfRHIC | 0.0038913 | 0.0060446 | 0.0016919
ES&FD_Exp | 0.016211 | 0.0037529
QuenchProtect | 0.0076595
PPS_RHIC | 0.0046603
InstRHIC | 0.0023532
RfAGS | 0.0019841 | 0.0032145 | 0.0023994
ElecRHIC | 0.0027685 | 0.0058908
ACG_RHIC | 0.0028916 | 0.0028916
HumanError | 0.0040297
RadMonIntlk | 0.0031992 | 0.0015842
Weather | 0.014289

## Slide 16
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- AUGUST 2025
| 07/29/25/2 to 08/05/25/1 | 08/05/25/2 to 08/12/25/1 | 08/12/25/2 to 08/19/25/1 | 08/19/25/2 to 08/26/25/1 | 08/26/25/2 to 09/02/25/1
PS_RHIC | 0.010543 | 0.0082938
PPS_RHIC | 0.0094262
QuenchDetect | 0.0024943
FoilChange | 0.0019281
PS_Experiment | 0.0022647
TndmSource | 0.0037796
RadMonIntlk | 0.0019281
VacRHIC | 0.0032441
ES&FD_Exp | 0.0045754 | 0.0068707 | 0.0033665
RfRHIC | 0.0029839 | 0.010712 | 0.0023259
CryoRHIC | 0.0047437 | 0.0036266
Weather | 0.0017904

## Slide 17
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM - SEPTEMBER 2025
| 09/02/25/2 to 09/09/25/1 | 09/09/25/2 to 09/16/25/1 | 09/16/25/2 to 09/23/25/1 | 09/23/25/2 to 09/30/25/1
PS_RHIC | 0.021018 | 0.036413 | 0.013044
PS_Bstr/BtA/LtB | 0.005001
CntrlsHdRHIC | 0.0065682 | 0.0056463
PPS_RHIC | 0.0088728 | 0.040654
RfRHIC | 0.029084 | 0.0063377
ElecRHIC (power dip) | 0.0072596
ES&FD_Exp | 0.0027886
TMP7 | 0.0032265
InstRHIC | 0.0037796
ES&FD_AtR | 0.0028808
QLI | 0.0092646 | 0.0042866
VacRHIC | 0.037404

## Slide 18
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- October 2025
| 09/30/25/2 to 10/07/25/1 | 10/07/25/2 to 10/14/25/1 | 10/14/25/2 to 10/21/25/1 | 10/21/25/2 to 10/28/25/1 | 10/28/25/2 to 11/04/25/1
PS_RHIC | 0.0079004 | 0.025912 | 0.014332 | 0.0044921
ES&FD_AtR | 0.0044208 | 0.0019394 | 0.0019252
ES&FD_Exp | 0.0037506 | 0.003993 | 0.0080715 | 0.01051
PPS_RHIC | 0.0041784 | 0.0099397 | 0.013405 | 0.0030518 | 0.0028236
VacRHIC | 0.0026097
QLI | 0.0031373
QuenchDetect | 0.0043923 | 0.014617
RfRHIC | 0.0065884 | 0.0040643
PS_AGS | 0.0068166
ElecRHIC | 0.0046205
TMP7 | 0.001697
PPS_Booster | 0.0021391

## Slide 19
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM -- NOVEMBER 2025
| 11/04/25/2 to 11/11/25/1 | 11/11/25/2 to 11/18/25/1 | 11/18/25/2 to 11/25/25/1 | 11/25/25/2 to 12/02/25/1
VacRHIC | 0.018352
ES&FD_Exp | 0.0023784 | 0.0018106 | 0.0078869
RfRHIC | 0.0030688 | 0.0035906
QuenchDetect | 0.0022556
CntrlsHdRHIC | 0.002409 | 0.0038974
PPS_AGS | 0.0048027
CoolACRHIC | 0.0023323
ElecRHIC (power dip) | 0.0079483
RadMonIntlk | 0.0020254
QuenchProtect | 0.0069356
PS_RHIC | 0.0091298 | 0.001918
CryoRHIC | 0.0047567

## Slide 20
Data courtesy C. Naylor
Chart: INTEGRATED FAILURES (GREATER THAN ONE HOUR) BY SYSTEM DECEMBER 2025
| 12/02/25/2 to 12/09/25/1 | 12/09/25/2 to 12/16/25/1 | 12/16/25/2 to 12/23/25/1 | 12/23/25/2 to 12/30/25/1
CryoRHIC | 0.0034326
ES&FD_AtR | 0.028773 | 0.0051085
ES&FD_Exp | 0.010399 | 0.010257
ElecRHIC | 0.0024634 | 0.013993
QuenchProtect | 0.0065623
QuenchDetect | 0.00737
QLI | 0.0021201
ES&FD_R | 0.0021201
PPS_RHIC | 0.0057547 | 0.015265
RfRHIC | 0.0038364 | 0.0020394
ACG_RHIC | 0.005896
PS_AGS

## Slide 21: Operations for BLIP
No BLIP
this past week
BLIP remains in shutdown.
Notes: BLIP plot from LogView /MCR/PerformanceStats/timeMeetingBLIP.lvdisp
Avail calc on nsrl time Excel sheet

## Slide 22: NSRL activity past weeks: 12/16-23/2025 (as of 0800)
Availability: 95.3%
NSRL running continues. Various setup and user beam time through the week.
10 EBIS species, Linac protons logged to NSRL
Including H2, d, d2
Run 26-A continues, schedule posted with users up to 30 January.
Chart: NSRL Hours
| Experimental Setup | Failure | Machine Development | Machine Setup | Operator Training | Scheduled Maintenance | Scheduled Shutdown | Science
NSRL Hours | 3.3 | 3.13 | 7.08 | 20.13 | 0 | 0 | 101.33 | 33.03
Notes: Graph and avail calc from nsrl time.xlsx

## Slide 23: Notes from Operations
Notable failures (>1 hr)
AGS MMPS cycloconverter fault, reset
RHIC ACS failure (3x), access for PLC I/O module replacement
Multiple AtR cooling faults, work on flow switches and filter bags
Blue injection kicker PS control hardware failure, 16-bit data converter replaced
Blue abort kicker misfire at injection, undiagnosed
Multiple power dips with weather stand-down and recovery of numerous systems complex-wide
AGS vertical tune quad PS failure; one unpolarized fill during repair
BtA quad PS qf10, qd11 multiple issues requiring retuning beam during repair
Delayed dump for Linac Mod 1 repair, 8618 tube replaced
sPHENIX equipment failure, pump room trip affecting racks in IR
Other notes
Work on AGS RF stations B, KL
Human error, RHIC MMPS current limit key configuration
Linac air compressor and nitrogen supply failure
Access for sPHENIX TPC issues
Other various trips, resets, tuning (12 resets logged, in addition to short failures)
Polarized beam development and other setup behind stores this week
Notes: NSRL plot from LogView /MCR/PerformanceStats/timeMeetingNSRL.lvdisp, select Dose Rate and ion species in use
Screenshot of AlarmSummary

## Slide 24: Reminder from Operations
MCR beam operations continue
  RHIC physics continues this week with polarized protons from LINAC
  NSRL users concluded until early January.
    Protons from Linac OPPIS
    Other ion beams from EBIS including H2, d, d2
  BLIP target production not scheduled until March.
    Beam in BLIP line ORC cancelled, due to pending restart authorization.
  AGS tuning, pp setup, other studies to continue behind stores as time allows.
  MCR shift coverage 24/7
    Please update calling lists for available staff!!
Data courtesy A. Nicola/C.S. Towle

## Slide 25: Keytree notes
Remote keytree systems require a database lookup to validate and authorize distribution of keytokens. The Keytrak system’s database is available for lookup functionality, but can no longer be modified in any way to add or edit qualified personnel.
Personnel affected:
New employees, experimenters arriving after 12/1/2024 could not be enrolled in database
Existing employees & experimenters, not already enrolled, cannot be added or modified in database
Remote keytrees affected:
Iris scan @NSRL
BNLID scan @AGS South gate, North Conjunction Gate
Iris scan @RHIC STAR, sPHENIX, IP12, IP4
Keytrees NOT affected:
BNLID scan @Linac, Booster
BNLID scan @RHIC IP2
MCR keytrees

## Slide 26: Keytree notes, continued
Mitigation:
Choose entrants that are currently enrolled to perform the work in the enclosure under Controlled Access (CA).
Obtain CA keys from MCR for remote gate watch
Arrange for local gate watch using other tokens for entrants.
Arrange for Restricted Access, and allow for delays incurred to sweep areas.
CAUTION
Follow procedures. Gate Watch staff will remain vigilant.
  No “piggy-backing” – one key/token per entrant.
  Access keys are not candy! Authenticated entrant retains possession.
Radiological Controls and the Access Controls System are not to be circumvented. This is (has been) enforced with disciplinary action (dismissal).

## Slide 27
Chart: Availability: RHIC Run-25
| week 25 | week 26 | week 27 | week 28 | week 29 | week 30 | week 31 | week 32 | week 33 | week 34 | week 35 | week 36 | week 37 | week 38 | week 39 | week 40 | week 41 | week 42 | week 43 | week 44 | week 45 | week 46 | week 47 | week 48 | week 49 | week 50 | week 51 | week 52 | week 1 | week 2 | week 3 | week 4 | week 5 | week 6 | week 7 | week 8 | week 9 | week 10 | week 11 | week 12 | week 13
Availability | 0 | 0.079436 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.98845 | 0.65826 | 0.72677 | 0.76798 | 0.72869 | 0.78479 | 0.77588 | 0.80406 | 0.83393 | 0.8186 | 0.92119 | 0.91526 | 0.82351 | 0 | 0.1301 | 0.77292 | 0.83203 | 0.53512 | 0.56851 | 0.76452 | 0.75982 | 0.82571 | 0.85626 | 0.88577 | 0.87648 | 0.89036 | 0.93274 | 0.79918 | 0.90839 | 0.83107 | 0
Average | 1 | 0.079436 | 0.079436 | 0.079436 | 0.079436 | 0.079436 | 0.079436 | 0.079436 | 0.079436 | 0.079436 | 0.26269 | 0.44008 | 0.52788 | 0.60781 | 0.638 | 0.75796 | 0.76087 | 0.76673 | 0.77511 | 0.77966 | 0.79388 | 0.80436 | 0.80597 | 0.80509 | 0.80025 | 0.79814 | 0.80053 | 0.79114 | 0.78779 | 0.78631 | 0.78472 | 0.78704 | 0.75037 | 0.75924 | 0.7991 | 0.80333 | 0.80874 | 0.80835 | 0.81244 | 0.81317 | 0.81317
Goal | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8

## Slide 28
Chart: Mean Time Between Failure, Mean Time To Repair, Average Failure Hours per Day (cryo start to cryo end date)
| MTBF | MTTR | Failure hr/day
Run7 | 5.3 | 2 | 6.3
Run8 | 6.3 | 1.2 | 3.8
Run9 | 5.5 | 1.3 | 4.42
Run10 | 5.5 | 1 | 3.8
Run11 | 5.7 | 1.2 | 4.5
Run12 | 6.9 | 1.1 | 3
Run13 | 6.6 | 1.2 | 3.5
Run14 | 8.6 | 1.3 | 3
Run15 | 8.1 | 1.1 | 3.1
Run16 | 6.4 | 1.2 | 3.6
Run17 | 6.8 | 1.2 | 3.7
Run18 | 9.3 | 0.9 | 2.1
Notes: Reliability comparisons to previous runs

## Slide 29
Chart: Mean Time Between Failure, Mean Time To Repair, Average Failure Hours per Day (cryo start to cryo end date)
| Run7 | Run8 | Run9 | Run10 | Run11 | Run12 | Run13 | Run14 | Run15 | Run16 | Run17 | Run18 | Run19 | Run20 | Run21 | Run22 | Run23 | Run24 | Run25 | Run25alt | Run25full
MTBF | 5.3 | 6.3 | 5.5 | 5.5 | 5.7 | 6.9 | 6.6 | 8.6 | 8.1 | 6.4 | 6.8 | 9.3 | 3.9 | 10.3 | 9.89 | 8.6 | 6.05 | 6.38 | 7.28 | 5.35 | 10.69
MTTR | 2 | 1.2 | 1.3 | 1 | 1.2 | 1.1 | 1.2 | 1.3 | 1.1 | 1.2 | 1.2 | 0.9 | 0.5 | 1 | 0.97 | 1.61 | 2.19 | 1.58 | 1.47 | 15.02 | 1.68
Failure hr/day | 6.3 | 3.8 | 4.42 | 3.8 | 4.5 | 3 | 3.5 | 3 | 3.1 | 3.6 | 3.7 | 2.1 | 3.1 | 2.1014 | 2.112 | 3.5903 | 5.7463 | 4.5433 | 4.0298 | 10.9368 | 2.9869
Notes: Reliability comparisons to previous runs

## Slide 30: Availability, expected beam hours
Notes: Normally hidden. Explains how the 85% from DOE comes in to play, & general availability metrics