## Slide 1: Photosensors for pfRICH and hpDIRC
Alexander Kiselev (BNL)
ePIC pfRICH Detector Subsystem Collaboration Leader
Incremental Preliminary Design and Safety Review of the pfRICH, dRICH and hpDIRC
April 1st - 2nd, 2025

## Slide 2: Outline
Photosensor choice for ePIC pfRICH and hpDIRC
EIC PED contract with Incom
Re-design and a first batch production of DC-coupled EIC HRPPDs
EIC HRPPD evaluation at Incom / JLab / BNL & INFN Trieste
Preparations for Photek MCP-PMT evaluation at the University of Glasgow
Summary & Outlook
2

## Slide 3: Charge Questions Addressed
Are the technical performance requirements appropriately defined and complete for this stage of the project?
Are the plans for achieving detector performance and construction sufficiently developed and documented for the present phase of the project?
Are the current designs and plans for detector and electronics readout likely to achieve the performance requirements with a low risk of cost increases, schedule delays, and technical problems?
Are the fabrication and assembly plans for the various particle identification detector systems consistent with the overall project and detector schedule?
Are the plans for detector integration in the EIC detector appropriately developed for the present phase of the project?
Have ES&H and QA considerations been adequately incorporated into the designs at their present stage?
Have the recommendations from previous reviews been adequately addressed?
3

## Slide 4: Photosensor Choice for pfRICH and hpDIRC
hpDIRC: 12*6*4 = up to 288 2” MCP-PMTs total
pfRICH: 68 4” HRPPDs total
 | pfRICH | DIRC
Baseline solution | DC-coupled HRPPD by Incom Inc. | Auratek MAPMT253 by Photek
Alternative solution | Auratek MAPMT253 by Photek | DC-coupled HRPPD by Incom Inc.
Spatial resolution | Sub-mm | Sub-mm
SPE timing resolution | s of the core part <50ps | <75ps RMS, including tail
Dark count rate | Few kHz/cm2 is acceptable | Few kHz/cm2 is acceptable
Occupancy per event | Small: could work with CC-coupled sensors | 200+ p.e. / event: need DC-coupled sensors
4
Charge Question 1
We would clearly prefer to have the same photosensor solution for both subsystems

## Slide 5
HRPPDs by Incom Inc.
5

## Slide 6: General Information
A uniquely large 120mm x 120mm DC-coupled MCP-PMT
104mm x 104mm active area (~75% geometric efficiency)
5mm thick fused silica (or sapphire) window with a bialkali photocathode
A pair of 10 mm diameter pore MCPs in a chevron configuration
Uniform anode pixellation: 32 x 32 square pads (pitch 3.25mm)
Anode plate vacuum side
Anode plate air side
HRPPD with a passive backplane
6
Substantially re-designed for use in EIC

## Slide 7: EIC HRPPD Assembly
Fused silica window
MCPs, spacers, etc
Side wall
Anode plate, a pre-routing
ceramic circuit board
Interface PCB
Compression interposers
(not shown)
4x4 spots, each with 8x8 square pads; 3.25mm pitch
8x8 pad pattern
compressed
from 3.25mm (inner side)
to 2.00mm (outer side) pitch
7

## Slide 8: 2023 EIC Contract with Incom Inc.
Perform a first stage of the design adaptation to EIC needs
In general, be as much conservative as possible in this first iteration
Focus on a mechanical and electrical interface, including a complete re-design of the anode plate
Produce five HRPPD tiles of a new design
Demonstrate production capability and reproducibility of key performance parameters
Gain, QE, Photon Detection Efficiency (PDE), Dark Count rates (DCR), timing resolution, …
First generation HRPPDs were “bare sensors” & lacked interfacing completely
First generation yielded only a couple of low QE devices out of a dozen of trials
Contract completed successfully in April 2024 (5+2 HRPPDs purchased)
Comprehensive test program ongoing at BNL / JLab / INFN Trieste
8
Link to EIC HRPPD specifications
Charge Question 1

## Slide 9: Incom’s Internal Testing Summary
HRPPD # | Operating Voltage Range (V/MCP) | Dark Rates @ 1X10^6 Gain (Hz/cm^2) | EIC Standard – Dark Rates
15 | 750-850 | 160 | ≤ 2E3Hz/cm^2 at ≥ 10^6 Gain
16 | 625-725 | 0.18 | 
17 | 600-700 | 19 | 
23 | 600-700 | 0.37 | 
24 | 575-700 | 69 | 
25 | 650-750 | <100 | 
26 | 650-700 |  | 
HRPPD # | Operating Voltage Range (V/MCP) | Gain Range at ≥ 100V Bias on Photocathode (PC)
15 | 750-850 | 6.5E5 – 1.4E7
16 | 625-725 | 9.2E5 – 1.3E7
17 | 600-700 | 1.0E6 – 1.8E7
23 | 600-700 | 2.9E6 – 3.7E7
24 | 575-700 | 6.8E5 – 2.2E7
25 | 650-750 | 1.4E6 – 3.2E7
26 | 650-700 | 8.0E5 – 8.7E6
HRPPD # | Transit Time Spread (lowest recorded, in ps) | EIC Standard – TTS
15 | 59 | ≤60 ps, validated at BNL utilizing a femtosecond laser
16 | 59 | 
17 | 58 | 
23 | 58 | 
24 | 70 | 
25 | 72 | 
26 |  | 
HRPPD # | QE @ 365nm (Avg./Max.) | Std. Dev.
15 | 32.9% / 34.0% | 1.1%
16 | 34.0% / 35.3% | 0.6%
17 | 35.2% / 36.2% | 0.9%
23 | 26.8% / 28.8% | 1.7%
24 | 36.5% / 37.3% | 0.5%
25 | 26.3% / 28.8% | 1.2%
26 | 32.8% / 34.0% | 0.9%
by Mark Popecki (Incom)
Link to a May 2024 LAPPD / HRPPD Workshop talk
9
Charge Question 3

## Slide 10: DCR, Gain & Afterpulsing from Incom’s Testing
Newly developed ALD process allows for a high gain at a remarkably low bias voltage
HRPPD #23: dark rates at mid-106 gain are below 1 kHz/cm2 with afterpulsing on a ~1% level
Dark Count Rate
gain
afterpulsing
10
Charge Question 3

## Slide 11: Quantum Efficiency from Incom’s Internal Testing
Tiles sealed in vacuum tank #1 are well above 30% (we required 27% min & >30% average)
Uniformity better than expected
Issues with a fraction of tiles produced in tank #2 are understood
Window support structure
11
Charge Question 3

## Slide 12: HRPPD Evaluation by EIC Groups
Primary QA at JLab
More systematic active area scans (including picosecond timing) at BNL
Magnetic field resilience studies at BNL in April 2025 and possibly later at INFN
Staffed by BNL, Incom, USC, INFN Trieste / Genova
Main objective: gain and timing performance recovery in a “typical” pfRICH and hpDIRC B-field
Aging studies at JLab / BNL / INFN Trieste
Side by side Photek Auratek & Incom HRPPD comparison in Glasgow
12
Charge Question 2
Link to HRPPD evaluation procedure document

## Slide 13: EIC HRPPD Test Stands
Three fully functional QA / test stations available at JLab / BNL & INFN Trieste
See previous slide for a distribution of responsibilities
There is also an MCP-PMT / HRPPD test stand at Glasgow …
... and a clone of BNL old HRPPD setup at Yale
JLab
INFN Trieste
BNL
13
Charge Question 2

## Slide 14: Highlights from Initial QA Process at JLab
Several HRPPDs showed issues with falling off HV pins and mechanical screws
Those were promptly fixed by Incom
Solutions for a “final EIC HRPPD design” are being developed
Remove HV pins; embed screws into the ceramic anode body
Examples of failed HV pins and screws
A present fix by Incom
14

## Slide 15: Quantum Efficiency Measurements at BNL
HRPPD 16 (sealing tank #1)
HRPPD 23 (sealing tank #2)
2D surface scans @ 365 nm
QE(l) measurements with a monochromator
HRPPD 23 / HRPPD 16 QE ratio
Pretty uniform across the active area
Origin of non-uniformity is well understood
All HRPPDs sealed in tank #1 have peak QE way above 30%
15
Charge Question 3

## Slide 16: Active Area Evaluation at BNL
Data extracted from QE scans
~104 mm in both projections, as expected [anode pixellation: 32x 3.25mm]
The actual high PDE area with a sufficiently high gain will be somewhat smaller
X-coordinate
Y-coordinate
Overall impression: QE is high enough & Incom’s data confirmed
16

## Slide 17: Elmo 780 Femtosecond Laser System at BNL
HRPPD
Scope
We make use of a very low intensity 3rd harmonic @ 390 nm
17

## Slide 18: Single Photon Timing Resolution Using Elmo Laser
Laser beam focused on a single HRPPD pad center; intensity tuned down to >95% empty events
HRPPD signal used for triggering (5 mV effective threshold)
To increase data taking efficiency
Signal waveform data taken with a Tektronix MSO66B scope (50 GS/s, 8 GHz ABW)
Leading edge fits [10% .. 90%] performed offline; Dt = tHRPPD - tFastPD is a plotted quantity
s ~ 20 ps core
distribution width;
RMS < 75 ps
~15 ps
18
Charge Question 3
point shown on the left plot

## Slide 19: B-field Studies at BNL
All optical components received
Mechanical setup in being finalized
A type 18D72 2.2 Tesla dipole with a 6” gap
Dark “box” setup
19
Planned for the second half of April 2025
Link to a B-field study proposal
ePIC MARCO solenoid field map
Charge Questions 1 & 2

## Slide 20: Aging Studies
https://indico.bnl.gov/category/605/
20
HRPPD aging studies dark box at BNL
Monte-Carlo charged particle fluence estimate
Similar setups at JLab / BNL / INFN Trieste
Experts in LAPPD aging (UT Arlington) participating
A coordinated campaign in April - May 2025
HRPPD
window
Filter
wheel
Reference
photodiode
Focusing optics
on XYZ stages
Charge Questions 1 & 2

## Slide 21: HRPPD Summary & Our Present Assessment
Yields: all ten Kyocera anodes were functionalized into working HRPPDs at the end
s/n 15 .. 17 (success)
s/n 18 .. 22 (failure, after an attempt to use a second sealing tank)
s/n 23 .. 29 (success, after establishing a way to recycle the lower tile assemblies)
QE: higher than expected and remarkably uniform across the active area
21
Gain: stable operation at ~107 and beyond
We will not need more than 106 in the experiment
Uniformity is relatively poor, with gain going up by a factor of ~2 towards the active area edge
Timing: looks better than expected
pfRICH should be able to use single photons as a complementary t0 measurement
Dark rates: very small, even at mid-106 gain
Overall, everything looks very promising & one more iteration is expected

## Slide 22
Auratek MCP-PMTs by Photek
22

## Slide 23: MCP-PMT / HRPPD Evaluation at Glasgow
2” Photek Auratek stock configuration
A set of adapters to “HRPPD world”
Model: MAPMT253 16x16 pixel Multi-anode MCP-PMT
10 mm pores, peak QE up to ~35%, TTS ~40ps RMS
Shipment to Glasgow was expected by mid February 2025
Adapter boards available (see pictures above), as well as DRS4 electronics
23

## Slide 24: MCP-PMT / HRPPD Evaluation at Glasgow
Laser on XY-stages
HV divider
DT5742B
Planacon MCP-PMT
Adapter Card
Borrowed Planacon XP85112-S-BA MCP-PMT from GSI
This is the one which was thoroughly tested at Erlangen by A. Lehmann
Will be used as a reference tube
Adapter board for connecting to CAEN V1742 digitizer and MCX cables manufactured
One of the EIC HRPPDs will be shipped to Glasgow for a side-to-side comparison
24
Charge Question 2

## Slide 25: Comments from July 2023 PID Review
-> See slide 20
-> See slides 23-24
-> Such a solution would likely be suboptimal:
25
In case realized on the HRPPD anode level:
Cost gain is expected to be small; we would not be able to maintain a common pool of pfRICH & hpDIRC spares
In case realized on a front-end board level:
6.50mm pixels would translate into ~4mrad SPE angular resolution (compare to ~5mrad overall with 3.25mm pixels)
One would have to maintain two different FEB designs …
… and deal with fast signal reflections because of non-trivial proximity pad grouping on HRPPD anode air side
Charge Question 7

## Slide 26
Summary
26

## Slide 27: Summary & Outlook
Incom HRPPDs were adopted as ePIC pfRICH baseline photosensor
An EIC PED contract with Incom Inc. was successfully completed in 2024
Re-design and production of 5+2 EIC HRPPDs
Evaluation effort is ongoing in several institutions working on the EIC Project
The results are very promising so far
A moderate final design update and a test production are anticipated in 2025
Photek Auratek MCP-PMT evaluation as a baseline photosensor for ePIC hpDIRC is awaiting a tube delivery to Glasgow
27
-> HRPPDs are becoming our preferred solution for both pfRICH and hpDIRC
-> This is our backup solution

## Slide 30: Mechanical & electrical interface
Topic | Critically important
Add recesses for screws (studs) to the ceramic anode | YES
Use PEEK screws? Decide on epoxying rather than brazing? | YES
Eliminate HV pins & make use of compression interposers | YES
Add more grounding pads to Samtec interposers | 
Add more HV leads for a bottom MCP (reduce inductance) | YES
Improve photocathode HV lead design | NO

## Slide 31: Robustness, longevity
Topic | Critically important
Improve window sealing interface | 
Switch to sapphire windows | YES
Consider developing >72C temperature sealing alloy | ?
Eliminate / minimize anode tarnish | YES

## Slide 32: Performance improvements
Topic | Critically important
Increase geometric efficiency to ~80% | NO
Increase specs for QE to ~28-30% min | NO
Further minimization of transfer gaps | MAY BE
Minor minimization of material budget (window, anode) | NO
Implement sapphire window anti-reflective coating | 

## Slide 33: Minor modifications
Topic | Critically important
Improve inner and outer anode plate pixellation (gaps, etc) | NO
Add more HV ground spots on the rear anode side | 
Consider a better electrical isolation of the indium seal | 

## Slide 34: Other options
Topic | Critically important
Try a photocathode QE shift towards an optical range | 
Add a high resistivity ALD coating to side walls (and anode?) | 
Increase anode segmentation (to 40 x 40 pads or so) | 
Add window edge reflective coating | 
Make use of 6 mm MCPs and / or increase bias angle to >130 | MAY BE
Implement anti-reflective coating on a window-PC boundary | MAY BE

## Slide 35: Steps towards converging on a single photosensor solution
(1) Complete evaluation of the first batch of EIC HRPPDs
Systematic scans: QE, PDE, gain, DCR, timing, …
Aging, B-field resilience, after-pulsing, high occupancy events
(2) Proceed with a final PED iteration with Incom
Repeat a systematic evaluation
(3) Perform a selective side to side comparison of HRPPD and Photek MCP-PMT
If Photek MCP-PMTs show no substantial performance advantage, take a note
(4) Evaluate functionality of an existing HGCROC3 backplane prototype with an HRPPD
(5) Assuming it provides a reasonable performance, build 5-7 full sets

## Slide 36: Steps towards converging on a single photosensor solution
(6) Install a pfRICH prototype with 4-5 HRPPDs on a Cosmic Ray Telescope at SBU
See this talk
(7) At some point move the on-board electronics onto the hpDIRC CRT prototype
And verify the performance on real life multi-hit muon events (cross talk, ringing, etc)
(8) Wait for FCFD evaluation boards
Perform intermediate bench top testing with final EIC HRPPDs (and Photek MCP-PMTs?)
(9) Build a full FCFD backplane
Repeat a fraction of HRPPD performance studies
(10) Perform beam test(s) for both pfRICH and hpDIRC