## Summary
|  |  | Geometry | [6 empty] | Electronics |  |  | Detector Cables | [9 empty] | Feed Cables | [9 empty] | Services Detector Gas | [6 empty] | Feed Gas | [6 empty] | Detector Cooling Tubes | [5 empty] | Feed Cooling Tubes | [5 empty] | Cooling Fluid | [5 empty] | Last Updated |  |  |  |  | Notes
|  |  | ID | OD | Length | z location from IP | BP Clearance | # of Segments | Radiation lengths | VTRX+ QTY | FEB QTY | RDO QTY | LV OD | LV QTY | LV Cable Type | HV OD | HV QTY | HV Cable Type | Signal OD | Signal QTY | Signal Cable Type | Length | LV OD | LV QTY | LV Cable Type | HV OD | HV QTY | HV Cable Type | Signal OD | Signal QTY | Signal Cable Type | Length | Dry Air ID | Dry Air OD | Dry Air QTY | Gas ID | Gas OD | Gas QTY | Length | Dry Air ID | Dry Air OD | Dry Air QTY | Gas ID | Gas OD | Gas QTY | Length | ID | OD | Insulation OD | QTY | Material | Length | ID | OD | Insulation OD | QTY | Material | Length | Fluid | Heat Disipation | Temp | Variation | Flow Rate | Pressure | First Name | Last Name | Date | Time | email
|  | Unit | mm | mm | mm | mm | mm |  | % |  |  |  | mm |  |  | mm |  |  | mm |  |  | m | mm |  |  | mm |  |  | mm |  |  | m | mm | mm |  | mm | mm |  | m | mm | mm |  | mm | mm |  | m | mm | mm | mm |  |  | m | mm | mm | mm |  |  | m |  | w | c | c | l/m | Pa
SVT | Inner
Outer
Disks, Forward
Disks, Backward
MPGD | Inner |  | 55 | 60.5 | 2470 |  |  | 32 | 1% | 128 | 128 | 0 | 4 to 5.5 | 128 | Alpha Wire (3464C, ... depending on length to DC/DC from 0.3 to 1 m) | 3.2 | 96 | HTC-50-1-1 (CEH50) | ~4 thick | 512 | KEL micro coaxial cable USL20-30SS-015 | 0.1 to 0.5 | 6 | 128 | Alpha Wire 2424C | 3.2 | 96 | HTC-50-1-1 (CEH50) | 2 | 128 | Optical fiber | to DAM (or via patch panel ?) | [10 empty] | 2 | 4 | 16 IN / 16 OUT | [5 empty] | copper ? | [7 empty] | H2O | 1100 | 18 | +/- 2 | 10 |  | Alain | Delbart | 2025-03-10 |  | alain.delbart@cea.fr | 2 sections of LV cables (from FEB to DC/DC & from DC/DC to LV PS). Lengths depending on location of patch panels. Thermal simulations to be done to define the cooling requirements and design the components of the system
| Outer |  |  |  | 3600 |  |  | 24 MPGD modules |  | 336 | 336 (24X14) |  | 4.8 | 336 | Alpha Wire (3464C, ... depending on length to DC/DC from 0.3 to 1 m) | 3.2 | 48 | HTC-50-1-1 (CEH50) | 2 | 336 | Optical fiber | 0.1 to 4 | 6 | 336 | Alpha Wire 2424C | 3.2 | 48 | HTC-50-1-1 (CEH50) | 2 | 336 | Optical fiber |  |  |  |  | 3 | 4 | 24 IN / 24 OUT |  |  |  |  | 3 | 4 | 24 IN / 24 OUT |  |  |  |  | 96 (4 per sector | Stainless steel/ | 4 | [6 empty] | H2O | 2800 | 15 | +/- 2 | TBD | TBD | Seungjoon | Lee | 2025-03-14 |  | sjlee@jlab.org
| Disks, Forward |  |  |  | 150 | 1650 | 5 | 8 quadrants | 1% | 48 | 48 |  | 4 to 5.5 | 48 | Alpha Wire (3464C, ... depending on length to DC/DC from 0.3 to 1 m) | 3.2 | 32 | HTC-50-1-1 (CEH50) | ~4 thick | 48 | KEL micro coaxial cable USL20-30SS-015 | 0.1 to 0.5 | 6 | 48 | Alpha Wire 2424C | 3.2 | 48 | HTC-50-1-1 (CEH50) | 2 | 48 | Optical fiber | to DAM (or via patch panel ?) |  |  |  | 3 | 4 | 8 IN / 8 OUT |  |  |  |  | 3 | 4 | 8 IN / 8 OUT |  |  |  |  | 8 | copper ? | [7 empty] | H2O | 410 | 15 | +/- 2 | 10 |  | Annalisa | D'Angelo | 2025-03-17 |  | annalisa.dangelo@roma2.infn.it
| Disks, Backward |  |  |  | 150 | -1225 | 5 | 8 quadrants | 1% | 48 | 48 |  | 4 to 5.5 | 48 | Alpha Wire (3464C, ... depending on length to DC/DC from 0.3 to 1 m) | 3.2 | 32 | HTC-50-1-1 (CEH50) | ~4 thick | 48 | KEL micro coaxial cable USL20-30SS-015 | 0.1 to 0.5 | 6 | 48 | Alpha Wire 2424C | 3.2 | 48 | HTC-50-1-1 (CEH50) | 2 | 48 | Optical fiber | to DAM (or via patch panel ?) |  |  |  | 3 | 4 | 8 IN / 8 OUT |  |  |  |  | 3 | 4 | 8 IN / 8 OUT |  |  |  |  | 8 | copper ? | [7 empty] | H2O | 410 | 15 | +/- 2 | 10 |  | Stefano | Gramigna | 2025-03-17 |  | stefano.gramigna@roma2.infn.it
TOF | Barrel |  | 625 | 700 | 1350 | -1175 | 1171.5 | 288 | 3% | 288 |  | 288 | 6.3 | 288 | Alpha-PN: 2424C SL005 | 1.5 | 288 |  |  | 1152 | fiber |  | 20 | 24 | 8 AWG | 14 | 18 | [31 empty] | H2O |  | 18 | +/-2 |  |  | Tonko |  | 2025-03-10 |  |  | zhangbu 03/20/25
| Disk, Forward |  | 100 | 650 | 70 | 1850 | 100 | 2 | 5% | 212 |  | 212 | 9 | 212 | 18 AWG TC | 2.4 | 212 |  |  | 212 | fiber |  | 10 | 44 | 14 AWG TC | 2.4 | 212 | [31 empty] | H2O |  | 18 | +/-2 |  |  | Tonko |  | 2025-03-10
RICH | Forward
Backward
DIRC | DIRC
EMCal | Barrel PB/SiFI | [8 empty] | N/A | 96 | 12 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 2 | 192 | 30 V/1A | 2 | 192 | 50/1A | 2 | 96 | Optical fiber | 100m | [36 empty] | Zisis.Papandreou@uregina.ca | Is this table with services still relevant? We were updating our numbers here: https://brookhavenlab.sharepoint.com/:x:/s/EICPublicSharingDocs/EdH38QZE9HpJrl039jn2-q4BbPvrMv7dTFiLV8--atclKw?rtime=NZnXV7JS3Ug
| Barrel AstroPix | [67 empty] | Zisis.Papandreou@uregina.ca
| Forward |  | 250.95 | 4115 | 275.2 | 3350 | 31.475 |  |  |  | 619 | [35 empty] | 4.826 | 6.35 |  |  | Copper | [13 empty] | Joshua | Harvey | 2025-03-10 |  | jharvey@bnl.gov
Backward
Magnet | Magnet
HCal | Barrel |  | 3645/3885 | 5380 | 6392.5 | 0 | N/A | 32 |  | N/A | 120 | 15 | 1mm | 64x120 | LED calib | 1mm | 64x120 | microcoax | 1mm | 64x120 | microcoax | 2m | 2 | 240 | 30 V/1A | 2 | 240 | 70V/1A | 3mm | 120 | Optical fiber | 100m | [20 empty] | 5mm | 6mm |  | 5W/FEB | Air | 20 | [6 empty] | Norbert | Novitzky | 2025-03-14 |  | novitzkyn@ornl.gov
Forward
Backward
Electronics | VTRX+
FEB
RDO
FarBackward | Lumi PairSpec
Lumi HR Calo
Low Q2-Tagger
FarForward | B0-Tracker
B0-Calo
| OffMomentum |  |  |  |  | 26000 | inside |  | < 5% per layer | [55 empty] | Alex | Jentsch | 2025-03-20 |  | ajentsch@bnl.gov
| Roman Pots |  |  |  |  | 30000 | inside |  | < 5% per layer | [55 empty] | Alex | Jentsch | 2025-03-20 |  | ajentsch@bnl.gov
ZDC
Polarimetery | H-Jet
pC IP4
pC IP6
Compton IP6 PhotonDet
Compton IP6 ElectronDet
Compton RCS PhotonDet
Compton RCS ElectronDet

## Example
These are emaples how to provide the cooling invormation
ToF Barrel | ToF Forward
Indicate the heat sources at the detector and how they are distributed over the detector volume
are the heat sources at the rim of the detector or over the entire volume or a mix
What is the heat source: ASIC, Stave, FEB, RDO, or ….
Power per each heat source
What simulations of the heat distribution have been done ?
What assumption went into the simulations?
What assumtions are made for the cooling system
Coolant
Flow
Are there limitations on the material budget the cooling can have
What are the requirements for the cooling system
Required Temperature stability
operation temperature of the detector (range if possible)
any limits on allowed vibration of detector
Are there other heat sources not directly I nthe detector volume which need cooling
Example: DC-to-DC converters, patch-panels, ….....
Provide the same information and requirements as for the cooling of the same subsystem
Double-sided staves
ASIC
Strip-Sensor

## SVT Inner
Layout
Heat sources
| Sensor element | Nominal power (mW) | Max power (mW)
| RSU | 152.3 | 200.3
| LEC | 1000 | 1000
| Full segment | 2527.6 | 3103.6
Sensor version
| L0 | 7582.8 | 9310.8
| L1 | 10110.4 | 12414.4
| L2 | 12638 | 15518
Detector element
| Half-Layer L0 | 15165.6 | 18621.6
| Half-Layer L1 | 20220.8 | 24828.8
| Half-Layer L2 | 50552 | 62072
| Layer L0 | 30331.2 | 37243.2
| Layer L1 | 40441.6 | 49657.6
| Layer L2 | 101104 | 124144
| TOTAL (mW) | 171876.8 | 211044.8
| TOTAL (W) | 171.9 | 211
IB
MOSAIX power domains
Supply purpose | Current (mA) | Voltage (V) | Power (W)
GSVDD | 227 | 1.2 | 0.272 | Services - include as 1V2 power (requiring BPOL2V5)
GAVDD | 540 | 1.2 | 0.648 | Global analogue - include as 1V2 power (requiring BPOL2V5)
GDVDD | 1482 | 1.2 | 1.778 | Global digital - include as 1V2 power (requiring BPOL2V5)
TXVDD | 300 | 1.8 | 0.54 | Serialisers - include as 1V2 power (requiring BPOL2V5)
PSUB | -1 | -1.2 | 0.001 | Assume 1mA current useage at 1V2 power (BPOL2V5, however different regulator will be used for this)
Segment Interface Board (SIB) |  |  | -> 1SIB for every MOSAIX segment
Contains 2 readout (4Tx) VTRx+ |  |  |  | @1V2 | @2V5
Power consumed on board |  | 0.395 | W | 0.12 | 0.275
Add 2% resistive losses |  | 0.403 | W
Total SVT sgements |  | 68 |  | L0 = 4*3 | L1 = 4*4 | L2 = 8*5
Total power consumed in SIBs |  | 27.397 | W
Segment Control Board (SCB) |  |  | -> 1SCB for every SIB (MOSAIX segment) - Assuming SCB contains BPOL12 and BPOL2V5
Contains 1 SC (1Tx+1Rx) VTRx+, 1 lpGBT, bpols |  |  |  | @1V2 | @2V5
Power consumed by above |  | 0.903 | W | 0.662 | 0.241
Power to transfer to SIB |  | 0.403 | W | 0.122 | 0.281
Power for MOSAIX |  | 3.24 | W | 3.24
** Also need to convert 10V-2V5 and 1V2 **
2V5 power needed |  | 0.521 | W | power into control board |  |  | 0.744 | W | 1 | BPOL12 needed per board
1V2 power needed |  | 4.025 | W | power into control board |  |  | 7.187 | W | 2 | BPOL2V5 needed per board
Convertion power losses |  | 3.386 | W | Power needed (3% res losses) |  |  | 8.17 | W | per board (from DPB)
Power consumed on board |  | 4.289 | W |  |  |  | 0.817 | A @ 10V | (10.3V, to drop 3% over 10m @ 0.76A of copper wire require 1.14 mm diameter, 17 AWG) |  |  | or 1.42 mm diamter of aluminiuim wire, 15 AWG
Add 2% resistive losses |  | 4.374 | W | [5 empty] | https://www.omnicalculator.com/physics/wire-size
Total power consumed in SCBs |  | 297.462 | W
Detector Power Board (DPB) |  |  | -> 1DPB for (up to) |  | 14 | Segment Service Modules, SSM (= 1SCB + 1SIB) |  |  |  |  | ITS3's standard seem to be 1DPB to 12SSMs, AKA 1/2 ITS3 barrel (3*L0 seg + 4*L1 seg + 5*L2 seg)
Contains 1 SC (1Tx+1Rx) VTRx+, 1 lpGBT, bpols |  |  |  | @1V2 | @2V5 | @10V |  | Maybe consider separating L0+1 from L2 for ePIC, so 'innerDPB' = 14 SSMs (half barrel for L0+1), and 'outerDPB' = 10 SSMs (quarter barrel for L2)
Power consumed by above |  | 0.903 | W | 0.662 | 0.241
Power for SCBs |  | 114.373 | W |  |  | 114.373 |  | Assuming DPB sends 10V only to the SCBs
** Also need to convert 48V-10V, 2V5 and 1V2 **
10V power needed |  | 114.373 | W | power into control board |  |  | 127.081 | W | 2 | BPOL48 needed per board
2V5 power needed |  | 0.241 | W | power into control board |  |  | 0.382 | W | 1 | BPOL12 needed per board
1V2 power needed |  | 0.662 | W | power into control board |  |  | 1.314 | W | 1 | BPOL2V5 needed per board
Convertion power losses |  | 13.501 | W | PSU supply needs (5% res losses) |  |  | 135.216 | W | per board
Power consumed on board |  | 14.404 | W |  |  |  | 2.817 | A @ 48V | (54V, to drop 10% over 100m @ 2.82 A of copper wire require 1.66 mm diameter, 13 AWG) |  |  | or 2.10 mm diamter of aluminiuim wire, 11 AWG
Add 2% resistive losses |  | 14.692 | W
Number of DPBs needed |  | 6 |  | L0+L1 | 2 | L2 | 4
Total power consumed in DPBs |  | 88.155 | W | Assuming the worst case of 14 SSMs/DPB (4 out of 6 should only be 10SSMs)
Total SVT RDO power (to cool) |  |  |  |  | Total SVT RDO power (to supply, @54V/PSU channel)
1045.059 | W |  |  |  |  | 1871.414 | W
Air cooling (to be designed)
|  | ITS3 | SVT
| Cooling system | Air based cooling
| Air flow (m/s) - L0 | 8
| Air flow (m/s) - L1 | 8
| Air flow (m/s) - L2 | 8
| Required temperature stability | ??
| Operation temperature of the detector | The sensor’s operating temperature must be <30°C, and the maximum temperature gradient in the pixel matrix region must be <5°C to guarantee a stable operation.
Maximum temperature of components on the half-barrel
| Limits on vibration | Measured 0.4 μm at 8 m/s airflow velocity

## SVT Outer
Layout | This still assumes curved stave surfaces, to be confirmed, but thermal management for square staves is similar
Red are high power density devices (LEC of the LAS and AncASIC)
Yellow are low power density (RSUs of LAS)
White are areas with no power generations
Layout on bottom is similar with module positions complementary to top module positions
Heat sources
| 5 RSU LAS (L4) |  | 6 RSU LAS (L3) |  |  | See Iain Sedgwick https://agenda.infn.it/event/43344/contributions/253346/
| nominal P [mW] | max P [mW] | nominal P [mW] | max P [mW] |  | See Georg Viehhauser https://agenda.infn.it/event/43344/contributions/253347/
Sensor | 1267 | 1723 | 1422 | 1943
AncASIC | 443.4 (35%) | 775.3 (45%) | 497.7 (35%) | 874.4 (45%)
FPC | 253.3 (20%) | 516.8 (30%) | 284.4 (20%) | 582.9 (30%)
L4 stave
5RSU LAS per stave | 16
LAS power | 27.6 | W
AncASIC power | 12.4 | W
FPC power | 8.3 | W
Total power per stave | 48.2 | W
Number of staves | 70
Total power L4 | 3377 | W
L3 stave
6RSU LAS per stave | 8
LAS power | 15.5 | W
LDO power | 7 | W
FPC power | 4.7 | W
Total power per stave | 22.5 | W
Number of staves | 46
Total power L3 | 1037 | W
Air cooling (to be designed)
Flow [cfm] | assume ΔT=15°C, cp,v = 0.00129 J/Kcm3
| per Stave | total
L3 | 2.4 | 110.5 | flow numbers are indicative and need confirmation
L4 | 5.2 | 365.9
required temperature stability |  | ±1°C | preliminary, contingent on cooling design and stave thermo-mechanical studies
operation temperature of the detector |  | coolant input temperature to be defined (maximum room temperature, but slightly lower temperature would be beneficial, maybe 15°C)
coolant output temperature = input temperature + 15°C
Maximum temperature of omponents on the staves
| LEC | RSU | AncASIC
Maximum temperature [°C] | 65 | 40 | 80
Maximum temperature variation [°C] | 10 | 1 over 3 mm, 10 over full length | 10
Heat sources not connected to the staves (i.e. not in the active areas):
Components in chain (same as for the disks):
dissipation [W]
FPC interface board (FIB) | 0.201 | 1 FIB per FPC group of up to 4 EIC LAS
FIB Control Board | 21.401 | 1 FIB control board for up to 64 FIBs
FPC Control Board | 1.73 | 1 FPC Control Board for up to 16 FPC groups
Components in OB:
Section | Structure | structures/section | FPC groups/structure | e-side FIBs | h-side FIBs | e-side FIB-CBs | h-side FIB-CBs | e-side FPC-CBs | h-side FPC-CBs
OB - L3 | stave | 46 | 2 | 46 | 46 | 1 | 1 | 4 | 4
OB - L4 | stave | 70 | 4 | 140 | 140 | 3 | 3 | 12 | 12
| Count | Dissipation [W]
Total FIBs | 372 | 74.772
Total FIB-CBs | 8 | 171.208
Total FPC-CBs | 32 | 55.36
Total |  | 301.34 | Assume liquid cooled

## SVT Disks
Layout
LEC & AncASIC: high power density
RSU & FPC: low power density
Relevant presentations:
https://agenda.infn.it/event/43344/contributions/253348/attachments/130492/194194/ePIC_Cooling_1_21_25.pdf
https://agenda.infn.it/event/43344/contributions/253353/attachments/130605/194405/ePIC_SVTDiscs_1_23_25.pdf
Heat sources
| 5 RSU LAS |  | 6 RSU LAS
| nominal Power [mW] | max Power [mW] | nominal Power [mW] | max Power [mW]
Sensor | 1267 | 1723 | 1422 | 1943
AncASIC | 443 (35%) | 775 (45%) | 498 (35%) | 874 (45%)
FPC | 253 (20%) | 517 (30%) | 284 (20%) | 583 (30%)
Numbers of 5 & 6 RSU LAS per disk present one of several possible layouts. The total number of sensors is not expected to vary more than +/- 10% if an alternative layout choice is used.
Power (W)
|  |  | Nominal Power |  |  |  | Max Power
Disc | # 5-RSU LAS | # 6-RSU LAS | LAS | AncASIC | FPC | Total | LAS | AncASIC | FPC | Total
ED0 | 102 |  | 129.21 | 45.23 | 25.84 | 200.27 | 175.72 | 79.08 | 35.14 | 289.95
ED1 |  | 258 | 366.88 | 128.41 | 73.38 | 568.66 | 501.29 | 225.6 | 100.26 | 827.15
ED2 |  | 266 | 378.25 | 132.39 | 75.65 | 586.29 | 516.84 | 232.59 | 103.37 | 852.8
ED3 |  | 264 | 375.41 | 131.39 | 75.08 | 581.88 | 512.95 | 230.84 | 102.59 | 846.38
ED4 |  | 266 | 378.25 | 132.39 | 75.65 | 586.29 | 516.84 | 232.59 | 103.37 | 852.8
HD0 | 102 |  | 129.21 | 45.23 | 25.84 | 200.27 | 175.72 | 79.08 | 35.14 | 289.95
HD1 |  | 258 | 366.88 | 128.41 | 73.38 | 568.66 | 501.29 | 225.6 | 100.26 | 827.15
HD2 |  | 266 | 378.25 | 132.39 | 75.65 | 586.29 | 516.84 | 232.59 | 103.37 | 852.8
HD3 |  | 266 | 378.25 | 132.39 | 75.65 | 586.29 | 516.84 | 232.59 | 103.37 | 852.8
HD4 |  | 266 | 378.25 | 132.39 | 75.65 | 586.29 | 516.84 | 232.59 | 103.37 | 852.8
[6 empty] | 5051.2 |  |  |  | 7344.56
Air Cooling (to be designed)
c_V = 0.00123 J/(cm3*K) |  |  | ΔT numbers are based on 15℃ air inlet temperature
specific heat @ constant V |  |  | Since LEC and RSU are connected, unclear as to whether they can support different max temperatures. So, LAS max temperature set to be RSU max
Air flow calculations below are based on the LAS, AncASIC, and FPC ΔT. FPC ΔT is assumed to be the same as the AncASIC
Other considerations: See SVT Outer Tab
Calorimetric calculation only. Numbers to be verified with ANSYS simulations and thermal data measurements
| LEC | RSU | AncASIC
Max Temp (℃) | 65 | 40 | 80
Conservative Estimate Air Flow (cfm)
LAS ΔT = 15℃, AncASIC ΔT = 55℃
|  | Nominal Power |  |  | Max Power
Discs | # LAS | LAS | AncASIC | FPC | LAS | AncASIC | FPC
ED0 | 102 | 14.8 | 1.41 | 0.81 | 20.12 | 2.47 | 1.65
ED1 | 258 | 42.01 | 4.01 | 2.29 | 57.4 | 7.05 | 4.7
ED2 | 266 | 43.31 | 4.13 | 2.36 | 59.18 | 7.26 | 4.84
ED3 | 264 | 42.99 | 4.1 | 2.34 | 58.74 | 7.21 | 4.81
ED4 | 266 | 43.31 | 4.13 | 2.36 | 59.18 | 7.26 | 4.84
HD0 | 102 | 14.8 | 1.41 | 0.81 | 20.12 | 2.47 | 1.65
HD1 | 258 | 42.01 | 4.01 | 2.29 | 57.4 | 7.05 | 4.7
HD2 | 266 | 43.31 | 4.13 | 2.36 | 59.18 | 7.26 | 4.84
HD3 | 266 | 43.31 | 4.13 | 2.36 | 59.18 | 7.26 | 4.84
HD4 | 266 | 43.31 | 4.13 | 2.36 | 59.18 | 7.26 | 4.84
Total (per) |  | 373.18 | 35.62 | 20.36 | 509.71 | 62.56 | 41.7
Total (all) |  |  | 429.16 |  |  | 613.97
Heat sources not connected to the disks (i.e. not in the active areas)
Components in chain (same as for OB):
dissipation [W]
FPC interface board (FIB) | 0.201 | 1 FIB per FPC group of up to 4 EIC LAS
FIB Control Board | 21.401 | 1 FIB control board for up to 64 FIBs
FPC Control Board | 1.73 | 1 FPC Control Board for up to 16 FPC groups
Components in disks:
Section | Structure | structures/section | FPC groups/structure | e-side FIBs | h-side FIBs | e-side FIB-CBs | h-side FIB-CBs | e-side FPC-CBs | h-side FPC-CBs
ED0 | half-disk | 2 | 13 | 26 |  | 1 |  | 3
ED1 | half-disk | 2 | 43 | 86 |  | 2 |  | 8
ED2 | half-disk | 2 | 43 | 86 |  | 2 |  | 8
ED3 | half-disk | 2 | 43 | 86 |  | 2 |  | 8
ED4 | half-disk | 2 | 43 | 86 |  | 2 |  | 8
HD0 | half-disk | 2 | 13 |  | 26 |  | 1 |  | 3
HD1 | half-disk | 2 | 43 |  | 86 |  | 2 |  | 8
HD2 | half-disk | 2 | 43 |  | 86 |  | 2 |  | 8
HD3 | half-disk | 2 | 43 |  | 86 |  | 2 |  | 8
HD4 | half-disk | 2 | 43 |  | 86 |  | 2 |  | 8
| Count | Dissipation [W]
Total FIBs | 740 | 148.74
Total FIB-CBs | 18 | 385.218
Total FPC-CBs | 70 | 121.1
Total |  | 655.058 | Assume liquid cooled

## MPGD Disks
MPGD Endcap Disks (MPGD ECT)
Total of 16 quadrants (4 quadrants X 4 disks)
6 FEB/quadrant = 96 FEB in total |  |  |  | Heat sources are ONLY the Front End Boards (FEBs) distributed all around the outer ring of the discs
Heat sources on FEBs: ASIC, LDO, LpGBT, VTRx
Extimated heat |  | ~6.8 W / FEB (8.5 W with 25% safety margin) |  |  | given current figures for electronics' power output
Heat per quadrant = 50 W |  |  | Total Heat = 800 W
Cooling requirement similar to MPGD Outer and MPGD Inner:
Cooling System Assumtions:
| Coolant | Demineralized Water+anti-corrosive + 5% ethanol (anti-weed) at 18°C ?
| Flow | ~few l/minute
Cooling System Requirements:
| Required Temperature stability : |  |  |  | not critical for FEBs (+/- 2 °C)
| operation temperature of the detector (range if possible) |  |  |  | 15-25 °C ? to be evaluated to guarantee Gas temperature stability inside the disks volume (+/- 1 °C ?), preventig condensation
any limits on allowed vibration of detector
Are there other heat sources not directly in the detector volume which need cooling?
DC-to-DC converters !
cooling requirements not yet evaluated

## MPGD Outer
| Barrel Outer Tracker (BOT) |  |  |  | Vew from IP |  | Service line from sectional vew (between DIRC and CF tube)
Radial Envelope: R = 735 mm to 760 mm
Two MPGD modules per sector, Total 12 sectors
RED area will be service lines
14 FEB/module |  | 28 FEBs/Sector | 336 FEBs total | 12 Service Sectors (1 sector serving 2 MPGD modules)
~6.8 W / FEB (8.5 W with 25% safety margin) |  |  |  | 238 Watt per sector
Heat sources on FEBs: ASIC, LDO, LpGBT, VTRx
Cooling requirement will be similar to Inner MPGD.
Weight estimation is ongoing
FEB | 300 g
MPGD

## MPGD Inner
CyMBaL Barrel
Z- Outer tiles
One of the 16 outer tiles tiles
One of the 16 inner tiles tiles
Z+ Outer tiles
Heat sources are ONLY the Front End Boards (FEBs) distributed all around the 2 sides of the circumference of the Barrel (2x64 FEBs)
~6.8 W / FEB (8.5 W with 25% safety margin) ---> ~2 x 5.5 kW on each side of the barrel
Heat sources on FEBs: ASIC, LDO, LpGBT, VTRx
Simulation of ther heat cooling not yet done
Could be derived from the one of the T2K/ND280 HA-TPC readout FEBs which have very similar heat dissipation (but to be more compact for CyMBaL)
T2K/ND280 HA-TPC readout FEBs : 7W/FEB, cooled at ~25°C +/- 0.2 °C by 3 l/minute 18°C water flow in copper tubes on aluminum cooling plate (negative pressure system)
What assumtions are made for the cooling system
| Coolant | Water+anti-corrosive at 18°C ?
| Flow | ~few l/minute
Are there limitations on the material budget the cooling can have
What are the requirements for the cooling system
| Required Temperature stability : |  | not critical for FEBs (+/- 2 °C) but
| operation temperature of the detector (range if possible) |  |  | 15-25 °C ? to be evaluated regarding Gas temperature stability inside the tiles volume (+/- 1 °C ?)
any limits on allowed vibration of detector
Are there other heat sources not directly I nthe detector volume which need cooling
DC-to-DC converters !
cooling requirements not yet evaluated

## RICH Forward
Describe the Detector layout: figures of the layout detailed and large area are critical
dRICH
The heat sources in the dRICH are primarily located inside the photodetector boxes
We distinguish two operation modes: physics, annealing
During physics operation, these are the main elements of the detector that contain heat sources
RDO board
FEB board
MasterLogic board
The heat sources within the RDO board are
the voltage regulators
the main FPGA (Xilinx Artix Ultrascale+)
the auxiliary PFGA (Microchip Polarfire)
The heat sources withing the FEB board are
the voltage regulators
the ALCOR ASIC chip
The heat sources withing the MasterLogic board are
the SiPM bias voltage regulators
the elements of the control logic
The total power consumption for each dRICH photodetector box is 1.8 kW and it is split as follows
720 W for the FEB analog voltage
620 W for the FEB and RDO digital core voltage
840 W for the FEB and RDO digital IO voltage
170 W for the SiPM bias voltage regulators
25 W for the ancillary power of the MasterLogic boards
It can be assumed that the heat sources are evenly distributed within the dRICH photodetector boxes
During annealing operation, these are the main elements of the detector that contain heat sources
PDU head
FEB board
MasterLogic board
The heat sources within the PDU head are
the SiPM sensors
the heat exchanger
the pipes
The heat sources within the FEB are
annealing diodes
The heat sources withing the MasterLogic board are
the elements of the control logic
The total power consumption for each dRICH photodetector box is 3.5 kW and it is split as follows
3.3 kW on the SiPM sensors
200 W on the annealing diodes
The heat sources are not evenly distributed within the dRICH photodetector boxes.
Annealing is done on one PDU every 8, therefore there might be strings of PDUs being annealed at the same time.
Nonetheless, one has to consider the thermal power being brough inside the photodetector boxes by the heating fluid.
All the heat exchangers of the PDU head will be at a static temperature of T = 100 C, as well as the pipes for the fluid.
The dissipated heat has been measured on real components (FPGA/regulators) or calculated via electronics simulation tools (ASICs)
Indicate the heat sources at the detector and how they are distributed over the detector volume
are the heat sources at the rim of the detector or over the entire volume or a mix
What is the heat source: ASIC, Stave, FEB, RDO, or ….
Power per each heat source
What simulations of the heat distribution have been done ?
What assumption went into the simulations?
What assumtions are made for the cooling system
Coolant
Flow
Are there limitations on the material budget the cooling can have
What are the requirements for the cooling system
Required Temperature stability
operation temperature of the detector (range if possible)
any limits on allowed vibration of detector
Are there other heat sources not directly I nthe detector volume which need cooling
Example: DC-to-DC converters, patch-panels, ….....
Provide the same information and requirements as for the cooling of the same subsystem

## RICH Backward
pfRICH LV and cooling system layout and estimates are based on a configuration with four hypothetical 256ch EICROC ASICs.
Respective slides have not been changed since July 2023 PID review. Will be reworked once more information on FCFD becomes available.
Cabling and other services are summarized in | https://brookhavenlab.sharepoint.com/:x:/r/sites/EICPublicSharingDocs/_layouts/15/Doc.aspx?sourcedoc=%7B06F1F7D1-F444-497A-AE5D-37F639F6FAAE%7D&file=ePIC.Interface%20Control%20Document%20for%20Services.052023.xlsx&action=default&mobileredirect=true&DefaultItemOpen=1
See also slides 6-11 in a TIC meeting summary of pfRICH leadership meeting with the EIC Project on 10/16/2024: |  | https://indico.bnl.gov/event/25169/contributions/97905/attachments/58137/99879/ayk-2024-10-28-tic-meeting-pfrich.pdf

## DIRC
The hpDIRC system largely follows the effort of the pfRICH system in terms of readout and cooling (modulo HRPPD vs Photek MCP-PMT decision).
Plan to work on this worksheet in early April.

## ECal Barrel PbScifi + AstroPix
Detector description
End of sector box ESB (located at the end of the sectors)
Electron-going side: conflict wit DIRC lens/expansion box requires BIC ESB to pull out longitudinally by 1.5 cm and then pull out radially for servicing - 1.5 cm confirmed with the Project
Hadron-going side will pull our longitudinally for servicing
Cooling comments and assumptions
UManitoba is currently looking at: a) thermal conduction of Astropix through the lead SciFi and b) ESB cooling for SiPM stability and ETC and CALOROC heat removal.
We may use a thermal pad behind the SiPM layer or a metal-core PCB for heat removal
Forced air in the ESB? Still considering this for pushing out heat from AstroPix trays if needed (i.e. if heat conduction through the PbScFi is not sufficient)
Regina can build the cooling lines. Assuming a cooling snake that runs by all the SiPMs, we could fit a 1-cm-diameter copper line between SiPM rows.
NASA: roughly 1 W per ETC card (the FPGA basically), so 4 W per box; Connector at the end of each the Astropix tray, then we can use cables to go to the ETC cards which could sit in the volume under the CALOROC (brown volume in above figure) and receive heat removal from the pad or metal-core PCB
Regina-design SiPM summing and amplification, 8 W total per ESB box (upper limit) but some of the electronics processing will be undertaken by CALOROC and there for the 8W should drop
HGCROC chip: 4W each and operate at 40 C without cooling (from interface document)
Coolant |  | Chilled water at 5C on chiller, 7 C on SiPMs/ N2 or dry air to push out condensation
Flow |  | Not known yet
Material Budget |  | Possibly 1-cm-diameter copper inside the ESB could by PVC lines outside' not verified yet
Cooling sources
SiPMs | Their summing and amplification boards (OPamps)
ETCs | FPGA
CALOROC | ASIC
AstroPix | ASIC | 2 mW/cm^2
AstroLinx | hybrid PCB | power regulators
Cooling requirements
SiPMs | Temperature stability similar to GlueX-BCAL: +/-2 C with chilled water; gain is stabilized using a custom circuit that adjusts the bias voltage based on the measured temperature
ETCs | heat removal; keep below 40 C?
CALOROC | heat removal; keep below 40C?
AstroPix | heat removal; maintain room temperature (20 +/- 3 C)
AstroLinx | heat removal; maintain room temperature (20 +/- 3 C)
Very preliminary estimate of power per ESB

## ECal Forward
Detector Layout | [6 empty] | Cooling System Description
Heat Sources |  |  | North Half |  | South Half
Component | Heat Generation | Units | Quantity | Total Heat Generation | Quantity | Total Heat Generation | Units | Notes
FEB (Full Board) | 4.5 | W | 277 | 1246.5 | 274 | 1233 | W | Applied on face in contact with thermal pad on heat sink.
FEB (Half Board) | 2.25 | W | 44 | 99 | 24 | 54 | W | Is it safe to assume heat generated by half FEB is half of full FEB?
SiPM Board | 0.17 | W | 2392 | 406.64 | 2288 | 388.96 | W | Total heat applied by 16 SiPMs on each SiPM board. SiPMs are main source of heat. Remember: there are 8 SiPM boards per full FEB.
Status of SImulations?
Current work in progress
Temperature Requirements
Maintain FEB temperature 2-3C above dew point (20C)
Necessity of actively cooling the SiPMs to-be-determined
What assumtions are made for the cooling system?
Very little natural convection (1-10 W/m^2 K) due to light tight cover
System should experience negligable vibration due to (probably) low flow rate and cooling mass. However, flow-induced-vibration should be investigated.
Cooling Other Heat Sources?
Almost certainly not necessary, will keep in mind.
Detector Heat Sources & Geometry
Indicate the heat sources at the detector and how they are distributed over the detector volume
are the heat sources at the rim of the detector or over the entire volume or a mix
fECal
Outflow
Inflow
SiPM Board
FEB
Thermal Pads
Aluminum Heat Sink
Copper Tube (0.25in OD, 0.19in ID)
SiPM

## ECal Backward
Detector technology: PWO crystals (sensitive to temperature) with SiPM readout (in contact with crystals). FEBs located at 0.5 - 1 m from crystals.
Crystals cooling based on cold water plates around the outer and inner radii of the detector
FEBs will be thermally isolated from crystal assembly & cooled separately: water and/or air cooling under investigation
Heat source: FEB
| Estimate: | 1000 W total for all FEBs (location away from crystals)
Heat source: SiPM
| Estimate: | 10 W for all SiPM in contact with crystals
ANSYS steady and transient state simulation performed
Cooling requirements:
+/- 0.1 deg C stability on crystals
Crystal temperature: 15-25 deg C
No further need of cooling outside the detector

## HCAL Backward
Cooling requirements
FEEs are going to be placed in front, each one produces 4.5 W
With 2352 channels per layer=23520 total (368 ASICs) the dissipated power is 1656 W. In principle, by integrating signals and reducing the number of channels it can be reduced to ~0.5 kW. On the other hand by extending the detector by +50% and separate readout from extra layers the max power is ~2.4 kW.
To keep the nearby EEEMCal cool we may need to couple it to steel or use some other heatsink.
No heat simulations done.
No need for coolant.
SiPMs in nHCal will be cooled by the surounding steel, so this should keep them stable at room temperature.
For FEEs cooling with fans not feasible, because of magnetic field.

## Low Q2 Tagger
Maximum Power (Watts)
Timepix4 ASIC | 9
2x6 stave | 108
6x6 layer | 324
4 layer station | 1296
Total 2 Stations | 2592
Timepix4 with a Silicon sensor will be the technology for the tracker.
Each Timepix4 chip can require 9 Watts, this may be reduced to suit the requirements in precission of the Low-Q2 tagger.
There is no need to cool the tracker below a reasonable room temperature but a controlled environment is preferable.
Current design has 6x6 Timepix4 ASICs in a layer arranged in 2x6 staves. There are 4 layers per station and 2 stations in total, giving 288 chips. The design is evolving to require fewer detectors but this is still under investigation

## B0 Tracker
AC-LGADs with EICROC readout is the technology.
Full dimensions are still a bit in flux due to the beam pipe + B0 magnet + mounting rail dimentions.
Head load is expected to be between 1-2 mW/ch, with each "stave" having 1024 x 3 = 3072 channels, or between 3-6 Watts of power dissipation per stave (assuming Roman pot staves tiles onto the B0 geometry).
In the present understanding we need about 150 sensors per B0 layer (50 staves). This means 150 - 300 Watts power dissipation per layer.

## B0 Calo
transverse view of the calorimeter crystals, each is 2 x 2 x 20 cm3
Each crystal will have 4 SiPM
a coarse estimate of 4.5 mW per SiPm leads to a total of about 2W
other elements of the FEE are not yet accounted for
the crystals should operate at 20℃ ± 5℃
but to limit fluctuation in response should be stable at the level of 0.1 ℃
(needed accounting & engineering is still lacking)

## OffMomentum
AC-LGADs with EICROC readout is the technology.
Full dimensions are about 20cm vertical and ~ 10cm horizontal.
Head load is expected to be between 1-2 mW/ch, with each "stave" having 1024 x 3 = 3072 channels, or between 3-6 Watts of power dissipation per stave.
To tile the full OMD region, we need ~ 22 staves per layer, so ~ 60 to 150 Watts of power dissipation is expected, per OMD layer.

## RomanPots
AC-LGADs with EICROC readout is the technology.
Full dimensions are about 26cm horizinal and ~ 12cm vertical, but this is in the "closed" position.
Head load is expected to be between 1-2 mW/ch, with each "stave" having 1024 x 3 = 3072 channels, or between 3-6 Watts of power dissipation per stave.
This amounts to about between 100-200 Watts of total power dissipation per layer of Roman pots.

## ZDC
Cooling for the crystal ECal | HCAL
Precision cooling for the crystals and power cooling for the electronic readout will be achieved using water-cooling circuits. The cooling needs for individual crystals will be addressed through a thermal conduit extending from the crystal's rear face, through metal inserts, to the interface plate and support elements. A front thermal screen will also be installed to maintain the crystals at a controlled temperature. | No Cooling needs

## Compton RCS ElectronDet
e