## Slide 1
The hpDIRC Detector for the ePIC Experiment
Greg Kalicy
April 1st, 2025
EIC Project PID Review

## Slide 2
The hpDIRC has been selected for the barrel PID of the EIC ePIC Detector
Expected PID performance meets ePIC requirements (Yellow Report), separation: ≥ 3 s.d. π/K up to 6 GeV/c, ≥ 3 s.d. e/π up to ~1.1 GeV/c
Key elements, simulation and focusing lenses, validated in particle beams in 2018
Main remaining steps towards production readiness and TDR in 2026:
  Validation of reusing BaBar DIRC radiator bars
  Evaluation of sensors and readout ASIC (synergy with pfRICH and electronics/DAQ)
  Completion of mechanical design and integration
No challenges expected for ES&H and QA
hpDIRC DSC includes nine institutions with well-established expertise and interest in work packages
Plans for remaining studies, QA preparations, and construction match project schedule
Executive Summary

## Slide 3
Introduction………………………………………………………………
  ePIC requirements for hadronic barrel PID
  DIRC principle
High-performance DIRC for ePIC…………………………...….
  R&D milestones
  Preliminary design
  Simulation and expected performance
  Validation with prototype in particle beams and CRT
Main components………………………………………………….….
  Specifications, validation plans
Mechanical design and detector integration…………..…
ES&H and QA…………………………………………………………….
hpDIRC collaboration and responsibilities..............…...
Schedule……………………………………………………………………
Outline
Charge Question 1
Charge Question 1, 2, 3
Charge Question 4
Charge Question 5
Charge Question 6, 7
Charge Question 3, 4
Charge Question 4

## Slide 4
EIC physics requires clean and efficient PID over a wide angular range
Hadronic PID required to separate:
  Electrons from charged hadrons -> electron ID mostly provided by calorimeters, supplemented for lower momenta by DIRC/RICH detectors
  Charged pions, kaons, and protons from each other
Requirements for barrel PID (-1 ≤ η ≤ +1):
  Main separation power requirement: ≥ 3 s.d. π/K up to 6 GeV/c
  System needs to be radially compact (impact on cost of outer systems)
  Minor changes to detector position and dimensions do not impact hpDIRC performance (sensor in B-field and detector integration)
  Low demand on detector services (simplified integration and operation)
hpDIRC capable of reaching required performance at 6 GeV/c for 0.5 mrad tracking angular precision
Charged PID Requirements in ePIC
Summary of PID requirements in ePIC
Charge Question 1
Performance requirement definition

## Slide 5
Detection of Internally Reflected Cherenkov Light
Pioneered by the BaBar experiment at the SLAC National Accelerator Laboratory
Fused silica bars or plates used as radiator and light guide
Detector surface is outside active volume
Cherenkov angle is conserved during internal reflections and reconstructed from detected photons
Ultimate deliverable: PID likelihoods
BaBar DIRC achievement: 3 s.d. π/K separation up to 3.5 GeV/c
Accumulated ePIC hpDIRC hit pattern (Geant4)
DIRC Concept

## Slide 6
  Make DIRC less sensitive to background (main challenge for BABAR and SuperB)
      decrease size of expansion volume, replace water as medium, add focusing optics;
      find a way to place photon detector inside magnetic field.
  Investigate alternative radiator shapes (plates, disks), develop endcap device
  Push DIRC π/K separation to higher momentum
      improve angular resolution of tracking system, mitigate multiple scattering impact;
      use photon detectors better PDE, improve Cherenkov angle resolution per photon.
BABAR DIRC
Limited in BABAR by: Improve for future DIRCs via:
    ▪ size of bar image ~4.1 mrad ▪ focusing optics
    ▪ size of PMT pixel ~5.5 mrad ▪ smaller pixel size
    ▪ chromaticity (n=n(λ)) ~5.4 mrad ▪ better time resolution
    9.6 mrad 5-6 mrad per photon → 1 mrad per particle (EIC goal) in reach
SUPERB, BELLE II, PANDA & EIC
DIRC Cherenkov angle difference vs. momentum
Advancing DIRC Performance

## Slide 7
10+ years ago: DIRC good candidate for hadronic particle in EIC detector barrel – if π/K momentum coverage achieved by BaBar DIRC is increased by 50%
R&D for a high-performance EIC DIRC started in 2011 (synergetic with PANDA DIRC) (Funded by DOE/BNL/JLab as RD2011-3, eRD4, eRD14, eRD103, EICGENRandD2022_12)
EIC DIRC R&D Milestones:
  2012: First multi-layer high-refractive index lens concept to avoid photon loss at air gaps
  2012: First 2-layer and 3-layer prototype lenses produced by industry
  2014: Simulation showed that lens-based design is expected to reach 1mrad Cherenkov angle resolution, equivalent to 3 s.d. π/K separation at 6 GeV/c
  2015: First successful CERN beam test with muti-layer spherical lens
  2017: Identified sapphire and PbF2 as radiation-hard material candidates for lenses
  2018: Validated 3-layer spherical lens performance and Geant4 simulation with PANDA DIRC prototype with particle beam at CERN
  2019: First radiation-hard lens prototypes fabricated by industry
  2022: hpDIRC selected as barrel PID solution for EIC detector
  2024: Transportation and start of BaBar bars assembly, separation of first bar
High-Performance DIRC R&D
Initial 3-layer lens concept
radiation hardness of sapphire
Cherenkov angle per particle, CERN 2015
radiation hard lens prototypes
Charge Question 3
Detector and readout design

## Slide 8
Extending DIRC π/K separation coverage to 6 GeV/c
  Concept: fast focusing DIRC, utilizing high-resolution 3D (x,y,t) reconstruction.
  Radiation-hard 3-layer spherical lens to reduce bar image size and shape imaging plane;
  Lifetime-enhanced MCP-PMTs with fine anode segmentation to reduce pixel size;
  Fast photon timing for chromatic dispersion mitigation and background rejection;
  Narrow bars for robust performance in high-multiplicity jet events;
  Compact expansion volume to simplify integration into central detector.
  Benefit from additional ePIC detector improvements:
    high-precision tracking, expect 0.5mrad polar angle resolution;
    post-DIRC tracking layer (EMCal AstroPix) for multiple scattering mitigation.
  Predicted performance for central rapidity range -1.5 ≤ η ≤ +1.5: 3σ π/K separation up to at least 6 GeV/c (Cherenkov angle resolution ≤1mrad), supplemental e/π separation up to 1.1 GeV/c.
Schematic view of the EIC hpDIRC
High-Performance DIRC Overview
Charge Question 3
Detector and readout design

## Slide 9
Radiator bars:
  Barrel radius: 780 mm, 12 sectors
  10 long bars per sector, 4500 mm x 35 mm x 17 mm (L x W x T)
  Long bar: 4 bars, glued end-to-end,
  Short bars made from highly polished synthetic fused silica
  Flat mirror on far end
Focusing optics:
  Radiation-hard 3-layer spherical lens (sapphire)
Expansion volume:
  Solid fused silica prism: 25 x 35 x 30 cm3 (H x W x L)
Readout system:
  MCP-PMT Sensors (Photek/Incom)
  ASIC-based Electronics (FCFD)
hpDIRC Components: (CFRP bar box and readout box, flat mirror are not shown)
Radiator bar
Focusing
Lens
Expansion Volume
Sensors
Readout Electronics
hpDIRC Design
Charge Question 3
Detector and readout design

## Slide 10
Stand-alone Geant4 Simulation
  Realistic optics, geometry, and material properties – based on prototypes and experimental data, wavelength-dependent material properties and processes
  Validated with test beam data
  Used for design optimization studies and to test novel design options
Full ePIC Simulation:
  Functionality of Stand-alone simulation imported and integrated
  Repeat physics background and detector impact study in the full software stack
Reconstruction and PID methods:
  Geometrical (BABAR-like), robust and fast method based on Look-Up Tables, delivers Cherenkov angle per particle and Single Photon Resolution (useful for calibration and in prototype tests), does not depend on precise time measurement
  Time Imaging (Belle II TOP-like), uses Probability Density Functions (analytical or simulation-based), makes optimum use of precision of position and time information
Pythia events in hpDIRC simulation
hpDIRC Simulation

## Slide 11
Recent hpDIRC related studies:
  Confirmed robust performance in magnetic field, and using physics events (Pythia) to include backgrounds, multiple tracks per bar (WSU)
  Performance with new iterations of ePIC angular track resolution maps (GSI)
  Verifying optimal sensor coverage (CUA, GSI)
  Study of bar imperfections impact on the performance relevant for the BaBar bars qualification (GSI)
  Impact of bars misalignments on performance (Jazan)
  Machine Learning approach to reconstruction (W&M)
  Preparations for hpDIRC prototype operation at CRT (SBU, GSI, ODU)
hpDIRC Simulation
time imaging reconstruction
0.5 mrad tracking, 100 ps photon timing
π/K separation power at 6 GeV/c
Geant4
→ Performance requirements reached: ≥ 3 s.d. π/K separation at 6 GeV/c for all angles
Simulation studies performed with
Stand-alone Geant4 simulation
Single particles from particle gun
6 GeV/c momentum
No magnetic field, no other ePIC subsystems
0.5 mrad tracking resolution

## Slide 12
Expected hpDIRC Performance vs. Tracking
π/K separation power at 6 GeV/c
Geant4
time-based imaging reconstruction
→ High-precision angular resolution crucial for
reaching required hpDIRC performance.
Simulation studies performed with
Stand-alone Geant4 simulation
Single particles from particle gun
6 GeV/c momentum
No magnetic field, no other ePIC subsystems
Impact of tracking angular resolution on hpDIRC performance
π/K Cherenkov angle difference at 6 GeV/c: ΔΘc≈3 mrad
Yellow Report tracking requirement: 0.5 mrad resolution at 6 GeV/c
Charge Question 2
Performance and construction plans

## Slide 13
View from inside the Barrel DIRC dark box
Dark box for optics
(bar, lens, prism)
Rotation stage (remote controlled)
Frontend electronics (PADIWA) (air-cooled)
DAQ boards (TRB)
MCP-PMT array
Narrow fused silica bar, hpDIRC 3-layer spherical lens
30 cm-deep fused silica prism
2x4 PHOTONIS Planacon MCP-PMT array (larger pixels, slower readout electronics than EIC)
PiLas picosecond laser calibration system
7 GeV/c π/p beam equivalent to 3.5 GeV/c π/K
MCP-TOF system to cleanly tag π and p events
Schematic view of 2018 prototype
Performance validation: synergetic beam test with PANDA prototype at CERN PS in 2018
Validation of Simulation and Components
Charge Question 2
Performance and construction plans

## Slide 14
Measured Cherenkov angle resolution per photon (SPR), photon yield, and π/K separation in excellent agreement with expectation and Geant4 simulation
Achieved π/K separation power of Nsep=5.2 s.d. with time imaging reconstruction for PANDA configuration, will improve with smaller pixels, better PDE and timing
Same simulation/reconstruction code used for EIC high-performance DIRC -> Confidence in Geant prediction for hpDIRC performance
Photon yield
π/p separation power @ 7 GeV/c
SPR
observed hit pattern at polar angle 20°
Performance validation: 2018 prototype at CERN PS
Validation of Simulation and Components
Charge Question 2
Performance and construction plans
π/p separation

## Slide 15
Validation of Simulation and Components
Charge Question 2
Performance and construction plans
Cosmic Ray Telescope (CRT) at SBU
Facility to test incremental upgrades of prototype components, performance evaluation
Initial PANDA Barrel DIRC-based prototype to commission setup
Modular design will allow to add new ePIC hpDIRC components once they become available
Cherenkov Tagger to select muons above 3.5 GeV/c
Three tracking stations for high-precision 3D-track reconstruction (location optimized with simulations)
PicoSec detector for event timing (Jlab group committed prototype and personnel to project)
Geant4 simulation used to optimise setup arrangement
CRT setup CAD schematic
Cherenkov Tagger
Tracking Station 1
Tracking Station 2
Tracking Station 3
PicoSec
hpDIRC Prototype
1.5 m
4 m
Geant4 simulation of CRT setup

## Slide 16
Validation of Simulation and Components
Charge Question 2
Performance and construction plans
Trackers in the CRT
Large area PICOSEC
DAQs
Cherenkov
Tagger
Cherenkov
Tagger
hpDIRC
Prototype
hpDIRC
Prototype

## Slide 17
  BaBar DIRC decommissioned in 2010, SLAC/DOE made DIRC bars available for reuse, 4 bar boxes awarded to JLab and installed as GlueX DIRC in 2018, remaining 8 boxes awarded to JLab for potential use in EIC DIRC
    Bar boxes transported to JLab in April 2024
  Full-size bar boxes are too long, do not fit into EIC central detector, wedges deteriorate resolution: need to disassemble bar boxes for reuse
    Facility, setups, and tools developed, disassembly of first bar box in progress
  hpDIRC barrel requires total of 360 short bars (1.225 m length)
  Eight bar boxes currently located at JLab could yield up to 384 short bars, sufficient to cover rapidity range -1.65 ≤ η ≤ +1.65
  Additional 120 bars required for the light guide section, η ≤ -1.65, to couple to lenses
  Quality of bar surfaces, 25 years after initial production and disassembly, to be verified
    QA of first disassembled bars in underway
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
More in talk from Jochen Schwiening

## Slide 18
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
Successful transport of 8 DIRC bar boxes in April 2024
  Low altitude road from SLAC, CA to JLab, VA
  Goal: Kept shocks on Bar box below 1g
  (Transportation crates with inner suspension, shock-absorbing foam, hydraulic shocks, air shocks, shock-absorbing donuts, air-ride and temperature control trucks)
1
2
5
3
6
10
7
9
4
8
Exterior crate
Interior Basket
Barbox
Wedge
Inside of transportation crates
Loaded bar boxes ready for transport
8 bar boxes in two trucks
Bar boxes in Jlab ready for disassembly
Sample of vibrations during transport and location of accelerometers
Andrew Lumanog (JLab), Tyler Lemon (JLab), Greg Kalicy (CUA), Jochen Schwiening (GSI)

## Slide 19
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
Disassembly process in JLab:
Bar boxes are disassembled and bars are separated in clean tent
Cleaning station to remove residue glue, visually inspect bars
QA laser lab to inspect quality of the bars after disassembly
Measured bars are wrapped, tagged and stored in cabinets
Team working on disassembly:
3 JLab Technicians: Andrew Lumanog, Caleb Graham, David Edwards
2 Scientists: Greg Kalicy (CUA), Sourav Tarafdar (JLab)
JLab DSG Group: Tyler Lemon, George Jacobs, Mindy Leffel
Graduate Students: Shelby Arrigo (W&M)
Bar boxes in
transportation
crates
Clean disassembly tent
Storage
cabinets
QA laser lab
Cleaning station

## Slide 20
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
Bar box with disassembled aluminum shell
Bar box on CNC
Removing aluminum shell

## Slide 21
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
Heat guns softening glue joints between bars
Highly adjustable clamps to separate bars
Weight with short travel pulling on clamp

## Slide 22
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
Glue residue after separation
Glue residue mid-way through cleaning
Cleaned bar end
Bar end immersed in acetone bath
"Bathtub" in position for glue removal
Glue removal

## Slide 23
Reuse of BaBar DIRC Bars
Charge Question 4
Fabrication and assembly plans
Measurement with 442 nm laser
Laser setup built at JLab based on similar setup at GSI for PANDA Barrel DIRC
Reflection coefficient measurement allows to evaluate surface quality of the bars with sub nm precision
HeCd laser with two wavelengths is used (325 nm, 442 nm)

## Slide 24
Photek MAPMT 253
INCOM Gen III HRPPD prototype (front/back view)
Readout
hpDIRC readout: Microchannel-Plate PMTs + ASIC-based electronics
MCP-PMTs capable of meeting all hpDIRC requirements (A. Lehmann review talk at RICH2022)
Baseline sensor for hpDIRC: 2’’ Photek MAPMT 253 MCP-PMTs
Potential solution: DC-coupled Incom HRPPD Making use of synergy with pfRICH, optimizing cost and workforce
  See sensor and readout presentation tomorrow
Baseline front-end board: FCFD
Synergetic development with ePIC AC-LGAD and pfRICH systems
Low-power ASIC, 128 channels per board
Will deliver hit time, time over threshold
  See DAQ, sensor and, readout presentations tomorrow
A. Kiselev
Charge Question 3
Detector and readout design
FCFDv0
More in talks from Alexander Kiselev and David Abbott

## Slide 25
Mechanical Support and Integration
Charge Question 5
Detector integration plans
Engineering/integration supported for hpDIRC from EIC project:
Avishay Mizrahi (MIT)
Kris Clevland (JLab)
Forward
Backward
Bar Box
Prism Box
Support Ring
Connection to ePIC Barrel
More in talk from Alex Eslinger

## Slide 26
Mechanical Support and Integration
Charge Question 5
Detector integration plans
Bar Boxes:
CFRP Shell
  Approximately 4.6m Long
  Thickness: 1.5 mm – 3 mm
    (thicker in key locations)
Contains Quartz Radiator Bars, Mirrors, and Lenses
  10 long bars each formed by 3 short bars and a Light Guide section
  Radiator bars supported on rounded nylon buttons to minimize losses
Capped on each end
  Spring End Plate at Forward End
  Adjustable spring assemblies for each bar
  Maintain compression in glue joints
  Optically transparent window at backward end
  Counters force from springs at forward end
  Provides gas-tight boundary
F = ~900N
Window
Prism
Lens
Light Guide
Bar
Fixed Button
Spring Button
Bar
Mirror
Mirror Plunger
Spring
End Plate
Spring Shell

## Slide 27
Mechanical Support and Integration
Charge Question 5
Detector integration plans
Readout Box
Box shell
  Current Material: CFRP
  Current Thickness: 5 mm
Contains expansion prism, calibration laser system
Supports sensors, readout electronics, and associated services
Mechanical design solutions for two sensor options:
  MCP-PMT (baseline)
  HRPPD
Optical cookie connection between prism and bar box window
MCP-PMT
HRPPD
Light-Tight Cover
Alignment Grid
Sensors, Readout Electronics, Service connections
Expansion Prism
Calibration Prism
Prism Box Shell
Calibration Prism Cover
Carrier
Adjustable Spring Prism-Bar Box Connector

## Slide 28
The system does not use any flammable gases or cryogenic liquids
The bar boxes and readout boxes will be purged by boil-off dry nitrogen at a flow rate of a few liters per hour per box
Hazard: electrical shock Sensors require typical PMT high voltages (up to 2.5 kV) → HV connections are not accessible in normal operation mode; HV module will be keyed off, operation only by trained personnel
Hazard: laser radiation The hpDIRC will use a class 3B PiLas laser pulser to calibrate sensors and electronics → laser will be fully contained in optical fibers, operation only by trained personnel
Hazard: chemical leak → Cooling liquid will circulate at sub-atmospheric pressure (details TBD)
The final detector will be large, heavy, and delicate, precise rigging and procedures required for safe installation
Disassembly and bar QA efforts underway at JLab have process-specific safety protocols, passed several JLab safety reviews
The hpDIRC system will comply with all applicable ES&D and OSHA regulations
ES&H
Charge Question 6
ES&H and QA
Bars cleaning station
QA laser lab

## Slide 29
QA Plan
Quality assurance plans (100% components will be verified/tested)
Combination of process control/QA at vendor site and lab measurements
Radiator bars and light guides: vendor QA for mechanical properties, laser scanning system at JLab to monitor internal photon transport efficiency of disassembled BaBar DIRC bars and/or new DIRC bars
Sensors and electronics: laser pulser systems at CUA/JLab/USC (TBD) to measure gain, quantum and collection efficiency, dark count rate, etc
Lenses: laser lab at ODU to evaluate shape of focal plane
Prisms: vendor QA, checks at WSU
Bar boxes, prism boxes: vendor QA, checks at CRT (SBU or JLab)
Assembled DIRC module (bar box coupled to readout box, vertical slice): Cosmic Ray Telescope at SBU or JLab
Installed DIRC module in ePIC: picosecond laser pulser calibration system, cameras to monitor optical coupling between sensors, prisms, lenses
DIRC laser lab at JLab
Lens evaluation setup at ODU
DIRC lab and Cosmic Ray setup at SBU (photos, Geant4)
Charge Question 6, 7 Recommendation 3
ES&H and QA

## Slide 30
Schedule
hpDIRC technical schedule consistent with project schedule
hpDIRC scheduled for installation into ePIC in June 2031, expect hpDIRC readiness for installation well before that date
Charge Question 4
Fabrication and assembly plans

## Slide 31
ePIC hpDIRC group (DSC)
  Core formed by groups that have been involved in the BaBar, GlueX, and PANDA DIRC counters, and in the EIC DIRC R&D program, for many years, some since 2011
  Current number of people across all institutions actively involved is 25
Responsibility examples:
  Radiators: transport and disassembly of BaBar DIRC bar boxes, validation quality of disassembled bars, optional QA of new bars/plates for light guide section – JLab, CUA, W&M
  Bar boxes: gluing of bars and lenses, assembly of bar boxes, QA in Cosmic Ray Telescope – SBU, JLab
  Lenses: evaluation of focal plane, QA – ODU, CUA
  Sensors: QA, readout chain tests – USC, Glasgow
  Readout boxes: assembly, QA – WSU
  Simulation, reconstruction – CUA, GSI, W&M, WSU, Jazan
hpDIRC DSC
Charge Question 4
Fabrication and assembly plans

## Slide 32
ePIC hpDIRC: Fast focusing DIRC concept, developed over 10+ years of EIC R&D
Expected PID performance meets ePIC requirements (Yellow Report), separation: 3 ≥ s.d. π/K up to 6 GeV/c, ≥ 3 s.d. e/π up to ~1.2 GeV/c
Key elements, simulation and focusing lenses, validated in particle beam in 2018
Main remaining steps towards production readiness and TDR:
  Validation of reusing BaBar DIRC radiator bars
  Evaluation of sensors and readout ASIC (synergy with pfRICH)
  Completion of mechanical design and integration
Presented plans for ES&H and QA, modular integration concept
ePIC hpDIRC group includes nine institutions, 23 active participants with well-established expertise
Plans for remaining studies, QA, and construction match project schedule
Summary

## Slide 33
 | Charge Question | Slide Number
1 | Are the technical performance requirements appropriately defined and complete for this stage of the project? | 4
2 | Are the plans for achieving detector performance and construction sufficiently developed and documented for the present phase of the project? | 11-14
3 | Are the current designs and plans for detector and electronics readout likely to achieve the performance requirements with a low risk of cost increases, schedule delays, and technical problems? | 7-9, 24
4 | Are the fabrication and assembly plans for the various particle identification detector systems consistent with the overall project and detector schedule? | 17-23, 30
5 | Are the plans for detector integration in the EIC detector appropriately developed for the present phase of the project? | 25-27
6 | Have ES&H and QA considerations been adequately incorporated into the designs at their present stage? | 28-29
7 | Have the recommendations from previous reviews been adequately addressed? | 29
Charge Question Roadmap
Thank you all for your attention.

## Slide 34
Extra Slides

## Slide 35
Expected hpDIRC Performance
time imaging reconstruction
0.5 mrad tracking, 100 ps photon timing
pixel-based (geometric) reconstruction
Cherenkov angle resolution per photon (SPR)
π/K separation power at 6 GeV/c
Photon yield per particle
Geant4
Geant4
Geant4
pixel-based (geometric) reconstruction
→ Performance requirements reached: ≥ 3 s.d. π/K separation at 6 GeV/c for all angles
Simulation studies performed with
Stand-alone Geant4 simulation
Single particles from particle gun
6 GeV/c momentum
No magnetic field, no other ePIC subsystems
0.5 mrad tracking resolution
Charge Question 2
Performance and construction plans

## Slide 36
Photosensors
hpDIRC sensor requirements
Single photon sensitivity in ePIC magnetic field: 106 gain at ~ 1T
Fast timing for single photons: timing precision (rms) < 100 ps
Large active area ratio for tiled sensors: goal > 75%
High PDE in visible range: goal > 25% at 400 nm
Small pixels: anode pixel size < 3.5 mm
Tolerance for high photon rates: goal > 0.5 MHz/cm2
Tolerance for high occupancies: up to 200+ photoelectrons per particle, need DC-coupled anodes
Long lifetime: goal > 10 C/cm2
Photek MAPMT 253
6 μm pores 8x8 anodes (PANDA)
S. Krauss, RICH2022
PHOTONIS XP85012-S
10 μm pores 8x8 anodes (PANDA)
Expected number of photoelectrons per particle per 12 cm x 12 cm sensor
Geant4 single particles
Charge Question 3
Detector and readout design

## Slide 37
Geant4 simulation: photon yield
Barrel DIRC counters (PANDA, EIC) require focusing for wide range of photon angles
Conventional plano-convex lens with air gap limits DIRC performance
  Significant photon yield loss for particle polar angles around 90°, gap in DIRC PID
  Distortion of image plane, PID performance deterioration
Innovative solution:
  3-layer compound lens (without air gap):
  Creates flat focal plane – matched to fused silica prism shape
  Avoids photon loss and barrel PID gap
  Detailed radiation-hardness studies performed with 60Co source, neutron irradiation next
  Lanthanum crown glass (LaK33B) for PANDA, rad-hard sapphire or PbF2 for EIC
  Industrial fabrication of lenses demonstrated
  Performance of spherical 3-layer lenses validated with PANDA Barrel DIRC prototype
  layer of high-refractive index material (focusing/defocusing) sandwiched between two layers of fused silica
PANDA lens prototypes
Focusing Optics
Charge Question 4
Fabrication and assembly plans

## Slide 38
Radiation-hard 3-layer lens prototypes
Radiation hardness of sapphire
(𝛾 dose from 60Co)
  layer of high-refractive index material (focusing/defocusing) sandwiched between two layers of fused silica
Focusing Optics
Barrel DIRC counters (PANDA, EIC) require focusing for wide range of photon angles
Conventional plano-convex lens with air gap limits DIRC performance
  Significant photon yield loss for particle polar angles around 90°, gap in DIRC PID
  Distortion of image plane, PID performance deterioration
Innovative solution:
  3-layer compound lens (without air gap):
  Creates flat focal plane – matched to fused silica prism shape
  Avoids photon loss and barrel PID gap
  Detailed radiation-hardness studies performed with 60Co source, neutron irradiation next
  Lanthanum crown glass (LaK33B) for PANDA, rad-hard sapphire or PbF2 for EIC
  Industrial fabrication of lenses demonstrated
  Performance of spherical 3-layer lenses validated with PANDA Barrel DIRC prototype
Charge Question 4
Fabrication and assembly plans

## Slide 39
Radiation-hard 3-layer lens prototypes
Close collaboration with Prof. Charles I. Sukenik (ODU)
Geant4 simulation
Lens Focusing
Radiation hardness of sapphire
(𝛾 dose from 60Co)
  Detailed scans of lens focusing properties with laser in optical lab at ODU

## Slide 40
Expected hpDIRC Performance vs. B Field
time imaging reconstruction
0.5 mrad tracking, 100 ps photon timing
pixel-based (geometric) reconstruction
Cherenkov angle resolution per photon (SPR)
π/K separation power at 6 GeV/c
Photon yield per particle
Geant4
Geant4
pixel-based (geometric) reconstruction
Geant4
Simulation studies performed with
Stand-alone and ePIC software framework
Single particles from particle gun
6 GeV/c momentum
MARCO: 1.7 T magnetic field
0.5 mrad tracking resolution
→ No significant impact of magnetic field on hpDIRC performance at 6 GeV/c
Impact of magnetic field on hpDIRC performance
Charge Question 2
Performance and construction plans

## Slide 41
R. Abdul Khalek et al., Science Requirements and Detector Concepts for the Electron-Ion Collider : EIC Yellow Report, doi:10.1016/j.nuclphysa.2022.12447
Requirements from Yellow Report

## Slide 42
hpDIRC sensors
1.7 T ePIC Solenoid Magnetic Field Map
→ hpDIRC detector plane located in modest magnetic field of less than 0.5 T
hpDIRC in ePIC Magnetic Field

## Slide 43
2 T ePIC Solenoid Magnetic Field Map
→ hpDIRC detector plane located in modest magnetic field of less than 0.5 T
hpDIRC sensors
hpDIRC in ePIC Magnetic Field

## Slide 44
→ hpDIRC detector plane located in modest magnetic field of less than 0.5 T
1.7 T ePIC Solenoid Magnetic Field Map
hpDIRC in ePIC Magnetic Field
field in radial direction
field in longitudinal direction
hpDIRC
hpDIRC
sensors
sensors
https://wiki.bnl.gov/EPIC/index.php?title=Experimental_Solenoid

## Slide 45
→ hpDIRC hit pattern spread out over larger area, more pixels, in ePIC magnetic field
Impact of magnetic field on hit pattern (accumulated hit pattern, 5000 charged pions, 6 GeV/c, 30° polar angle)
hpDIRC in ePIC Magnetic Field
with magnetic field
without magnetic field

## Slide 46
→ hpDIRC hit pattern spread out over larger area, more pixels, in ePIC magnetic field
Impact of magnetic field on hit pattern (accumulated hit pattern, 5000 charged pions, 6 GeV/c, 35° polar angle)
hpDIRC in ePIC Magnetic Field
with magnetic field
without magnetic field

## Slide 47
Calculate average time of Cherenkov light emission per particle from difference between measured and calculated arrival time for each detected photon
Good precision for large number of photoelectrons, can be useful as TOF “stop time” if event T0 is known
standalone Geant4 simulation, single particles, no magnetic field, no background
hpDIRC Timing

## Slide 48
→ hpDIRC hit pattern spread out over larger area, more pixels, in ePIC magnetic field
Impact of magnetic field on hit pattern (accumulated hit pattern, 5000 charged pions, 6 GeV/c, 35° polar angle)
hpDIRC in ePIC Magnetic Field
with magnetic field
without magnetic field

## Slide 49
→ Expect less than 109 neutrons cm-2 and less than 0.1krad per year at DIRC lens location
hpDIRC in ePIC Radiation Environment
1 MeV equivalent neutron fluence - PYTHIA 18x275 GeV e+p minimum bias events at top luminosity (500 kHz min-bias rate, 6-month run)
hpDIRC
lenses
hpDIRC lenses
Accumulated ionizing radiation dose - PYTHIA 18x275 GeV e+p minimum bias events at top luminosity (500 kHz min-bias rate, 6-month run)
https://wiki.bnl.gov/EPIC/index.php?title=Radiation_Doses
1 MeV equivalent neutron fluence
Accumulated ionizing radiation dose (hadrons, e+/-, γ)
Radiation Dose and Neutron Flux from Minimum-Bias e+p events

## Slide 50
Four materials studied up to 2 Mrad:
  sapphire
  lead fluoride (PbF2)
  lanthanum crown glass (S-YGH51)
  polycarbonate (PC)
Sapphire confirmed to be extremely radiation hard. PbF2 showed very small deterioration.
Initial photo-annealing and luminescence tests.
hpDIRC Radiation Tests

## Slide 51
Monochromator
Co60 Chamber
hpDIRC Radiation Tests

## Slide 52
Validation of the BaBar DIRC bar reuse:
  BaBar bar box transfer from SLAC to JLab and disassembly
  Validation of mechanical and optical bar quality in QA laser setup
hpDIRC studies in simulation:
  Study of the hpDIRC performance with background
Evaluating components and prototypes in DIRC labs (ODU, JLab, GSI, BNL)
hpDIRC prototype program:
  Modular hpDIRC prototype in Cosmic Ray Telescope at SBU
  Incremental hpDIRC optical components integration and evaluation
  Adaptation and evaluation of sensors and readout electronics in hpDIRC prototype
Generic DIRC R&D explores innovate optical DIRC configurations to create opportunities for cost reduction, performance improvement, and complementarity
hpDIRC in F4A
DIRC lab/CRT space at SBU
DIRC barboxes in SLAC
BaBar DIRC bars
Hybrid DIRC in Geant4
10s of hpDIRC in CRT simulation
Remaining Studies

## Slide 53
Performance improvement
EIC detector barrel length requires additional fused silica bars or plate to connect BaBar DIRC bars to prism
Narrow bars could be obtained by cutting and repolishing BaBar DIRC bars or by ordering new bars from industry
At RICH 2016 J. Va’vra showed the “ultimate fDIRC” concept for SuperB with then best-in-class predicted DIRC performance Concept: use single short wide plate as transition light guide between BaBar DIRC bars and expansion volume
For EIC hpDIRC design: use plate as light guide between BaBar DIRC bars and prism, combine with lens focusing
Would significantly reduce cost compared to new narrow bars and potentially improve hpDIRC performance
GEANT4 visualization of hybrid of bars and plate in each sector
1225mm
1225mm
1225mm
400-700mm
J. Va’vra, RICH2016
“ultimate fDIRC”
doi.org/10.1016/j.nima.2017.02.044
EIC hpDIRC
Light Guide Section
Study supported via generic R&D as EICGENRandD2022_12

## Slide 54
Radiator geometry | Narrow bars (35mm) | Wide plates (450mm) | Narrow bars (53mm) | Narrow bars (35mm)
Barrel radius | 85cm | 115cm | 48cm | 72cm
Bar length | 490cm (4×122.5) | 250cm (2×125) | 240cm (2×120) | 488cm (3×122.5 + 1x90.5)
Number of long bars | 144 (12×12 bars) | 16 (16×1 plates) | 48 (16×3 bars) | 120 (12×10 bars)
Expansion volume | 110cm, ultrapure water | 10cm, fused silica | 30cm, fused silica | 30cm, fused silica
Focusing | None (pinhole) | Mirror (for some photons) | Spherical lens system | Spherical lens system
Photodetector | ~11k PMTs | ~8k MCP-PMT pixels | ~8k MCP-PMT pixels | ~74k MCP-PMT pixels
Timing resolution | ~1.5ns | <0.1ns | ~0.1ns | ~0.1ns
Pixel size | 25mm diameter | 5.6mm×5.6mm | 6.5mm×6.5mm | 3.2mm×3.2mm
PID goal | 3 s.d. π/K to 4 GeV/c | 3 s.d. π/K to 4 GeV/c | 3 s.d. π/K to 3.5 GeV/c | 3 s.d. π/K to 6 GeV/c
Timeline | 1999 - 2008 | Running (installed 2016) | TDR in 2017 | TDR-ready in 2024
PANDA BARREL DIRC
BABAR DIRC
BELLE II TOP
EPIC HPDIRC
*Preliminary design
*
Barrel DIRC Overview

## Slide 55
e/π separation at low momentum
Yellow report effort identified need for supplemental e/π suppression from PID systems to support EM calorimeter at lower momentum
hpDIRC e/π performance at low momentum very different from high-momentum domain, dominated by multiple scattering (MS) and EM showers in DIRC bars
Expected performance (without any MS mitigation): ≥ 3 s.d. e/π separation at 1.1 GeV/c (caveat: non-Gaussian tails)
Study of potential improvements from DIRC “ring center fit” and use of additional tracking point from AstroPix layer outside DIRC radius starting (also expected to further improve high-momentum π/K separation)
Example from Geant4 e/π log-likelihood difference 1 GeV/c momentum, 30° polar angle, 0.8mrad tracking resolution
Geant4, stand-alone, single particles, no B field
hpDIRC expected e/π separation at 1.2 GeV/c without MS mitigation
Expected hpDIRC Performance
Charge Question 2
Performance and construction plans