## Slide 1: Investigations into Multi-Energy Extraction at MedAustron
Katrin Holzfeind, Elisabeth Renner (TU Wien)
Fabien Plassard, Markus Wolf (MedAustron)

## Slide 2: MedAustron
Located in Wiener Neustadt, Austria
Based on PIMMS [1] study
Synchrotron/ main ring
Clinical treatment
  Proton: 62 –252 MeV
  Carbon: 120 – 402 MeV/u
© MedAustron

## Slide 3: MedAustron
Located in Wiener Neustadt, Austria
Based on PIMMS [1] study
Synchrotron/ main ring
Clinical treatment
  Proton: 62 –252 MeV
  Carbon: 120 – 402 MeV/u
Non-clinical research in first irradiation room (IR)
  Low flux
  Helium
  800 MeV p+
  Different research fields!
TU Wien Technical University Vienna
© MedAustron

## Slide 4: Multi-Energy-Extraction (MEE)
Motivation
  Ion beam therapy: Bragg-peak is used to irradiate cancerous tissue layer by layer [1]
  Depth of Bragg-peak depends on energy of ions
  Currently: Single energy extraction
  Duration of treatment can be reduced significantly
  Efficiency of machine can be improved
© MedAustron

## Slide 5: Multi-Energy-Extraction (MEE)
Motivation
  Ion beam therapy: Bragg-peak is used to irradiate cancerous tissue layer by layer [1]
  Depth of Bragg-peak depends on energy of ions
  Currently: Single energy extraction
  Duration of treatment can be reduced significantly
  Efficiency of machine can be improved
Study goals
  Bunched RFKO
  Beam dynamics during energy changes and effects on the spill
  Identify performance constraints
This presentation: establishing and benchmarking of setup
© MedAustron

## Slide 6: Previous Studies at MedAustron
Summary of studies done by Florian Kühteubl [2], Dale Prokopovich and Fabien Plassard [3]
MEE test with COSE
  10 energy steps of 1 MeV [4]
RFKO
  Optic settings with similar spill characteristics as betatron driven extraction for clinical use [4]
  Loss mitigation studies for bunched RFKO
    Reduce bucket height
    Increase tune-distance by using radial loop offset
930 V
140 V

## Slide 7: Developments for RFKO MEE
Newly implemented MEE
  Custom waveform for power converters of all main ring magnets and the RF cavity
  Enables up- and downwards operation for arbitrary energy steps

## Slide 8: Developments for RFKO MEE
Newly implemented MEE
  Custom waveform for power converters of all main ring magnets and the RF cavity
  Enables up- and downwards operation for arbitrary energy steps
New RFKO setup: µTCA-crate
  Combination of different signals, amplitude modulation and triggering are possible [5]
  First benchmark measurements
    Binary phase shift keying with same band width but appropriate central frequency applied for extraction

## Slide 9: Developments for RFKO MEE
Newly implemented MEE
  Custom waveform for power converters of all main ring magnets and the RF cavity
  Enables up- and downwards operation for arbitrary energy steps
New RFKO setup: µTCA-crate
  Combination of different signals, amplitude modulation and triggering are possible [5]
  First benchmark measurements
    Binary phase shift keying with same band width but appropriate central frequency applied for extraction
Implications of setup
  All main ring magnets are current regulated
  Possible mismatches in energy

## Slide 10: Developments for RFKO MEE
Newly implemented MEE
  Custom waveform for power converters of all main ring magnets and the RF cavity
  Enables up- and downwards operation for arbitrary energy steps
New RFKO setup: µTCA-crate
  Combination of different signals, amplitude modulation and triggering are possible [5]
  First benchmark measurements
    Binary phase shift keying with same band width but appropriate central frequency applied for extraction
Implications of setup
  All main ring magnets are current regulated
  Possible mismatches in energy  B-field regulation in near future [6]

## Slide 11: Developments for RFKO MEE
Newly implemented MEE
  Custom waveform for power converters of all main ring magnets and the RF cavity
  Enables up- and downwards operation for arbitrary energy steps
New RFKO setup: µTCA-crate
  Combination of different signals, amplitude modulation and triggering are possible [5]
  First benchmark measurements
    Binary phase shift keying with same band width but appropriate central frequency applied for extraction
Implications of setup
  All main ring magnets are current regulated
  Possible mismatches in energy  B-field regulation in near future [6]

## Slide 12: Developments for RFKO MEE
Newly implemented MEE
  Custom waveform for power converters of all main ring magnets and the RF cavity
  Enables up- and downwards operation for arbitrary energy steps
New RFKO setup: µTCA-crate
  Combination of different signals, amplitude modulation and triggering are possible [5]
  First benchmark measurements
    Binary phase shift keying with same band width but appropriate central frequency applied for extraction
Implications of setup
  All main ring magnets are current regulated
  Possible mismatches in energy  B-field regulation in near future [6]
  Quadrupoles, sextupoles and correctors  offsets in tune, chromaticity and momentum offset possible
  Hysteresis effects

## Slide 13: Developments for RFKO MEE
Implications of setup
  All main ring magnets are current regulated
  Possible mismatches in energy  B-field regulation in near future
  Quadrupoles, sextupoles and correctors  offsets in tune, chromaticity and momentum offset possible
  Hysteresis effects
Open questions
  Speed limitations
    Ramping of dipoles: definition of waveform (steps, round off)
    Repeated ramping of sextupole necessary?
  Sensitivity to mismatches
    Caused by hystersis effects?
  Suitable RFKO excitation

## Slide 14: Xsuite Simulations
Time dependent 6D particle tracking [7]
Beneficial for our studies
  MEE: ramping of elements
  RFKO: fast-changing exciter element
Special characteristics in PIMMS
  Radial regulation loop for off-momentum operation
  ZetaShift shifts particles i.r.t. synchronous particle

## Slide 15: System Implementation Tests: ΔE = 1 MeV
1 MeV ≈ 1 mm depth difference
* Adjusted values to visualize behavior over time.

## Slide 16: System Implementation Tests: ΔE = 1 MeV
RF frequency follows current of main dipoles
* Adjusted values to visualize behavior over time.

## Slide 17: System Implementation Tests: ΔE = 1 MeV
No losses during energy change
* Adjusted values to visualize behavior over time.

## Slide 18: System Implementation Tests: ΔE = 1 MeV
Ramp of resonant sextupole and gated application of RFKO signal
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 19: System Implementation Tests: ΔE = 1 MeV
Transfer line to irradiation room (IR) configured for second energy
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 20: System Implementation Tests: ΔE = 1 MeV
Transfer line configured for third energy
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 21: System Implementation Tests: ΔE = −20 MeV
Slow ramp (200 A/s) of main ring dipoles with low RF voltage
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 22: System Implementation Tests: ΔE = −20 MeV
Losses as sextupole ramps up
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 23: System Implementation Tests: ΔE = −20 MeV
Transfer line configured to first energy
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 24: System Implementation Tests: ΔE = −20 MeV
Transfer line configured to second energy
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 25: System Implementation Tests: ΔE = −20 MeV
Transfer line configured to second energy  Peak in IR
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 26: System Implementation Tests: ΔE = −20 MeV
Transfer line configured to second energy  Peak in IR
Mismatch in energy due to current regulation  low transport efficiency
RFKO signal on
RFKO signal off
* Adjusted values to visualize behavior over time.

## Slide 27: System Implementation Tests: ΔE = ± 4 MeV
Comparison of up- / downwards operation
RF voltage kept constant at 140 V

## Slide 28: System Implementation Tests: ΔE = ± 4 MeV
Losses in both up- and downwards operation
RF voltage kept constant at 140 V
* Adjusted values to visualize behavior over time.

## Slide 29: System Implementation Tests: ΔE = ± 4 MeV
RF voltage kept constant at 140 V
RF voltage ramped to ~ 300 V during energy change
As expected: No more losses!
* Adjusted values to visualize behavior over time.

## Slide 30: Longitudinal Tomography after Energy Change
Kept at high RF voltage (930 V) during energy change
Observed:
  Quadrupolar-mode oscillations
  Splitting of beam in beamlets

## Slide 31: Longitudinal Mismatches at Flat Top
If bunched at flat top (~ seconds)
  Coherent oscillations
  Observed with RF regulation loops on and off; different characteristics
  With loops:
    Mostly quadrupolar-mode oscillations
    Occasionally splitting of the beam in beamlets
  Investigation ongoing
    Dependence on chromaticity?
    WHY?

## Slide 32: 930 V
140 V
Longitudinal Tomography at Flat Top
[8] Longitudinal tomography

## Slide 33: Summary and Outlook
So far:
  Workflow for arbitrary energy steps established
  First exploratory measurements with new MEE and new RFKO exciter control system
  Compatibility with Xsuite simulations
Next steps:
  Finding cause of and mitigating coherent longitudinal oscillations
  Faster energy changes and effects on the spill
  Different RFKO excitations signals
Thanks to all contributors at TBU MedAustron and students at TU Wien!

## Slide 34: Summary and Outlook
So far:
  Workflow for arbitrary energy steps established
  First exploratory measurements with new MEE and new RFKO exciter control system
  Compatibility with Xsuite simulations
Next steps:
  Finding cause of and mitigating coherent longitudinal oscillations
  Faster energy changes and effects on the spill
  Different RFKO excitations signals
Thanks to all contributors at TBU MedAustron and students at TU Wien!
Thank you for your attention!
Questions and input are appreciated!

## Slide 35: References
[1] L. Badano, M. Benedikt, P. J. Bryant, M. Crescenti, P. Holy, A. T. Maier, M. Pullia, S. Rossi, P. Knaus. „Proton-Ion Medical Machine Study (PIMMS), 1,” CERN-TERA Foundation-MedAustron Oncology Collaboration, CERN (2000), https://cds.cern.ch/record/385378.
[2] F. Kühteubl. „Slow extraction developments at MedAustron,” 5th Slow Extraction Workshop, Wiener Neustadt, Austria (2024).
[3] F. Plassard. „Beam loss optimisation for slow extraction at MedAustron,“ 5th Slow Extraction Workshop, Wiener Neustadt, Austria (2024).
[4] F. Kühteubl. „Slow extraction optimisation for the MedAustron synchrotron,” Dissertation, Technische Universität Wien (2024), https://doi.org/10.34726/hss.2024.88260.
[5] M. Wolf. „Development of an RF system for slow extraction signal generation and measurement,” 5th Slow Extraction Workshop, Wiener Neustadt, Austria (2024).
[6] T. Margreiter. „High-Power Modular GaN Based Power Supply for MedAustron Scanning Magnets,” IPAC(2024).
[7] G. Iadarola, R. De Maria, S. Łopaciuk, A. Abramov, X. Buffat, D. Demetriadou, L. Deniau, P. Hermes, P. Kicsiny, P. Kruyt, A. Latina, L. Mether, K. Paraschou, G. Sterbini, F. F. Van Der Veken, P. Belanger, P. Niedermayer, D. Di Croce, T. Pieloni, L. Van Riesen-Haupt, M. Seidel. „Xsuite: An Integrated Beam Physics Simulation Framework,” JACoW HB2023 (2024), TUA2I1.
[8] S. Hancock. „ A simple algorithm for longitudinal phase space tomography,” CERN-PS-RF-Note-97-06 (1997).