## Slide 1: Status of rfko extraction development at medaustron
Fabien Plassard, E. Renner, K. Holzfeind, M. Wolf, T. Margreiter, M. Kausel, D. Prokopovich, M. Pivi, V. Rizzoglio, A. Wastl, C. Schmitzer

## Slide 2: OVERVIEW OF RFKO Implementation STUDIES
Current Slow extraction method at MA:
Betatron Core
Coasting beam at extraction
RF phase jump to obtain large and uniform momentum spread
Betatron core pushes the beam energy towards resonance condition via chromatic tune shift
Spill ripple correction with empty RF bucket Channelling

## Slide 3: OVERVIEW OF RFKO Implementation STUDIES
Current Slow extraction method at MA:
Betatron Core
Coasting beam at extraction
RF phase jump to obtain large and uniform momentum spread
Betatron core pushes the beam energy towards resonance condition via chromatic tune shift
Spill ripple correction with empty RF bucket Channelling
Future Slow extraction method at MA:
RF-Knock Out
Bunched beam at extraction
RFKO exciter coherently increase horizontal emittance till 3rd order resonant condition
Extracted flux quality depends on RFKO excitation signal
Compatible with MEE and DIC

## Slide 4: OVERVIEW OF RFKO Implementation STUDIES
Testing shifts with Schottky monitor as RFKO exciter
Synchrotron optics optimization
Exciter signal optimization
Spill intensity ripple and extracted beam properties
Spill quality studies
Testing of the RFKO cavity exciter
Adaptation of MA control system
Data workflow
Compatibility to MEE, DIC
Hardware and Control system
Safety and Impact analysis
Update of safety system (energy verification)
Risk analysis and testing
Regulatory compliancy
Synchrotron dipole B-field regulation
Pre-requisite project

## Slide 5: RFKO Hardware setup
RFKO exciter manufactured and tested on bench but not installed
Use of Schottky monitor plates for RFKO extraction testing
Use of the mTCA based LLRF crate for signal generation
  Compatible interface with our control system and intended use
  Same LLRF for Injector, Synchrotron, RFKO and beam diagnostics
  See “Development of an RF system for slow extraction signal generation and measurement” SlowEx 2024 workshop by Markus Wolf

## Slide 6: Beam extraction optimization
Pre-computed Amplitude Modulation for constant flux extraction
No amplitude feedback based on extracted dose implemented yet.

## Slide 7: Beam extraction optimization
LLRF loops during extraction
Voltage optimization: Reduce the amplitude of the energy offset by reducing the bucket height closer to the beam to contain the oscillations
Synchrotron optics optimized for RFKO (Tunes, chromaticity, Sext. Strength, mom. Offset)  see F. Kühteubl Thesis (“Slow extraction optimisation for the MedAustron synchrotron”)
See “Beam loss optimization for slow extraction at MedAustron”
(SlowEx 2024)

## Slide 8: Beam extraction optimization
Observation of strong intensity ripples on the first part of extracted beam
The effect is strong with proton and bunched beams

## Slide 9: Beam extraction optimization

## Slide 10: Beam extraction optimization

## Slide 11: Beam extraction optimization
Observation of strong intensity ripples on the first part of extracted beam
  First part of the spill shows strong frequency signature at twice the synchrotron frequency
  Contribution from the observed longitudinal quadrupolar oscillation at flat-top energy
  Collective effect contribution?
Ongoing investigation….
Longitudinal tomography p252MeV
Courtesy of K. Holzfeind

## Slide 12: Energy Verification with RFKO
Measurements were conducted to ensure that the energy of the delivered beam is within the intended limits
Establish a workflow for energy verification during clinical operation with RFKO:
  Define the limits of the relevant beam parameters and their correspondance to extracted energy and range
Courtesy of V. Rizzoglio and D. Skrijelj

## Slide 13: Risk Analysis of RFKO extraction
Measurements and simulations conducted to assess the risks associated with RFKO extraction
Test all relevant failure scenarios during beam extraction with RFKO
Example of RF-KO extraction simulation during single fault scenario
Courtesy of E. Renner
Courtesy of A. Wastl

## Slide 14: B-field regulation with RFKO
Ongoing implementation of B-field regulation at MA for the Synchrotron dipoles
Pre-requisite for the RFKO extraction:
  Long B-field stabilization time at FT energy in current-regulation mode
  Stable field at flat-top preferred with RFKO to prevent beam condition drift during extraction
First RFKO extractions in B-field regulation at MA
T. Margreiter
(“Magnetic field control in the MedAustron synchrotron” IPAC25)

## Slide 15: B-field regulation with RFKO
T. Margreiter
(“Magnetic field control in the MedAustron synchrotron” IPAC25)

## Slide 16: Conclusions
Progress in machine preparation for RFKO implementation:
  RFKO exciter manufactured
  New version of MA control system with compatibility to RFKO extraction
  Safety and RFKO impact analysis performed
Important issues need to be takled before the start of RFKO commissioning:
  Understand the beam instabilities observed before and during extraction
  Understand the mechanism of high intensity ripple at the start of the spill
Optimization of tailored RFKO signals for spill for proton and carbon beams for their entire energy range

## Slide 17: THANK YOU!