## Slide 1: STAR Time projection Chamber and Heavy Flavor Tracker (HFT)
Flemming Videbæk
Brookhaven National Lab
Notes: Some of the key detectors in STAR TPC since 200 HFT 2014-2016; intro

## Slide 2: The STAR Detector at RHIC
Barrel EM Calorimeter
FTPCs
Time Projection Chamber
Silicon Trackers
SVT & SSD
Endcap Calorimeter
Magnet
Coils
TPC Endcap & MWPC
Central Trigger Barrel & TOF
Beam Beam Counters
The central STAR detector with the TPC is the key element of STAR

## Slide 3: Principles of TPC Operation

## Slide 4: TPC: Principles of Operation
420 CM
Secondary electrons, resulting from a primary track, drift in parallel and uniform electric and magnetic fields inside the TPC. These electrons are ‘captured’ as a function of time by a 2D readout system on the endcaps to form a 3D image.
Voltage : - 28 kV at the central membrane 135 V/cm over 210 cm drift path

## Slide 5: The STAR Magnet (room temperature Aluminum coils)
A charged particle in a uniform field follows a (perfect) circular path

## Slide 6: 3D: The Time Projection Chamber
Time Projection Chamber  full 3-D track reconstruction
x-y from wires and segmented cathode of MWPC
z from drift time
in addition dE/dx information (charge of ionization)
Diffusion significantly reduced by B-field.
Requires precise knowledge of vD  LASER calibration + p,T corrections
Drift over long distances  very good gas
quality required
Space charge problem from positive ions, drifting back to central membrane  use a gated grid
Gate open
Gate closed
STAR TPC

## Slide 7: Sector Wire Geometry
Blown up picture of sensitive area on readout chambers

## Slide 8: Picture of Chambers and pad plane (iTPC)
The STAR TPC was upgraded in 2019 with improved chambers that increased acceptance and accuracy (Beam Energy Scan)
Here pad plane and wire planes during final assembly

## Slide 9: Actual installation
Unpacking and cover

## Slide 10: Electronics
Doubling #channels per FEE card. 2 SAMPA per FEE
55 FEEs per inner sector
16 FEEs per RD0
4 RDOs per inner sector
Pre-production RDO-- installed
iFFE
Readout Board

## Slide 11: Au on Au Event at CM Energy ~ 130 GeV*A
Data Taken June 25, 2000.

## Slide 12: Physics Goal Heavy Flavor
Heavy quarks at RHIC
Charm M~1280 MeV
Bottom/beauty M~ 4180MeV
Heavy quarks usually created in pairs c,c-bar b,b-bar
Forms mesons, baryons that decays rather quickly with few mm.

## Slide 13: STAR HFT Physics Motivation
Extend the measurement capabilities in the heavy flavor domain, good probe to QGP:
Direct topological reconstruction of charm hadrons (small c decays, e.g. D0  K )
Method: Resolve displaced vertices (~120 mm)
The STAR detector
HFT
@ RHIC
200 GeV Au+Au collisions @ RHIC
dNch/d ~ 700 in central events
Notes: SV
Highlight detecting short lived particles –
you have heard about the physics in previous talks

## Slide 14: How are measurements done?
Measure points on the tracks
Determine momentum by fitting to a circle (helix)
Project to primary vertex,
Determine decay point and angles.
From the momenta and masses of decay products calculate mass of decaying meson
Notes: Note to scale -

## Slide 15: Heavy Flavor Tracker (HFT)
TPC Volume
Magnet
Return Iron
Solenoid
Outer Field Cage
Inner Field Cage
EAST
WEST
FGT

## Slide 16: HFT was installed into STAR
SSD
IST
PXL
HFT

## Slide 17: Heavy Flavor Tracker (HFT)
SSD
IST
PXL
HFT
Detector | Radius (cm) | Hit Resolution R/ - Z (m - m) | Radiation length
SSD | 22 | 20 / 740 | 1% X0
IST | 14 | 170 / 1800 | <1.5 %X0
PIXEL | 8 | 12/ 12 | ~0.4 %X0
 | 2.5 | 12 / 12 | ~0.4% X0
SSD
Existing single layer detector, double side strips (electronic upgrade)
IST One layer of silicon strips along the beam direction (r-φ) , guiding tracks from the SSD to PIXEL detector.
PIXEL
two layers
18.4x18.4 m pixel pitch
10 sectors, delivering ultimate Pointing resolution that allows for direct topological identification of charm.
Monolithic active pixel sensors (MAPS) technology

## Slide 18
PXL Detector Mechanical Design
Mechanical support with kinematic mounts (insertion side)
Insertion from one side
2 layers
5 sectors / half (10 sectors total)
4 ladders / sector
Aluminum conductor Ladder Flex Cable
Ladder with 10 MAPS sensors (~ 2×2 cm each)
carbon fibre sector tubes (~ 200µm thick)
20 cm

## Slide 19: PXL Sensor
Monolithic Active Pixel Sensor technology
Ultimate-2: third generation sensor developed for the PXL detector by the PICSEL group of IPHC, Strasbourg
High resistivity p-epi layer
  Reduced charge collection time
  Improved radiation hardness
    20 to 90 kRad / year - 2*1011 to 1012 1MeV n eq/cm2
S/N ~ 30
MIP Signal ~ 1000 e-
928 rows * 960 columns = ~1M pixel
Rolling-shutter readout
  connects row by row to end-of-column discriminators
  185.6 ms integration time
  ~170 mW/cm² power dissipation
  Configurable via JTAG
  2 LVDS data outputs @ 160 MHz

## Slide 20: PXL Hit Position Resolution
Ultimate-2 sensor geometry
  pixel size: 20.7 m X 20.7 m ~6 m geometrical resolution
  3-pixel av. cluster size ~3.7 m resolution on center-of-mass
Position stability
  Vibration at air cooling full flow: ~5 m RMS
  Stable displacement at full air flow: ~30 m
  Stable displacement at power on: ~5 m
Global hit resolution: Dx ~ 6.2 m
Sector in the metrology setup
Metrology survey
  3D pixel positions fully mapped and related to kinematic mounts
PXL insertion
Kinematic mounts
Novel insertion approach
  Inserted along rails and locked into a kinematic mount inside the support structure
  Capability to fully replace PXL within 12 hour
Dx ~ 6.2 m
r1 = 2.8 cm
r2 = 8 cm
HFT DCA pointing resolution:
( 10  24/p) m

## Slide 21
After assembling sectors in a half shell sector tooling balls measured with touch probe relative to kinematic mount coordinate frame
Since the shells are supported the same way in the CMM and in the STAR installation the relative pixel position mapping is not disturbed
duplicate, truncated PXL support tube with kinematic mounts

## Slide 22: PXL Material Budget
Thinned Sensor
  50 mm
  0.068% X0
Flex Cable
  Aluminum-Kapton
  two 32 mm-thick Al layers
  0.128% X0
    Copper version  0.232% X0
Carbon fiber supports
  125 mm stiffener
  250 mm sector tube
  0.193% X0
Cooling
  Air cooling: negligible contribution
Total material budget on inner layer: 0.388% X0
  (0.492% X0 for the Cu conductor version)
HFT DCA pointing resolution:
( 10  24/p) m
Curved sensor
40-60% yield after thinning, dicing and probe testing

## Slide 23: Intermediate Si Tracker
Details of wire bonding
24 ladders, liquid cooling
Prototype Ladder
S:N > 20:1
>99.9% live and functioning channels

## Slide 24: Silicon Strip Detector (SSD)
The ladders and Si-sensors was from existing detector
Upgrade readout system with new ladder cards on detector, RDO cards, and cooling system
20 Ladders
Ladder Cards
Add pictures

## Slide 25: HFT performance
Pointing resolution was determined in Au+Au collisions at 200 GeV in transverse plane and longitudinal. Measurements are in agreement with expectations.
Important also to take into account the vertex resolution vs. centrality, particular for the most peripheral collisions

## Slide 26: D0 spectra
We have achieved very good significance for D0
Low pT is difficult due to the pileup and miss matched hit points
Notes: Of course bot d0 and d0-bar

## Slide 27: Physics of the Heavy Flavor Tracker at STAR
Direct HF hadron measurements (p+p and Au+Au)
  (1) Heavy-quark cross sections: D0±*, DS, ΛC , B, …
  (2) Both spectra (RAA, RCP) and v2 in a wide pT region: 0.5 - 10 GeV/c
  (3) Charm hadron correlation functions, heavy flavor jets
  (4) Full spectrum of the heavy quark hadron decay electrons
Physics
  (1) Measure heavy-quark hadron v2, heavy-quark collectivity, to study the medium properties e.g. light-quark thermalization
  (2) Measure heavy-quark energy loss to study pQCD in hot/dense medium e.g. energy loss mechanism
  (3) Analyze hadro-chemistry including heavy flavors
Notes: Need to go through formatting

## Slide 28: Timeline
R&D 2003- LBL with IHPC, Strasbourg for MAPS sensors
2008 Project start ; 2010-2014 construction
2013 Commissioning run
2014-16 Physics
  14: AuAu
  15: pp
  16: AuAu dAu
2017 Removed from STAR and in storage
2014-2023 Physics analysis and papers

## Slide 29: Next Generation MAPS
For the upgrade of the ALICE detector in the current RUN-3 at LHC, the collaboration developed a next generation MAPS sensors
  Faster readout time (~5 micro sec)
  Better radiation hardness
  Pixels size 30*30 micron
See e.g. https://arxiv.org/pdf/2111.08301.pdf

## Slide 30: MVTX - sPHENIX Inner Most Detector for HF physics
MVTX parameters: L = 271 mm
Service Barrel (SB)
Cylindrical Shell
Structure (CYSS)
End-Wheels (EW);
- 3 layers
3-layer sensor barrel
- 48 staves, 432 chips
R (mm) | min | mid | max
Layer 0 | 24.61 | 25.23 | 27.93
Layer 1 | 31.98 | 33.35 | 36.25
Layer 2 | 39.93 | 41.48 | 44.26
North
🡪
North
New BP
Carb. Fib. Comp./CFC = EW, CYSS, SB
Total Weight: ~24 kg
(sensor part: ~3 kg)
Notes: A quick reminder of the MVTX overall design -

## Slide 31: sPHENIX MVTX Photos
Viewed from South right before the installation: One can clearly see the service lines inside the Service Barrel (SB):
Blue: data cable, Green: Cooling air, Red/Black/White: LV power
Viewed from North after installation
MVTX North EndWheel

## Slide 32
Pixel size: 28 um x 27um
1024 columns (~3cm)
512 rows
(~1.5cm)

## Slide 33
North-end view of MVTX
half-detector @BNL 2-88
First outer layer assembled

## Slide 34
MAPS detector has matured a lot, and next generations with be used at eRHIC in EPIC detector and at another upgrade to ALICE
TPC are used at collider experiments and also in neutrino physics experiments
More questions?

## Slide 35: ideas to explore
What are the applications for MAPS in current and planned experiments?
Radiation damage is important to understand at RHIC, LHC, and space. How are this being investigated?
Where does radiation have a positive impact?
Where has Nuclear Physics impacted society with inventions and applications?

## Slide 36: Info on PXL detector mechanics and construction
NIM Article
  arXiv 1710.02176
  NIM A97(2018) 60
Howard Wieman
  HFT PXL mechanics
  Forum on Tracking Detector Mechanics
  30 June-2 July 2014 at DESY, Hamburg
  https://indico.cern.ch/event/287285/contributions/1640694/attachments/534386/736809/PXL_mechanics.pdf

## Slide 37
Spatial mapping of the pixels
Pixel locations determined with CMM equipment to within 10 m prior to installation in STAR
Programmed CMM Measurement method:
All pixels located on a sector with respect to 3 sector tooling balls
  2 Lithography points on the chips measured with optical head
  Chip surface profile measured with 11 x 11 point pattern using a Feather Probe. Using a touch probe permits picking up over hung surfaces
Full sector measurement takes ~ 8 hours
PXL sensor surface profile from survey: ± 30 m > PXL hit error,
and the chip to chip surface deviation along the ladder surface is still larger, but all of this is then corrected with the spatial map
Parametriuzed by 5*5 spline fct

## Slide 38
Final operating pointing performance in STAR
Cosmic ray result
Excellent half to half pointing, sub 30 m
But after half to half alignment which was required to correct poor reproducibility of kinematic mount seating
Check alignment with very low luminosity AuAu also
X
X
X
X
X
X

## Slide 39: D0 reconstruction
Use topological variables
Use particle ID

## Slide 40: Grid Leak Walls checked
The narrow end uses a simple brass wall to terminate the array of wires
  Ground Potential
  No bias required
  Anode wires > 10 mm away
The wide end uses a G10 PCB board to terminate the array
  Inside wall is grounded to the strongback.
    Anode wires > 10 mm away
  Outside wall is used to suck ions off the Anode wires of the outer sector
    No shield wall on outer
    So requires ~750 volts (DC) to suck ions across the gap
    Outer Sector Anode wires > 2 cm away
    3 Pads & 3 voltages (-115 V, ~750 V & ground)
Courtesy Jim Thomas

## Slide 41: Some words on efficiency
We have applied two methods to evaluate efficiencies
Fast simulation data driven to replicate multi dimensional basic observables
Geant embedding with displaced geometries
Methods agrees well

## Slide 42: Inner Detector Support (IDS)
MSC
Pixel Insertion Tube
Pixel Support Tube
IDS
East Support Cylinder
Outer Support Cylinder
West Support Cylinder
PIT
PST
ESC
OSC
WSC
Shrouds
Middle Support Cylinder
Inner Detector Support
Carbon Fiber Structures provide support for 3 inner detector systems and FGT.
All systems are highly integrated into IDS.
Installed for Run-12
MSC installed for Run-13 for PXL

## Slide 43: D0 spectra
Precise measurements extends to low pT and non-central collisions. Data from 2014
Results consistent with 2010/11 TPC only analysis
Data from submitted paper: arXiv 1812.10224

## Slide 44: D0 RAA
RAA < 1 for 0-10% for all pT
Suppression at high pT increases for more central collisions
Same trends as at LHC and for light mesons
Notes: Add references

## Slide 45: D0 RCP
Significant suppression at high pT
Matches theoretical calculations
D0-bar/D0 slightly greater than one
  Possible due to slight baryon asymmetry at RHIC

## Slide 46: D0 cross section
Total D0 cross section is nearly independent of centrality, and smaller than in p+p. For pT > 4 GeV/c it decreases with centrality

## Slide 47: Λc reconstruction
More than 50% improvement in signal significance with TMVA methods
Includes 2016 dataset

## Slide 48: Λc/D0 centrality dependence
Λc enhancement increases towards more central Au+Au collisions
Large contribution to total charm cross sections in HI

## Slide 49: Λc/D0 pT dependence
Significant enhancement of Λc/D0 compared to p+p.
Comparable with light flavor B-to-M ratios
Consistent with charm quark hadronization via coalescence
Notes: Ko PRC79(2009) 044905.
Greco Eur.Phys.JC(2018) 78:348
SHM PRC79(2009( 044905

## Slide 50: Strange Charmed quarks

## Slide 51: Total charm cross section
D0 yields measured down to pT=0 GeV/c
For D+-, Ds Levy fits to measured spectra for extrapolation
For Λc three model fits to data are used and included in systematics
Total charm cross section is consistent with p+p values
Redistributed on meson and baryons

## Slide 52: Ds, D0 enhancement
Strong enhancement of Ds/D0 observed in central A+A w.r.t fragmentation baseline
->Strangeness enhancement and coalescence
Enhancement larger than model, particular at high pT

## Slide 53: Non prompt D0, B-decay
Strong interaction of charm with the medium.
What about bottom?
RAA of non-prompt D0 extracted.
Improved signal significance using BDT
Will get results form combined data set

## Slide 54: B study from different channels

## Slide 55: D0 Elliptic Flow
Published D0 v2 from 2014 dataset
Clear mass ordering for pT < 2 GeV/c
Follows NCQ scaling in mid-central (10-40%) collisions
Phys.Rev.Lett 118, 212301 (2017)

## Slide 56: D0 Elliptic Flow
D0 v2 measurements extended to 0-10% centrality with combined 2014 and 2016 data set
Significant flow observed

## Slide 59: D0 directed flow (v1)
First evidence for non-zero v1 for D’s
Much larger than K+-K-
May imply effect of strong EM fields in collision system

## Slide 62: Summary
The HFT upgrade was overall successfully and has hopefully lead the way for future high resolution vertex detectors (ALICE ITS, sPHENIX, EIC)
Strong Modification of charm hadron spectra and hadro chemistry in AA collisions as witnessed by the different observations
  Total charm cross section is conserved.
  Substantial energy loss , coalescence
  Gain significant flow & may have achieved thermal equilibrium in medium (v2)
  Observed non-zero directed flow
Indication for suppression at high pT for B->D0 ,… measurements
Thanks to the many STAR collaborators that has worked on HFT construction and analysis over the years

## Slide 63
The STAR Upgrades and BES Phase II
iTPC Upgrade:
Rebuilds the inner sectors of the TPC
Continuous Spatial Coverage
Improves dE/dx
Extends h coverage from 1.0 to 1.5
Lowers pT cut-in from 125 MeV/c to 60 MeV/c
EPD Upgrade:
Allows a better and independent reaction plane measurement critical to BES physics
Improves trigger
EndCap TOF Upgrade:
PID at h = 1.1 to 1.5
Provided by CBM-FAIR Phase-0
Major improvements for BES-II
Endcap TOF
TPC
i
E

## Slide 64: All detectors now in place for BES II