## Slide 1: OB stave update
Work from all the UK groups, material from Marcello Borri, James Glover, Sam Henry, Adam Huddart, Jian Liu from recent UK F2F meeting and otherwise

## Slide 2: OB Stave Construction
¼ Length Stave End
QSP
L4 X 8 Modules
L3 X 4 Modules
Notes: We form a skeleton structure with a central carbon fibre I-beam and K9 foam cross braces
K9 foam blocks are located to align with the Ancillary ASIC and LEC of the modules for enhanced heat transfer.
Serial power chain
  Flexible Printed Circuit (FPC) is structurally integral to the stave, is part of the gas containment boundary and supplied power and data transfer.

## Slide 3: Bonding trials
Wire bonds:
  Sensor to module carrier/bridge FPC
  AncASIC to module carrier/bridge FPC
TAB bridge FPC to main FPC
Bonding tests (using prototype FPCs from LTU)
Optimisation ongoing
TAB:
Lessons learned:
Alignment important (to 50 µm)
Reducing step height important
Different aluminium foils
No major difference in bonding behaviour
Required parameter adjustment beyond standard settings
Foil bond strength showed higher variation compared to gold board
Back support critical
Wire bonding:

## Slide 4: Modules
Started to think about assembly of module prototypes
  Will start with flat modules
Some questions:
  Automated vs manual assembly?
  What glue to use (will start with Araldite, but need to watch viscosity
  Thinking of getting some dummy Si with interconnection pads

## Slide 5: Where we are with stave prototypes
Not where we wanted to be
  Still only quarter L4 stave prototypes built (3 off)
  We are tantalizingly close to building our first L3 stave (for some time)
    Waiting for mould tooling required for full length stave (held up in Oxford Physics finance department)
Of the quarter L4 stave prototypes one has been equipped with heaters (on a 50 µm steel shim) and is used to shake down our setup
  Air flow (working on control of air flow)
  Heating
    Our first set of heaters don’t have the right power distribution - redesigning
  External temperature measurement (TC and thermal imaging)
  Internal (air flow) temperature measurement (TC)
  Vibrations (capacitive displacement sensors)

## Slide 6: Pictures
TI camera
Air in
Air out
Capacitive displacement sensors
Front view
Rear view

## Slide 7: Heaters
1.8 W/quarter stave, 30 l/min air in lower channel
Design
Measured

## Slide 8: Some preliminary results - Vibrations
Only for indication – need full length L4 stave
Vibrations
  fs = 5 kHz, fl = 68.1 l/min, T = 10 s
  dRMS1 = 4.2 µm, dRMS1 = 4.2 µm
  , for longitudinal (beam) modes
  , for transverse (torsional) modes
, for longitudinal (beam) modes
, for transverse (torsional) modes

## Slide 9: Some preliminary results - Thermal
Only for indication – need full length L4 stave and correct heater geometry
Convection Coefficients
  Natural convection - 7 W/m2 K from external surfaces
  Forced convection - 48.44 W/m2 K from internal surfaces
  Forced convection - 650 W/m2 K from K9 adjacent surfaces
Effect of K9 foam

## Slide 10: First thoughts on stave mounting
Start to think about stave mounting
  Sequence:
    Outer support cone
    L4 staves mounted
      RDOs dressed on outside of cone
      Air connections
    Test L4
    Inner support cone
    L3 staves mounted
      RDOs dressed on outside of cone
      Air connections
    Test L3
Support Cone
Support Cone
RDO
RDO
RDO
RDO

## Slide 11: Stave Mounting
Currently fastener access from outer side, these could only be accessed through cutouts in support cone
  No space beyond the end module for screw holes
  Looking at options for underside/cone end mounting
As modelled:
  Coolant flows from one end of the stave (co-current flow),
    Sacrificing floating end could allow for counterflow with an equivalent manifold block on both sides
  Common mounting face for structural connection and coolant inlet
  Coolant channel integrated into mounting manifold
Fixed/Pinned for thermal expansions

## Slide 12: Air supply first thoughts
Staves manifold with a grouping of 6 staves
6 manifolds per half
QTY 12 - 10 mm ID
Ambient
  Pressure drop – 0.7 Bar/m
    Supply may be 2-3 bar
  Flow velocity 300 m/s
    Near sonic  much larger feed pipes required
Pressurised
  Supply Pressure - 8 Bar
  Velocity 38 m/s
  Pressure drop – 0.08 Bar/m

## Slide 13: Software
Status
  Full OB geometry created from CAD files
  SVTOB_UK branch: https://github.com/eic/epic/tree/SVTOB_UK available to all, with scripts to convert CAD to GDML, README
  All significant bugs, issues identified so far now understood and fixed
  In use by Oxford and Birmingham team
  Hit maps and material scans for latest CAD files plotted
Next steps
  Update standard epic geometry to better reflect material thickness and active area, perhaps with curved staves
  Optimise full version to improve speed
  Benchmarking scripts to monitor performance
Hit maps (106 muons) for single stave
L3
L4
Material scan L4 only

## Slide 14: Further material

## Slide 16: ¼ Length Stave Prototypes
Preformed components
  Carbon Fibre I-beam pre-cured in a 2-part mould
  Flexible Printed Circuits (FPC) preformed (with heat and pressure) into C channel

## Slide 17: Stave Inner Structure
⌀ 3 mm Silver Steel rods
Aluminium Internal Formers
K9 Foam Heatsink

## Slide 18: Outer Skin
Airtight boundary formed by FPC, Carbon Fibre face sheet and modules
FPCs Mounted along outer edge, joined to K9 Foam with Prepreg Adhesive Film
Carbon fibre face sheets aligned on inner structure
Silicone rubber sheet added on top and bottom surfaces as an intensifier

## Slide 19: Post-Cured Stave Assembly

## Slide 20: Pressure Loss Contributions
Chart: Pressure (Bar)
| 0 | 27.5 | 39 | 129.85 | 141.35 | 242.4 | 253.9 | 267.45
 | 0.14553 | 0.02164 | 0.01764 | 0.014876 | 0.010876 | 0.0081118 | 0.0041118 | 0
Subassemblies used to isolate contribution
59 ltr/minute pressure drop contributions
  Inlet - 0.124 bar
  K9 Foam – 400 Pa
  Outlet – 400 Pa
Scaled for full L4 stave
  Pressure at first K9 foam block
    59 lpm  0.062 Bar
    Scaled to 110 lpm  0.207 bar
0.0216 Bar
Inlet Block
Inlet & Outlet Blocks
Inlet, Outlet & K9