## Slide 1: Simulations for Detector Cooling Requirements
Girish Gowda
EIC User Group & ePIC Joint Collaboration Meeting

## Slide 2: Introduction – Detector Cooling
Cooling Performance Considerations :
Temperature Stability
Minimising thermal noise
Preventing Radiation damage
Cooling Design Considerations :
Heat Load
Material Cost
Maintenance & Reliability

## Slide 3: ePIC Detector – Cooling Requirements

## Slide 4: ePIC Detector – Cooling Requirements

## Slide 5: Forward EMCAL
Sectional View
fECal

## Slide 6: Forward EMCAL
Sector Geometry
Exploded View
FEB
SiPM
Acrylic
Epoxy
Tungsten Scifi
SiPM Board

## Slide 7: Forward EMCAL – CFD Simulation
FEB Heat load (W) | SiPM Heat load (W)
4.5 | 0.17
CFD Data :
Technical Data :
Half FEB Heat load (W) | SiPM Heat load (W)
2.25 | 0.17
Materials
Copper thickness (10 %) in FEB
Copper thickness (5%) in SiPM Board
FEB and SiPM Board thickness : 2 mm
Simulation
Copper tubes, heat sink & thermal pad geometry not considered
Cooling Requirements:
Maintain FEB temperature 3 to 5 C above dew point
Cooling the SiPMs to be determined

## Slide 8: Forward EMCAL – CFD Simulation
SiPM Temperature
FEB Temperature
Ambient Temperature (C) | Half FEB Heat load (W) | SiPM Heat load (W)
20 | 2.25 | 0.17
Heat Transfer Coefficient (W/m^2 C) | FEB Temperature (C) | SiPM Temperature (C)
5 | 48.0 | 24.1
1 | 143.8 | 30.8

## Slide 9: Forward EMCAL
Sector Geometry
Exploded View
Copper tube (inch)
0.25
FEB
SiPM
Acrylic
Epoxy
Tungsten Scifi
Heat Sink
Thermal Pad
Copper tube
SiPM Board

## Slide 10: Forward EMCAL – CFD Simulation
Water flow rate (litre/min) | Flow velocity (m/s) | Tube Length (m) | Pressure drop (psi) | Temperature rise (C) | FEB Heat load (W) | SiPM Heat load (W)
0.83 | 0.76 | 6 | 2.1 | 2 | 4.5 | 0.17
CFD Data :
Technical Data :
Water flow rate (litre/min) | Flow velocity (m/s) | Tube Length (m) | Pressure drop (psi) | Temperature rise (C) | Half FEB Heat load (W) | SiPM Heat load (W)
0.83 | 0.76 | 0.21 | 0.07 | 0.04 | 2.25 | 0.17
Materials
Copper thickness (10 %) in FEB
Copper thickness (5%) in SiPM Board
FEB and SiPM Board thickness : 2 mm
Simulation
Copper tubes, heat sink & thermal pad geometry considered
Cooling Requirements:
Maintain FEB temperature 3 to 5 C above dew point
Cooling the SiPMs to be determined

## Slide 11: Forward EMCAL – CFD Simulation
SiPM Temperature
FEB Temperature
Ambient Temperature (C) | Flow velocity (m/s) | Half FEB Heat load (W) | SiPM Heat load (W)
20 | 0.76 | 2.25 | 0.17
Heat Transfer Coefficient (W/m^2 C) | Water Inlet Temperature (C) | FEB Temperature (C) | SiPM Temperature (C)
5 | 22 | 25.9 | 23.8

## Slide 12: Forward EMCAL – CFD Simulation
Heat Transfer Coefficient :
Ambient Temperature (C) | Flow velocity (m/s) | Half FEB Heat load (W) | SiPM Heat load (W)
20 | 0.76 | 2.25 | 0.17
Heat Transfer Coefficient (W/m^2 C) | Water Inlet Temperature (C) | FEB Temperature (C) | SiPM Temperature (C)
5 | 22 | 25.9 | 23.8
1 | 22 | 27.51 | 28.19
FEB Temperature
SiPM Temperature

## Slide 13: Forward EMCAL – CFD Simulation
Relative Humidity :
SiPM Temperature
FEB Temperature
Ambient Temperature (C) | Flow velocity (m/s) | Half FEB Heat load (W) | SiPM Heat load (W)
20 | 0.76 | 2.25 | 0.17
Relative Humidity (%) | Dew Point Temperature (C) | Water Inlet Temperature (C) | FEB Temperature (C) | SiPM Temperature (C)
100 | 20 | 22 | 25.9 | 23.8
60 | 12 | 14 | 20.22 | 22.85

## Slide 14: Forward EMCAL – CFD Simulation
Design :
SiPM Temperature
FEB
SiPM Board
FEB
SiPM Board
SiPM Temperature

## Slide 15: Summary
Suggestions :
Modifying geometry from straight tube to serpentine tube will enhance heat transfer and reduce FEB temperature further.
Adding fins or turbulators in heat sink will enhance heat transfer and aid in cooling of FEB.

## Slide 16: Barrel EMCAL – Cooling
Computational design of BIC with 48 sectors, showing one of the sectors in detail
University of Manitoba – Thermal Analysis of BIC single sector

## Slide 17: Barrel EMCAL – Cooling
Computational design of the one of the sectors of the BIC
Material distribution for the sector

## Slide 18: Barrel EMCAL – Cooling
Thermal simulation
Boundary condition :

## Slide 19: Barrel EMCAL – Cooling
Baseline design of Cold Plate
Thermal simulation with volume flow rate of 2gal/m

## Slide 20: Barrel EMCAL – Cooling
New design of Cold Plate

## Slide 21: Barrel EMCAL – Cooling
Thermal simulation Results

## Slide 22: Barrel EMCAL – Cooling

## Slide 23: Question & Comments