## Slide 1: Electronics & DAQ WG
Fernando Barbosa (JLab)
25 September 2025

## Slide 2
  Today’s Topics:
  Recent Tests on bPOL DC-DC converters – Gerard, Norbert
  dRICH Power Distribution - Robert
  Need to:
    Develop Power distribution diagrams per sub-detector
      Realizing synergies and finalizing designs
      Update power infrastructure needs and quantities
        bPOL quantities (bPOL48/12, controller ASIC and modules)
        Chassis, modules (LV, BIAS, HV)
        Cabling
  Upcoming meetings:
  29 September 2025 – TIC: concurrence on the 32/36 ch CALOROC choice
Welcome & News

## Slide 3
  Comments from reviewers related to power:
  1) Power Budgeting - Clearly define the expected total power consumption of the detector, including all subsystems, to support infrastructure planning and risk mitigation.
  2) Power Distribution Tree - Provide a detailed power-tree diagram showing power flow from sources to endpoints, including estimated losses at each stage.
  4) Wire Gauges & IR Drops - Specify wire gauge selections and analyze voltage drops (IR drops) across all power lines to ensure reliable operation under load. Examine the effect of IR drops during powered steady state operation when unexpected clock drop out occurs on detector mounted electronics including DCDC converters and ASICs.
  5) Power-Up/Down Sequences - Document safe power-up and power-down sequences to protect sensitive electronics and maintain system integrity.
  Develop power distribution system diagrams from main supply to point-of-use, with wire gauges, fusing, IR drops, power budgeting – per sub-system/sub-detector.
  Refine cooling requirements for input to infrastructure and integration.
  Develop document on power sequencing and behavior analysis from possible failure modes.
Further notes on Power Diagrams from IPDR3