## Slide 1: OB module prototyping Prepared with input from Birmingham, Daresbury, Liverpool, and Oxford
19 December 2025
UK EIC SVT Working Meeting, University of Oxford
Jian Liu
on behalf of the OB Module Team

## Slide 2: Module concept
Currently working on prototyping for flat module assembly
Adapting this concept to a curved stave would require
  Updates to the assembly jigs to match the curvature
  Fabrication of the module in a curved geometry
Module description
2 LAS
2 AncASIC
1 FPC
Glue Si dies to FPC
Wire bonding Si dies to FPC
Updates covered today
Interconnection
Prototyping jigs
Dummy silicon parts

## Slide 3: Interconnections
Wire bonds
  Sensor  module Bridge-FPC
  AncASIC  module Bridge-FPC
spTAB
  Bridge-FPC  Main-FPC

## Slide 4: spTAB – 1st trials
Initial setup
  First-generation FPCs, PCBs, and jigs
  Early optimisation of bonding parameters
  Initial mounting procedure development
Goal
  Establish feasibility of spTAB bonding
  Identify critical mechanical and process limitations
PCB differential lines
FPC differential lines
Errors identified
  Differential lines did not align as expected
    FPC mounted upside down
    Aluminium foil stretched more than designed
    Bonding performed on top foil layer instead of bottom
    Consequence
      Foils pushed close to mechanical limits
      Significant damage to aluminium traces
Key findings
  FPC–PCB alignment is critical (~50 µm tolerance)
  Step height between FPC and PCB must be minimised
Bond trials (here)
Initial PCB mounting (here and here)

## Slide 5: spTAB with updated design
Updated jig from
  Updated jig ensures the FPC fits only in the correct orientation
Updated PCBs
  Solder resist removed beneath the FPC
    Minimises height difference between FPC aluminium layers and PCB pads

## Slide 6: spTAB – 2nd trials
Assembly approach
  Nylon stand-offs used as weights to improve FPC alignment
  UV-curing glue dots used to hold FPC flat during bonding
  Additional glue applied along FPC edges after bonding
Bonding successful, no visible damage, and basic electrical continuity confirmed

## Slide 7: Main-FPC
The assembled M-FPC is currently under testing (see the report presentated on Monday)

## Slide 8: Wire-bonding on the foils
Two types of foils from LTU received in March (details see the presentation by LTU)
Equipment used in Liverpool
  Hesse BJ820 Wire Bonder
  Dage 4000 Pull Tester

## Slide 9: Wire-bonding configurations
Substrates
Gold boards (reference)
Single-layer and multi-layer aluminium foils
Mounting methods
Taped flat (no vacuum)
Vacuum-mounted directly to jig
Vacuum-mounted to foils glued on FR4 PCB
Vacuum-mounted using a diffuser
Cleaning
Safewash (default)
Acetone
Methanol
Wires
CCC (used only in initial tests): Al-1%Si, 25 µm diameter, EL % 1-4, TS 15-18g
Heraeus (default): Al Si-M, 25 µm diameter, EL > 1%, BL 15-17 cN
Heraeus H74-41 (softer): Al 25µm, AlSi-S, EL 1,0-4,0%, BL 14-16g
Gold board
Taped foil
Foils on PCB
Foils on jig
Foils with diffuser

## Slide 10: Bonding parameter
Standard settings
  Ultrasonic: 22%
  Bond force: 22 cN
  Deformation: 40%
  Overtravel: 25 µm
25 µm wire with 100 µm spacing, 1500 µm bond length → ~30° pull angle
Two diffuser types
  Thick (0.225 mm) and Thin (0.068 mm)
A gold reference board was bonded to establish baseline pull strength and standard deviation
Gold board

## Slide 11: Wire-bonding failures
Bond force or ultrasonic power below 22 resulted in high failure rates
Thick diffuser on single-layer foils
  Dark rings observed around many bond feet
  Indicates insufficient local support
  Suggests suboptimal vacuum contact
Thin diffuser
  Significantly fewer wire failures
  Dark regions around bond feet disappeared
  Residual shadow on single-layer foil
    Indicates incomplete vacuum flattening
Maintaining flatness is challenging for single-layer foils due to flexibility

## Slide 12: Pull test matrix with the thinner diffuser
Thin diffuser provided a significant improvement over the thick diffuser
Single-layer foils
  Failures still observed at higher parameter settings
Multi-layer foils
  Performance remained good
Results validated with repeated testing
Initial test
Repeated test
Detailed results can be found here.

## Slide 13: Wire-bonding comparison
50% Deformation
Better bonding quality when foils were glued to the PCB
The thinner diffuser significantly improved bonding quality, though some stability challenges remain
Multi-layer foils consistently offered easier and more reliable bonding, regardless of method
The tests demonstrated that acceptable mean strength and standard deviation values can be achieved
15 µm Overtravel

## Slide 14: Cleaning
Cleaning solvent
  Safewash (PCB cleaning agent)
  Acetone (recommended by LTU)
  Methanol (used for some ATLAS module frames)
A soft toothbrush was used to gently clean the surface and remove residues such as fingerprints
Standard PCB cleaner previously used is unsuitable due to foil fragility
Delamination was observed in the multilayer foil during the cleaning process!
  Most likely caused by the drying process with the compressed air gun  should be avoided

## Slide 15: Cleaning comparison
Acetone
Methanol
Safewash
Safewash appears to be the most effective cleaning method
No significant difference is observed before and after cleaning with acetone or methanol
Note
  The observed impact depends on the initial condition of the foil, and the tests were performed over several months
Bonding parameters with the thin diffuser (0.068mm)
Overtravel 25 µm
Deformation 50%
Ultrasonic % - 25%
Bondforce 25 CN

## Slide 16: Wire comparison
50% Deformation
15 µm Overtravel
Heraeus (default, older): Al Si-M, 25 µm diameter, EL > 1%, BL 15-17 cN
Heraeus H74-41 (softer, new): Al 25 µm, AlSi-S, EL 1,0-4,0%, BL 14-16g
Tests performed using the thin diffuser (0.068mm)
No major difference between the default and the softer wire

## Slide 17: Prototyping jigs
FPC jig
LAS jig
Design completed in October

## Slide 18: Prototyping jigs – mechanical draws
Tolerancing
Flatness: 10 µm
Surface finish: Ra=0.8 µm
The drawings and step files are with the Liverpool workshop
Some tolerances are still under discussion
The manufacturing takes about 20 - 25 days
Expected to start in January
Deploy for assembly  March onwards

## Slide 19: Pure dummy silicon
Pure silicon dummy sensors prepared at CERN  removed from backing foil  shipped to UK
Packaging
  Membrane boxes
  Gel-paks
No breakages observed
Modules moved within membrane boxes
Removal from gel-paks using vacuum jig and suction cups was successful
  Silicon is flexible but does not break
  Using more than one suction cup would improve stability

## Slide 20: Dummy silicon with pads – design
Processes from the supplier
  Procure Silicon wafer 625 µm
  Photomask fabrication
  ~1 µm Al patterning and etching
  Backgrinding to 50 µm and 300 µm respectively
  Dicing & Chip tray packing
  Shipping
5RSU dimensions
Thickness: 50 µm
Dicing Requirements
  Length: 114.330 mm
  Width: 19.567 mm
AncASIC dimensions
Thickness: 300 µm
Dicing Requirements
  Length: 15 mm
  Width: 1.5 mm
Metalization
Material: Aluminum
Thickness: around 1 µm
Pads
  Pad size: 91 µm × 144 µm
  Pad pitch: 100 µm

## Slide 21: Dummy silicon with pads – fabrication
The supplier is completing the drawings
Expected delivery date: by the end of February
AncASIC: 500-600 chips
5RSU LAS: ~50 chips

## Slide 22: Summary
Flat module prototyping is currently ongoing
Interconnection
  spTAB-bonded M-FPC prototype is under testing
  Wire-bonding results show that acceptable mean strength and standard deviation can be achieved with LTU foils
  Further bonding optimisations to be done with more realistic designs and final-grade components
Manufacturing of the prototype jigs ongoing
Pure dummy silicon sensors are available; dummy LAS and AncASIC are expected in January/February
Assembly practice is expected to start in March

## Slide 23: Backup

## Slide 24: Packaging
Nanosystems JP packaging details
  Dummy LAS: “given the three different chip sizes, instead of fabricating a custom case from scratch, we propose a standard method where we will remove the chips from the dicing ring and pack them using a cushioning system as shown below. The pink layer will serve as a cushion, and a black spacer will secure the chips between the layers. This is a standard packaging method we frequently use.”
  Dummy AncASIC: “a similar-sized chip tray (15 × 1.5 mm) is available, which we will use for packaging.”

## Slide 25: Standard bonding parameters

## Slide 26: Module assembly