## Slide 1: SLDO ASIC Test System
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD
Presentation Outline
User Manual Preparation.
Test and Commissioning of the test system.
Conclusion & Initial Chip Tests.

## Slide 2: SLDO Test System Operating Instructions
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 3: Communication
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 4: Commands
Command | Function
0x12 1 | CONTROL_4 ON
0x12 2 | CONTROL_4 OFF
0x13 1 | CONTROL_3 ON
0x13 2 | CONTROL_3 OFF
0x14 1 | CONTROL_2 ON
0x14 2 | CONTROL_2 OFF
0x15 1 | CONTROL_1 ON
0x15 2 | CONTROL_1 OFF
0x16 1 | CONTROL_0 ON
0x16 2 | CONTROL_0 OFF
Command | Function
0x07 1 | DIS_LATCH_Q_N ON
0x07 2 | DIS_LATCH_Q_N OFF
0x08 1 | SEL_EXT_DISABLE ON
0x08 2 | SEL_EXT_DISABLE OFF
0x09 1 | EXT_DISABLE ON
0x09 2 | EXT_DISABLE OFF
0x10 1 | ACTIVATE_OCP_N ON
0x10 2 | ACTIVATE_OCP_N OFF
0x11 1 | RESET_OCP_N ON
0x11 2 | RESET_OCP_N OFF
Command | Function
0x01 1 | Relay 1 ON 40mA Load
0x01 2 | Relay 1 OFF
0x02 1 | Relay 2 ON 500mA Load
0x02 2 | Relay 2 OFF
0x03 1 | Relay 3 ON 900mA Load
0x03 2 | Relay 3 OFF
0x04 1 | Relay 4 ON 100nF CAP
0x04 2 | Relay 4 OFF
0x05 1 | Relay 5 ON 1uF CAP
0x05 2 | Relay 5 OFF
0x06 1 | Relay 6 ON 10uF CAP
0x06 2 | Relay 6 OFF
0x23 1 | Relay 7 ON 10nF CAP
0x23 2 | Relay 7 OFF
Over Current Load. 0x24 1 | Relay 8 ON 1.6A Load
0x24 2 | Relay 8 OFF
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 5: R_IN Value
Load on ASIC Output | Board | Value to send To Rheostat | Resistance Required
500mA | 1 | 1285 | 5.213K
900mA | 1 | 1167 | 2.906K
40mA | 1 | 1594+ (Fit 47K to PL2) | 58K
500mA | 2 | 1285 | 5.22K
900mA | 2 | 1167 | 2.911K
40mA | 2 | 1594 +(Fit 47K to PL2) | 58.02K
500mA | 3 | 1285 | 5.22K
900mA | 3 | 1167 | 2.909K
40mA | 3 | 1594 | 58K
500mA | 4 | 1285 | 5.219K
900mA | 4 | 1167 | 2.911K
40mA | 4 | 1594 + (fit 47K to PL2) | 58K
500mA | 5 | 1285 | 5.218K
900mA | 5 | 1167 | 2.916K
40mA | 5 | 1594 + (fit 47K to PL2) | 58K
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 6: Program Update
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 7: Just In Case
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 8: Test and Commissioning
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 9: Configuration and Initial Tests
DL559815A Initial configuration: -
Board 1 IP address set to 192.168.1.100
Board 2 IP address set to 192.168.1.102
Board 3 IP address set to 192.168.1.101
Board 4 IP address set to 192.168.1.104
Board 5 IP address set to 192.168.1.105
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 10: Capacitance Load Test
Andrew Hill
STFC Daresbury Laboratory
Keckwick Lane
Daresbury
Warrington
Cheshire
WA4 4AD

## Slide 11: Resistive Load Tests

## Slide 12: Temperature Tests
A voltage of 18.5V was applied to the Carrier Board Connector SK5.
The current was limited to 0.45A.
As the temperature rose:-
110.6 Ω was measured.
123.4 Ω was measured.
To achieve 105º C we will need to use a hot plate in conjunction with the heater resistors.

## Slide 13: Conclusion