## Slide 1: Scintillating fiber quality control at anl
June 18th 2026
Noémie pilleux

## Slide 2: Scintillating fiber reception
Receiving 28000 fibers/month until the beginning of October
247,253 x 4.55m = 1,125,000 m of fibers
Received 5 crates from 2 production batches
90 days for QC after reception
Goal would be testing 280 a month (1%)
Currently sampling ~50 a month but scaling up

## Slide 3: Quality control parameters
Dimensions
4.54 < length < 4.56 m
0.99 < diameter < 1.01 mm
Attenuation length
>4m
Batch-to-batch variation < 10%
Light yield
From a Sr-90 source
Larger than 3.5pe with a standard PMT
Batch-to-batch variation <15%
Emission spectrum
  Blue/green

## Slide 4: Measurement setup
Several iterations of the setup inspired by our colleagues at UoS and UofR, thanks to Tom and Kevin
Seoyun Jang
Fibers in a groove in a plate inside a dark box
  SiPM readout, can be switched for photo-spectrometer
LED measured at specific location
→ attenuation
LY measurement can be done in the same setup
  Both ends polished, one blackened
LED
Fiber optics bringing LED directly to the tested fiber

## Slide 5: Fiber polishing

## Slide 6: quality control results (1)
Dimensions are within specs:
  Fibers are all really close to 4.54 m long
  Diameter close to 0.99 mm
  Skewed towards lower-end of the specs.
Attenuation lengths are all longer than 4m
  Shape of the short component still being assessed
  Cross-check with colleagues at UofR confirming this likely comes from our setup.

## Slide 7: Quality control results (2)
LY assessed qualitatively, as we only have a “weak” source, low statistics, for now:
  ~5-8pe at 200cm
  above 20 pe at 20cm
  with SiPM readout (factor 0.5-0.6 on a PMT).
All emission spectra are very consistent

## Slide 8: MEP early career group effort
Tons of tests thanks to amazing students and postdocs filling a full week of tests!

## Slide 9: Future steps
Intense Sr-90 source for LY measurements
Replicating tests with intense LED + photodiode to address current questions about the short component of the attenuation
Duplication and move of the setup
2 students in July, 1 student August-December, 1 student in the fall will help with testing and systematic checks

## Slide 10: Attenuation length – double exponential fits

## Slide 11: Attenuation length – fiber to fiber comparisons
Fit to the average from all fibers
Fit fixing the attenuation length to the QC limit