## Slide 1: ePIC Electronics Summary of ASIC Internal Reviews
Fernando Barbosa (Jlab)
L3 Manager WBS 6.03.01.07
Joint ePIC/EICUG Meeting
11-17 July 2026
Glasgow
Chief Electrical Engineer

## Slide 2: Outline
ePIC Channels, Scale & Technology
ASIC Review Goals
ASIC Reviews:
  EICROC
  ALCOR
  SALSA
  CALOROC
  FCFD
Post-Review Notes
Summary

## Slide 3: ePIC Channels, Scale & Technology
Initiated as R&D, now as PED:
Discrete COTS – SiPM, 14-bit
CALOROC – SiPM, 10-bit
ALCOR – SiPM, 1 pe
EICROC – AC-LGAD, pixel
FCFD – AC-LGAD, strip
FCFD Variant – HRPPD/MCP-PMT
SALSA - MPGD
eRD109/PED
eRD109/PED
eRD109/PED
CERN
KIT, NASA (GSFC), ANL
U. Glasgow
Electronics Readout benefits from large synergies among 26 systems, reduced to:
  One (1) Discrete COTS
  Six (6) ASICs

## Slide 4: ASIC Review Goals
Comprehensive assessment of the ASICs from design to production:
  Specifications & Requirements - frontend, processing, interfaces, power and packaging (including form factors where applicable).
  Digital Interfaces - commands & controls
  Status & Steps to Completion - with schedule and to include testing/validation with detectors
  Production Outlook - to include final production review, production timeline and QA/QC testing
  Documentation - Functionality (e.g. User’s Manual/Spec sheets) and QA/QC Plans, and with planned availability for FEB final designs and Engineering Articles, bump/wire bonding.
ASIC | Date | Contributors | Institutions
EICROC | 1 April 2026 | Christophe de la Taille Alexandre Soulier | OMEGA/CNRS/IN2P3
ALCOR | 22 April 2026 | Fabio Cossio G. Dellacasa | INFN Torino
SALSA | 6 May 2026 | Damien Neyret | CEA Saclay IRFU U. of Sao Paulo
CALOROC | 13 May 2026 | Frederic Dulucq El Berni Mowafak Pedro Dumas | OMEGA/CNRS/IN2P3
FCFD | 20 May 2026 | Artur Apresyan Datao Gong Sergey Los | FNAL

## Slide 5: EICROC Review – 130 nm/AC-LGAD pixel
Preliminary measurements of the EICROC1 are very encouraging, after a work-around was implemented due to a design issue with readout selection inverters. A spare wafer will be modified to fix the inverters, and to provide about 60 chips to be available for tests with sensors/detectors in July 2026.
Specifications for readout and rates from all EICROC detectors will be provided to sub-system developers soon.
Test Boards will be available soon, requiring further coordination with BNL and Tonko.
OMEGA/CNRS/IN2P3 will be responsible for the production of the EICROC ASICs and to be provided as In-Kind contribution to ePIC. QA/QC of the EICROC ASICs will be the responsibility of, and performed by, ePIC; OMEGA/IJCLab will provide documentation from the recently delivered ALTIROC ASIC, for guidance.
EICROC2 is the production version.
Version |  | Status | Notes
EICROC0A | 4x4 | Under test | Improved testability
EICROC0B | 4x4 | Under test | Peak and Wilkinson ADC
EICROC1 | 32x32 | Under test | inverter issue
EICROC1, rev | 32x32 | Submitted, available July 2026 | revised metal layer
EICROC2 | 32x32 | Submission Q1 2027 | Design with Digital on Top, high rate, derandomizer, zero-suppress

## Slide 6: ALCOR Review – 110 nm/SiPM
ALCOR-64 (V3) has increased amplifier bandwidth for better intrinsic resolution of ~84 ps and discriminators have programmable hysteresis.
Slew Rate (SR) mode can resolve pulses ~2 ns apart.
Maximum rates of ~1 Gb/s per four FEBs well under the ASIC and RDO limits.
Thermal simulations are on-going to ensure reliable operation of the readout modules and keeping die temperatures reasonable at <65 °C.
QA/QC plans are being developed for ALCOR ASICs, FEBs/RDOs and PDUs (full module).
BGA packaging using UBM process from the current vendor is not robust, with shorts and opens. MPW with alternative process needed, and interfacing with IMEC.
Production likely delayed to Q1-Q2 2028.
Version |  | Status | Notes
ALCOR-32 | 32 ch | Complete | System tests continue
ALCOR-64 | 64 ch | Under test | UBM process packaging not robust
ALCOR-64, rev MPW | 64 ch | Submission Q3-Q4 2026 | New MPW with alternative packaging process
ALCOR-64 | 64 ch | Submission Q1-Q2 2028 | Production

## Slide 7: SALSA Review – 65 nm/MPGD
ESD input protection to be determined by each of the MPGD sub-detectors and depending on operational conditions.
Power has increased to 25-30 mW/ch from original estimate of 15 mW/ch.
Readout will use lpGBT ASICs.
SALSA1 Ch 1 frontend version selected with improved pole-zero, anti-saturation and lower noise (<0.5 fC below 350 pF).
12-bit SAR ADC from USP selected, with 11-bit ENOB.
Same BGA package and pinout for SALSA2 and SALSA3.
Preliminary documentation available to developers.
SALSA3 is final version submission.
Production in Summer 2029, with MPGD FEBs in Summer 2030.
Version |  | Status | Notes
SALSA1 | 4 ch | Complete | Tested with MPGDs, TID tests at CERN next
SALSA2 | 32 ch | Submission Q4 2026 | Most IP blocks included
SALSA3 | 64 ch | Submission Q1-Q2 2028 | Full functionality
SALSA | 64 ch | Submission Q3 2029 | Production of SALSA3

## Slide 8: CALOROC Review – 130 nm/SiPM
CALOROC1A and 1B received for tests: 1A has frontend similar to H2GCROC with TOT and 1B has gain switching. Testing 1B first for large sensor capacitance.
Two output lpGBT links (ch0-17 and ch18-35). Hit rate per channel ~100 kHz (4 samples).
Noise of 4.1 ADC counts from digital activity and likely from PCB layout and grounding. New PCB to be developed.
Preliminary gain uniformity and linearity look reasonable.
Power is 12 mW/ch, and as expected.
Fast commands and documentation well developed with QA/QC document complete. Datasheet version 1.4 available to users with pinout common to all CALOROCs.
Production yield of 70% expected, multiple causes.
CALOROC3 is production version.
Version |  | Status | Notes
CALOROC1A/B | 36 ch | Under test | Testing 1B first
CALOROC2 | 36 ch | Submission Q1 2027 | Version selected for final production and availability for system detector tests. First production quantities
CALOROC3 | 36 ch | Q1 2028 | Final production quantities

## Slide 9: FCFD Review – 65 nm/AC-LGAD strip
CFD uses a 50% fraction.
Jitter is 10 ps @ 20 fC.
Noise floor is reasonable at 1.5 fC or 1 LSB.
Rate requirement (guidance) of 10 Hz per channel – needs further refinement from bTOF detector.
Time resolution of ~40 ps (sensor + FCFD); Spacial resolution of 15-20 um.
Power of 6.2 mW/ch.
Resolution is dominated by sensor and 5-bit ADC suffices. Higher resolution ADC possible but with limited improvements.
One lpGBT channel per 128 FCFD channels for readout with daisy chain readout and multi-drop clock distribution.
Data format is 128 bits per hit, with shared header and trailer for multiple hits and link of 320 Mbps.
FCFD3.0 is production version.
Version |  | Status | Notes
FCFD1.1 | 6 ch | Completed | Frontend only. Multiple tests with strip sensors from Hamamatsu
FCFD1.2 | 6 ch | Submitted April 2026 | Full functionality: Discriminator, TDC, ADC, I2C, simplified readout. Tests starting in July 2026
FCFD2.0 | 32 ch | Submission Q2 2027 | Full functionality
FCFD3.0 | 32 ch | Submission Q2 2028 | Production

## Slide 10: Post-Review Notes
The FCFD Variant ASIC, for HRPPDs, is currently under development and it is based on the FCFD ASIC. Results from preliminary tests of the FCFD1.1 with a HRPPD prototype are very encouraging.
Recent developments (June 2026) regarding the EICROC and CALOROC ASICs:
      Per TSMC, no new submissions for 130 nm after 1 October 2026.
      Options: port designs to 65 nm/TSMC or 110 nm/UMC or …
      Concerns: validation of designs, radiation hardness and higher costs.
      CNRS is looking into options and funding.
      Impact: delayed schedules.
      A detailed plan will be available soon.
      Any new details will be communicated during the next two presentations by Pedro and Frederic.

## Slide 11: Summary
Considerable progress towards final ASIC designs has been realized. However, two issues have recently surfaced:
  ALCOR – BGA packaging using UBM process needs to be resolved.
  EICROC and CALOROC – porting the designs to a technology node other than 130 nm. It is noted that validations of the designs continue with dice from recent submissions in 130 nm.
ASICs continue to be tested during development, with multiple sources and their respective sensors (AC-LGAD pixel, AC-LGAD strip, SiPM, MPGD, HRPPD).
ASIC production QA/QC plans follow the CALOROC’s as a template. QA/QC of the EICROC, FCFD and FCFD Variant ASICs will require contributions from the ePIC collaboration and the project.