## Slide 1: SiPM Characterization
Henry Bilton

## Slide 2: Summary
Expose the SiPM to 0-8 photons and measure the output voltage. Putting the voltages on a histogram yields several spikes for integer photons
By fitting the curve with exponentially modified normal distributions, we can determine the single photon response.
By doing this for for several input voltages, we can fully characterize the SiPM

## Slide 3: 540 SiPM
Curve fitting for 540 SiPM at 55.5 volts
From above we can determine the breakdown voltage is 51.9V

## Slide 4: 541 SiPM
Curve fitting for 541 at 55.5 V
The breakdown voltage for this one is also 51.9 V

## Slide 5: 546 SiPM
Curve Fitting for 546 at 55.5 V
The breakdown voltage for this one is significantly higher at 52.5 V

## Slide 6: 547 SiPM
Curve Fitting for 547 at 55.5 V
The breakdown voltage for this one is in between at 52.2 volts

## Slide 7: 548 SiPM
Curve Fitting for 546 at 55.5 V
The breakdown voltage for this one is in between at 52.2 volts

## Slide 8: 535 SiPM
This SiPM is much noisier than all previous SiPMs
The 535 has a breakdown voltage of 51.7 volts

## Slide 9: Comparing 540 and 541 SiPMs
541 has slightly tighter peaks, indicating less noise and better resolution
Other SiPMs cannot be directly compared since they have different gains.
Orange: 541, Blue: 540

## Slide 10: Conclusion
Best: 546, 547, 548
Adequate: 540, 541
Subpar: 535, 536

## Slide 11: Comparing 547 and 548 SiPMs
Both curves have similar resolution and steepness.
The differences are likely due to differences in setup geometry.
Orange: 548, Blue: 547