## Slide 1: sPhenix Testing 1/29/18
C. Schultheiss, P. Rosas, Z. Altinbas
K. Yip, P. Orfin, B. Van Kuik
And a host of others.

## Slide 2: Major Events
Lost Voltage Tap 11
  Reconfigured Quench Detector
PS Compensation
  Stable on Magnet, Large Current Error, Needs More Work.
Ramped to 500 Amps twice
  First in steps at 0.5 Amps/sec
    Fast Discharge occurred at 469 Amps
  Ramped at 1.0 Amp/sec
    Descended via Slow Discharge

## Slide 3: Voltage Taps

## Slide 4: Voltage Tap VT11 Remakes Connection

## Slide 5: Ramps During Testing

## Slide 6: Ramp Used For Inductance Calculation

## Slide 7: Magnet Voltage for Total Inductance Calculation

## Slide 8: Magnet Voltage for Outer Coil Inductance Calculation

## Slide 9: Magnet Voltage for Inner Coil Inductance Calculation

## Slide 10: Inductance Calculation
Calculations from the 1 Amp/sec ramp
  Total Inductance 2.215 H.
  Outer Coil Inductance 1.112 H.
  Inner Coil Inductance 1.081 H.
  Delta from total -0.022 H. approx. -1%

## Slide 11: Fast Discharge Event

## Slide 12: Fast Discharge Quench Detector Overview

## Slide 13: Fast Discharge Expanded View

## Slide 14: Calculated Coil Difference and Quench Signal

## Slide 15: Calculated Coil Difference and Quench Signal Expanded View

## Slide 16: Quench Detector Response

## Slide 17: Fast Discharge Time Constant

## Slide 18: Energy Extraction Resistor
Assume measured L and tau are correct, then R=L/tau and 2.215/35.42 = 0.0625; subtract cable R of 666 uOhms and Energy Dump Resistor is 0.0619 Ohms.

## Slide 19: Conclusions
Fast Discharge caused by PS Voltage Upset
  Low amplitude 286 Hz oscillation
Measured Inductances are reasonable, better estimates later with higher ramp rates.
Slow Discharge required about 1341 seconds.
Fast Discharge required about 200 seconds.

## Slide 20: Issues
Issues still to be addressed:
  PS Compensation
  Difference between PS and Magnet Current
    At Flattop
      PS Current 775.71 Amps
      Mag Current 496.24 Amps
    Slope Error
  Offset in PS Voltage
    Appears to be -5.71 Volts
    At 9:58:09
      Mag Volt, PS Volt, Mag Curr, PS Curr are all zero
    At 9:59:50
      PS Volt jumps down to -5.71 Volts

## Slide 21: Offsets on Turn On

## Slide 22: PS and Magnet Currents and Voltages

## Slide 23: Issues (continued)
Time Displacement between Fast Logger Boards A and B
  Outer Magnet Coil and Inner Magnet Coil on Logger Board A
  Total Coil Voltage on Logger Board B
  Board A lags Board B by about 80 ms

## Slide 24: Logger Board Time Lag