## Slide 1: sPhenix Test Issues
Carl Schultheiss
2/28/2018

## Slide 2: sPhenix Issues
Front-End Electronics
  Dual Gain Channels
    Saturation at high gain
      Change time constant of absolute value circuit
        Capacitor value change
    Saturation at low gain
      Increase attenuation
        Resistor value change

## Slide 3: Logger Issues
Logger
  Time skew between channels
    Modified software improved skew, but didn’t eliminate it.
    Add a third A/D board and send the Quench Signal to all three boards to use for time alignment.
  Add PS setpoint signal to logger channels.

## Slide 4: Quench Detector Issues
PXIe hardware is not adequate for operations.
  Reads and Writes require a lot of overhead, this uses too much time for real-time systems.
    Prevented the dual-gain channels from being implemented in the quench detector.
    Required the channels be processed in blocks of twelve, this prevented the error counter (integrator) from being linear.

## Slide 5: Quench Detector Issues (continued)
  The data files from a quench can’t be sent automatically to the control system.
    Requires someone log into the quench detector and move them from the real-time hard drive to the Windows hard drive so they can be shipped to someone.
    This is not acceptable for operations.
  Replace the PXIe hardware.
Two voltage taps were intermittant during testing, VT10 and VT11.
  Repair if possible.

## Slide 6: Cryo Issues
Cooling of the Gas Cooled leads seemed marginal, but short running time.
  Insure slightly more gas flow for permanent installation.
Only one working temperature sensor in the magnet.
  Repair the others if possible.