## Slide 1
Meeting page: https://indico.bnl.gov/conferenceDisplay.py?confId=4307
Making sense of Strain Gauges (a little analysis)
Now, as the High Field Test is over, there are things we need to think about …
  Keeping vacuum at least until the Magnet is warmed up … but do we keep on vacuuming after that ?
  Do we want to loosen the jack ?
   Carl Schultheiss’ slides
    Remember to purchase better flow controller at the Current Leads that can deal with lower pressure.
    Do we want to think about fixing VT10/VT11 ?

## Slide 2
Trying to make sense of strain gauge …
On Feb. 5, soon after the intentional fast discharge at 2000 A. Most strain gauges were screwed ( volt/resistance close to 0 ) as there was a short developed.
This was not the first time (and each time it could be a different channel). Per Paul’s suggestion, we switched to just measuring the resistance across the strain gauge (2-wire), without using the constant current source to measure voltage/resistance (4-wire).
  We unplugged the current (flow) wires to avoid the short.
  But in this method, we’re sensitive to the cable/wire resistance between strain gauge and the readout. I actually asked Winston to put a hack (subtracting a 7.5  for all connections) to make the strain measured not too big.
  In any case, our thinking was that since we’re interested in the change of strain between 0 A to top current, we only care about its relative change during the ramp, assuming the temperature fluctuation during the ramp was minimal.

## Slide 4
Stress = strain  Young’s modulus in the linear-elastic region
  Young’s modulus ~ 200 GPa or ~29007 ksi for Inconel 718
  Yield ~1.1 GPa and Tensile Strength (Ultimate) ~ 1.375 GPa
Force ~ stress   (Tie-Rod)
  typically  (Tie-Rod) ~ 200 mm2 from BaBar documents
Anything that I wrote below can easily have 10% uncertainty.
Feb. 5, 2018 (when we still had the 4-wire implementation --- more precise/absolute) : ramping to 2000 A
  ~Biggest change @ strain-gauge 19 (TS21)
  See picture (next page)
  -0.00003  0.00005 ~0.00008  stress ~16 MPa

## Slide 5
Feb. 5: TS21 (more precise with constant current source)
Magnet Current

## Slide 6
Feb. 13, 2018 (when we only had the 2-wire implementation to measure the resistance --- less precise and only good for relative measurement) : ramping to 4830 A
  `
Feb. 13: TS21 (only good for relative measurement)
Magnet Current

## Slide 7
Feb. 13, 2018 (only good for relative measurements) :
TS21: 0.0016876  0.0021046  change of ~0.00042  stress ~83 MPa
At 2000A : 0.0016876  0.0017602  change of ~0.000073
  which is not too different from ~0.00008 measured on Feb. 5, 2018.
Force  stress and Force  I2 -------- *
0.00042(10%) / 0.000073(10%) ~ 5.75  0.81
(4830 A/2000 A)2 ~ 5.83
  The above-mentioned relations (*) are verified 
TS21 is axial.
How about radial ? For a good example,  TS08 (See next page)

## Slide 8
TS08 (only good for relative measurement)
Magnet Current
Also on Feb. 13, 2018, similarly for TS08
  TS08: 0.000290  0.000155  change of ~0.000445  ~89 MPa