## Slide 1: Microphonics control studies for high QL SRF Cavities Module commissioning of the bERLinPro SRF Gun
A. Neumann
Helmholtz-Zentrum
Berlin
Second Topical Workshop on Cryomodule Microphonics and Resonance Control, Brooklyn, NY

## Slide 2
Systems under study
CW modified TESLA cavity for BESSY-FEL (2005-2008)
  Studied mainly at HTS (HoBiCaT)
  QL from 5.106-2.108, Eacc=8-22 MV/m
  CW studies, heat load, microphonics compensation high QL LLRF
HoBiCaT
cryomodule
TTF-III coupler
SC solenoid
Cavity
Cathode
2-phase line
Drive laser
port
YAG-Screen
Faraday Cup
Slit
ICT
Stripline
YAG-Screens
Bending magnet
Beam dump+
Faraday Cup
Beam dump+
Faraday Cup
Nb Pb lead cathode DESY gun (1.6 cell), 2011-2012
  Studied with diagnostics beamline
  E0=12-25 MV/m
  QL=3.106-1.5.107
  Df/DP=100 Hz/mbar Df/DE0²=1 Hz/(MV/m)²
bERLinPro ERL medium power SRF gun cavity cold string
DESY-JLab-HZB coll.
SRF Photoinjector (ERL), today
  Studied with diagnostics beamline + cathode transfer
  E0=8-35 MV/m
  QL=1.107-2.8.107
  Df/DP=33 Hz/mbar Df/DE0²=3.4 Hz/(MV/m)²

## Slide 3
Microphonics: How does it appear?
2.5-3 mm Niobium walls
Deterministic,
narrow-band
sources:
Vacuum pumps
Stochastic
background
noise
Field amplitude variation:
Dynamic Lorentz force, Df/DEacc² = 1-3Hz/(MV/m)²  Ponderomotive instability
G. Bissofi
222 Hz
151 Hz
Response of the
Cavity-Helium vessel-Tuner
system:
Mechanical Eigenmodes
Mechanical oscillations of the Cavity:
Microphonics
Helium pressure fluctuations
Df/Dp = 50-60 Hz/mbar, SRF Gun:30-100 Hz/mbar
16 mbar±30 µbar
Is that the whole story?
Lowest modes:
Usually transverse
Heat transport dynamics in LHe
Thermo-acoustic oscillations (cryo valves…)
G.P. Gelata et al. International Journal of Thermal Sciences
Volume 58, August 2012, Pages 1-8

## Slide 4
Microphonics: What do we see?
The cavity is the best sensor of mechanical excitation itself, true?  What deformation affects actually the RF mode (TM0yz different than e.g. TM1yz)? Cavity design: Cavity sensitivities
Helium bath often main driver of microphonics, usually the static term Df/DP can be measured or the cavity design optimized for  What about dynamic response?
Detuning is often obtained by comparing forward and transmitted (reflected) wave of the TM010 p-mode (or FPC excited RF mode).  We only see what affects the RF mode, not all oscillations, do we care? Cavity oscillations w.r.t. beam motion?
Is this always what we consider microphonics (oscillations in acoustic regime)?
Can we compensate every oscillation which affects the wanted cavity RF mode using tuner with e.g. piezos?
Be aware of other contributions to appear in the signal:
  Transient beam-loading (hopefully repetitive, but what about beam losses in recirculating machines?)  beam arrival jitter, synchrotron oscillations
  Multipacting in more special cavities (e.g. SRF gun, coaxial parts)
  Loop oscillates if stability criterion is not met
  Coaxial FPC: Oscillation of inner conductor (cooling media)
   Some of them alter the measurement, some are real detuning, but eventually not tackled by tuners

## Slide 5: Detuning characterization of a TESLA cavity, short-term
sf = 1.56 Hz
SFFT
FFT
HoBiCaT: sf = 1 - 5 Hz (rms)  2-13° phase error (Aim: 10-2 °) „open loop“ „closed loop“
He pressure variations: fmod < 1 Hz
Cavity specific: First mode at 20 - 50 Hz
Spectrum can appear more populated  depends on cavity system
Excited Eigenmode
He pressure-
variations
Detuning (Hz)
Detuning (Hz)
Time (s)
Characterisation: Measurement results
For elliptical cavity, usually transverse mode

## Slide 6: Microphonics and thermal load on cavity tank
Simultaneous 2 cavity operation (TESLA cavities)
Heaters attached at each tank
Monitor microphonics due to thermal load on cavities
Onset of non-laminar Helium flow?
30mbar
Chimney limit

## Slide 7: Correlation of Helium pressure and detuning of cavity
Open loop measurement of cavity frequency and He pressure
50-60 Hz/mbar down to resolution limit of pressure meters (~1Hz)
Evidence that main contribution of microphonics mediated through superfluid Helium
TESLA cavity

## Slide 8: Longterm stability: Peak events how much, how often?
Microphonics recorded at HoBiCaT with TESLA cavity for 48 hours at Eacc=8MV/m
RMS Values around 1-5 Hz  Determines field stability and thermal loading of RF system (5 kW)
Peak values extend out to 17 σ!  Determines RF power installation (15 kW)
Peak events occur 10-20 times a day! (This was partly improved by changes to the control settings of the under-press. pumps.)
Expected field stability: 0.02 - 0.1°
For „comfort“ want to reduce the microphonics
Gaussian sub-range
0.8 Hz rms
Excited 1st mech. resonance: Transverse mode

## Slide 9: Time-frequency analysis by Wavelets
y
^
t/s
0
y
s. w
Morlet
Variation up to Df =10 Hz on a ~100 ms time scale
Spectrum of He-pressure variations of stochastic nature
Adaptive, „learning“ (dynamic) compensation mandatory
Need for classic feedback control
Df (Hz)

## Slide 10: Tuner qualification
Tested 3(4) different
tuner systems
 | Saclay I* | Saclay II | INFN Blade**
Mech. principle | 1-lever+flexures | 2-levers+flexures | Knee-lever+blades
Tuning resolution | 0.176Hz/step | 0.09 Hz/step | 2.6 Hz/step
Drive | Phytron / HD 1:88 | Phytron / HD 1:88 | Sanyo / PG 1:100
Max remanence | 30 Hz | 55 Hz | 380 Hz
Coarse tuning range | 750 kHz | 500 kHz | 720 kHz
Coercitive steps | 180 (no backlash) | 350-500 (backlash) | 100 (backlash)
Used piezo type | HV (0-1000V) | LV (-10-150V) | LV (0-200V)
Piezo tuning range | 750 Hz | 1420 Hz | 800 Hz
Group delay (df/dw) | 290 µs | 150 µs | 650 µs (138 µs)***
Lowest resonance | 40 Hz | 40 Hz (double) | 35 Hz
* Increased stiffness of piezo holder frame
** Several versions exist
*** 138 µs for 1.4 cell SRF gun with Cornell blade tuner
 Important for CW piezo based detuning control

## Slide 11: Model based controller: Fit of the system
Relevant for tuning
Fit: Parallel acting 2nd order systems
Evaluate response of higher modes at lower frequencies
>20 modes needed for fit
Systems complexity complicates use of model based feedbacks
(e.g. Kalman filter)
Transfer function as
look-up table or
Kalman approach tested
with cavity simulator
See talk A. Ushakov
TESLA Cavity
Individual for:
Each cavity tuner combination
Changes with coarse tuning or helium pressure
Affected by cav. mech. design
But: Only LF and piezo transfer functions accesible in operation!
Degeneration of 1st mode

## Slide 12: A tested scheme: Least-mean-square based adaptive feedforward
FFT
∙H-1
IFFT
t (s)
DU (V)
Compensating
signal
Dl (nm)
t (s)
External mechanical  Mostly by helium system, unknown!
oscillations
t (s)
Df (Hz)
Detuning
of the cavity
Calculation
of optimal
FIR filter
parameters
For white noise excitation
The FIR filter would be
H-1piezoDf
A. Neumann et al, Phys. Rev. ST Accel. Beams 13, 082001
or IIR

## Slide 13: Piezo based compensation results
SFFT
Open
loop
sf = 2.52 Hz
Feedback only
sf = 0.89 Hz
Feedback
and Feed-
forward
sf = 0.36 Hz
Results:
QL=6.4.107
Multi-resonance control:
Piezo resolution seems
to limit control of
neighboring modes  transfer of energy
Single-resonance control:
14 deg phase mod. down to 2 deg  sub-promille stability in theory possible
Resonances: Control voltage of mV regime required before amplifier

## Slide 14: LLRF studies with U Cornell: Limits of QL
TESLA cavity
Eacc= 10-12 MV/m
Tbath= 1.8 K
PI piezo loop Integrator vs. LF detuning
8/9-p filter optimized
sf = 5-10 Hz, Dfpeak= 15-25 Hz
LF detuning  IOT beam instable Cavity field trip
f1/2=3.25 Hz
log(sf)
Areas with sf>0.1 were blanked out
QL | sf (deg.)
5.107 | 0.008°
1.108 | 0.0093°
2.108 | 0.0236°
1st MEM
f1/2=13 Hz
f1/2=6.5 Hz
Pf=325W
at 12 MV/m

## Slide 15: SRF Gun for bERLinPro: Stability issues
Thermal short  high static losses (20 W) trial to cool via filling line, no phase separator, thus flash gas lead to bubble formation beating the cavity up to 3 kHz (PLL-mode)!
SRF gun cavities have high sensitivity to Lorentz force, up to 3 kHz tuning for target field required, higher probability of ponderomotive instabilities
Several cooling media attached to cavity and ancillaries: Cooling of normal conducting cathode via 80K helium gas  Vibration of cathode? Would act as a plunger modulating the TM010-p mode (depending on cathode position)
80 K cooling of HOM absorber….
Capacitive sensors for cathode positioning, BW 5kHz

## Slide 16
Some example during SRF gun operation
This instability was affected by DC bias voltage in cathode channel
This is used to mitigate Multipacting
Strong correlation with vacuum activity
Piezo in lowpass PI loop
Detuning
Multipacting?
No direct correlation to loop gain!
RF Phase
Forward Power
sf=9 Hz
sf  ∞
Poisson result:
Courtesy J. Völker
DC bias simulation
Choke
Half-cell

## Slide 17
SRF gun cavity LLRF operation
QL=2.83.107
Dfpeak=20 Hz sf=9.8 Hz 
sf=0.03 deg.
sA/A=1.5.10-4
To be controlled!
Piezo tuner transfer function
Phase spectrum
Detuning spectrum
Stability paid by 4xpower overhead

## Slide 18
SRF gun cavity LLRF operation: Limits
@25 MV/m, quench
would occur after few minutes
Insufficient
cooling of cathode
A lot of power dissipated in LHe bath  effect on microphonics?
First MEM
at 220 Hz
1st MEM
Helium bath activity?
DESY
Nb/Pb gun

## Slide 19
Message to be delivered
Low beam-loaded high QL operated multi-cell SRF elliptical cavities can be operate up to QL of 2.108 with stability below 0.02°  For better stability 5-7.107
Microphonics compensation can gain an order of magnitude and thus lower thermal load via FPC
Major contributions by excited mechanical eigenmode, often lowest transverse mode
Excitation most probably transferred via helium system
Special cavities like SRF guns demand for higher level of tuning control as they are more susceptible by design (half-cell)
Operation at high losses or close to quench limit will open up new surprises, higher level of microphonics
Future studies
Apply Kalman (A. Ushakov) and LMS feedforward control to SRF cavities as SRF gun, Booster 2-cell and Linac multi-cell + collaborate with DESY LLRF team at CMTB
Develop tuning strategies and firmware for high current and transient beam-loading cases (see talk P. Echevarria)

## Slide 20
Questions?
Thanks to all collaborators and partners of the past, present and future projects and co-workers at HZB:
Further readings:
A. Neumann, W. Anders, O. Kugeler, J. Knobloch (2010). “Analysis and active compensation of microphonics in continuous wave narrow-bandwidth superconducting cavities“, Phys. Rev. ST Accel. Beams 13, 082001.
O. Kugeler, A. Neumann, W. Anders, J. Knobloch (2010). “Adapting TESLA technology for future cw light sources using HoBiCaT”, Rev. Sci. Inst. 81 (7).
A. Ushakov, P. Echevarria, A. Neumann Developing Kalman Filter Based Detuning Control with a Digital SRF CW Cavity Simulator Proc. of IPAC 2018, Vancouver, Canada, WEPAK012

## Slide 21: Back-up slides down here

## Slide 22
CW operation: A electro-magnetic-mechanical -thermo-acoustic coupled problem?
Df (Hz) rms
PLHe (mbar)
T=1.5-1.9K
Pheater (W)
TLHe (K)
Qcrit (W/cm²)
Pheater (Dfmax) (W)
T. Peterson, TESLA-Report 1994-18
Cavity driven by LLRF at E0=15 MV/m
Piezo compensation in PI loop mode with low-pass filtering, QL=1.4.107
Additional power dissipated in LHe bath by heater (few cm²) within liquid
Microphonics recorded while heater is powered
Regimes of non-stratified flow?

## Slide 23: LLRF studies: Summary
QL | sf (Hz) | sf (deg) | sA/A | Pf (kW)
5.107 | 9.5 | 0.008 | 1.10-4 | 1.106
1.108 | 7.9 | 0.009 | 2.10-4 | 0.595
2.108 | 4.2 | 0.024 | 3.10-4 | 0.324

## Slide 24
Tuner dynamics: Higher order response?
Higher harmonics
Transfer function
Microphonics
background
substracted
Measured with first version of piezo frame (2005-2006)
Higher harmonic content most probably by piezo amplifier (even within drive signal?)
Measured at high excitation amplitudes (above 20 Hz)
Here complete detuning spectrum taken
at a given excitation frequency
Usually transfer functions taken with lock-in
amplifier to reduce noise content
(Stanford Research, SR850)

## Slide 25: Detuning spectrum versus bandwidth
For two different tuning schemes (Saclay I and INFN Blade) open loop measurements of
microphonics vs. QL were performed
Both tuners showed to have different transfer functions and thus detuning spectra on the same cavity type!
QL,Saclay: 3.107-4.108
QL,Blade: 7.105-2.107
Saclay: Excitation of 1st mechanical eigenmode sets in
Blade: Mechanical eigenmode
at 300 Hz, vacuum pump freq.