## Slide 1: LEReC Retreat Laser System Performance
by
Patrick Inacker

## Slide 2: LEReC Laser Parameters
Parameter | Target Value | Currently Achieved Value
Power on Gun Table | 15 W | 8 W
Green Power in Trailer | 26 W | 26 W
Stability (Green Power) | 5% PV | 12% PV
Stability (Pointing) | 8% Beam D. | 50% Beam D.
Beam Profile | Gaussian | Distorted
Temp. Profile | Modulated Flat Top | ???
Extinction Ratio (Pulsed) | >10^6 | >10^7
MPS Shut OFF extinction (Pulsed) | >100 | >10^7
MPS Shut OFF extinction (CW) | >100 | <80

## Slide 3: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz

## Slide 4: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz
Oscillator:
Has shown sensitivities to humidity fluctuations
Source (?) of 30-50Hz power modulations

## Slide 5: LEReC IR Power Monitor before and after Shielding of the Oscillator
Before Shielding
After Shielding
IR Power Monitor:

## Slide 6: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz
60W Amp. Pump:
Inappropriate cooling water flow
Required daily current adjustments to reach nominal output power

## Slide 7: 60W Amplifier effects of insufficient Flow
Returnline
Temp.
Flow
60W Amp ON
60W Amp OFF

## Slide 8: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz
Optics and Rod cooling loop:
Inappropriate cooling water flow
Room temperature related power fluctuations

## Slide 9: Old LEReC Laser Cooling System
Thermotek
T255
Flow: <0.5 l/min
Temp: 24°C/31°C
Thermofischer
AC125A50 Circulating Bath
Flow: <0.5 l/min
Temp: 16°C/18°C
Thermofischer 1400
Flow: 11 l/min
Temp: 37.5°C
60W Amplifier Pump Diode
Fiber coupling heat sinks
Beam dumps
Main Amplifier Rod
Main Amplifier Pump laser

## Slide 10: New LEReC Laser Cooling System Upgraded last week of May
Thermotek
T257p
Flow: 1.3 l/min
Temp: 29.5°C/30°C
Polyscience
LS51
Flow: 2 l/min
Temp: 18°C/18.5°C
Thermofischer 1400
Flow: 11 l/min
Temp: 37.5°C
60W Amplifier Pump Diode
Fiber coupling heat sinks
Beam dumps
Main Amplifier Rod
Main Amplifier Pump laser

## Slide 11: Results with new Cooling System
Old Cooling system
New Cooling system
IR Power Monitor:

## Slide 12: Results with new Cooling System 60W Amplifier
Returnline
Temp.
Flow
60W Amp ON Flow higher after removing algae
60W Amp OFF

## Slide 13: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz
Non-Critically PM SHG stage:
12-15% PV Power fluctuations driven by room temperature and IR power fluctuations

## Slide 14: Phase matching Technique comparison
Crititally PM Crystal
Non-Crititally PM Crystal
Experiment to examine the output stability of both techniques in an unstable environment
The IR Beam is Split after the Pockels Cell
Converted Beam Power is measured individually and simultaneously
Critically PM Crystal is cut for 1064nm light => low conversion efficiency, but relative power fluctuations for 1030nm cut crystal are equivalent

## Slide 15: Phasematching Technique comparison
Power Split:
Total: 49W
Non-Critical: 41W
Critical: 8W
Average Powers (W):
Non-Critical: 16.92
Critical: 0.07

## Slide 16: Phasematching Technique comparison
Power Split:
Total: 49W
Non-Critical: 9W
Critical: 40W
Average Powers (W):
Non-Critical: 3.41
Critical: 0.366

## Slide 17: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz
Intensity Control EOM:
Insufficient and unstable extinction ratio due to temperature fluctuations induced by room temperature and Laser power fluctuations.

## Slide 18: Measurement of the EOM Extinction Ratio
Extinction Ratio requirement: >100
EOM extinction ratio fluctuates with AC cycle.
Could be room temperature or laser heating related
Further measurements have to be conducted after system improvements
Possible replacement with AOM or regular Pockels Cell
Small Crystal Aperture and long crystal leads to wavefront distortions

## Slide 19: LEReC Laser Overview
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
Output
20-30W cw 9.1&704MHz
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically PM LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
Output
12-18W cw 9.1&704MHz
Temporal Shaping:
Small Crystal aperture leads to clipping and beam distortion
High power induces thermal lens

## Slide 20: Beam Quality Degradation
Gun Table
Relay Table
Wavefront distortion and clipping leads to non-uniform laser beam profile on the Gun table.
Sources need to be identified and removed to achieve the proposed Gaussian Profile.

## Slide 21: Remote Control and Diagnostics
Pockels Cell
- Analog Control through VME, Error Readback
Intensity Control
- RS232 Control over the rotation of a HWP
Steering
- Piezo Mirror in Trailer and on Relay Table
- Zoom Lens to adjust beam size on Gun Table
- 2 Axis motion for steering on Cathode on Gun
Table
- Position Feedback
Chillers
- T257p
- Polyscience LS51
- ThermoFischer 1400
Power supplies
- Newport (PreAmps)
- TDK Lambda (60W Amp)
- Main Amp
Sensors and Cameras
Temp: Main Amplifier Coolant
Power Amplifier Coolant
Main Amplifier Surface Flow: Main Amplifier
Power Amplifier
Dump Loop
Power: LP on Gun Table (Live) (3W Thermal Sensor)
LP on Gun Table (Flipper) (50W Thermal Sensor)
IR Monitor in Trailer
IR Dump Power in Trailer
Green Power monitor in Trailer Diodes: IR 9.1MHz Macrobunch Signal SHG Train Signal
PreAmp signal 60WAmp Signal
Cameras: Relay Table
Gun Table
Exit Table
Current Planned

## Slide 22: General Open Tasks
Finalizing Laser enclosures in the trailer
  Shot-to-Shot motion of the Laser beam due to airflow blowing at the beam in the trailer
Installation of more diagnostics
  Output power information after every amplification stage
  Dumped pump laser power
Upgrade of Cooling loops
  Separation of beam dumps from sensitive amplifier equipment
  Rework loops to remove mixed metal configurations

## Slide 23: Critical Questions Where are we and where should we be?
Key Parameters affected by laser performance To be completed by end of LEReC commissioning phase
(End of 2018)
Demonstration of 30mA requires sufficient Laser Power. - 8W on the Gun table have been demonstrated without aperture (1%QE 8W Laser = 33mA)
All Hardware Installed. - Beam stabilization through the Transport, as well as other Laser design changes are not yet installed
4. Demonstration of Beam parameters suitable for cooling. - Temporal Laser Pulse shape measurement still outstanding. Electron Pulse shape with Transverse Deflecting Cavity performed and temp. shaping setup optimized.

## Slide 24: Critical Questions
When can we expect stable laser intensity: Both short and long term?
When can we expect the thermal lens issues in the temp. shaping crystals to be resolved?
Should we expect 10W on the gun table or on the cathode as maximum?
What can be done to provide a more uniform laser profile on the gun table
Laser intensity stability in IR already incresed (BM: Short term 6% PV Long term 20% PV) (AM: Short term 4% PV Long term 18% PV) Main part of power fluctuations originate in SHG stage, not IR. (Still >12% PV Short term)
2,3,4: All connected to the same stage: Temporal Shaping Long crystals = Strong thermal lensing Very Large laser beam diameter = Clipping and beam quality degradation Stage is last in system. Improvements happen from the top to bottom of the system Possible solutions: Larger Aperture Crystals = Less clipping for same beam size YVO4 Crystals not available with larger aperture
BBO Available, but twice the crystal thickness necessary
Use of interferometers to reduce total crystal length (Misalignment robust interferometer in design stage)
BM = Before cooling system modifications
AM = After cooling system modifications

## Slide 25: Conclusion
Many open tasks to reach the originally proposed performance specifications
Appropriate steps have been taken during Run18 which led to improvements
  24/7 capability
  System protection
  Increased repeatability and reliability
  Increased power transport to the gun
Time remains very tight to enable remote control, performance improvements and diagnostic design, installation and comissioning before the end of LEReC commissioning phase

## Slide 26: Acknowledgements
Laser group: Michiko Minty, Brian Sheehy (retired), Zhi Zhao, Patrick Inacker & Linh Nguyen
LEReC AIP Project Team under the lead of Alexei Fedotov
CAD Support groups

## Slide 27: Measurement of the EOM Extinction Ratio
High Power Meter
Low Power Meter
Temp. Shaping
EOM
Intensity Control
EOM was set up to dump all light into the High power meter (Simulates Tripped MPS system)
Leaking light intensity was measured and compared to EOM in 100% transmission mode

## Slide 28: Temporal Shaping
YVO4 Crystals
Large beam in Temp. Shaping Crystals clips on the Crystal edges
Small beam in Temp. Shaping Crystals leads to extreme thermal lens
Crystal Length | Transparency
19.68mm | 88%
9.84mm | 96%
4.92mm | 98%
2.46mm | 99%
Total | 82%

## Slide 29: Temp. Shaping with interferometers
Not Robust
Robust
Angle Misalignment
No Misalignment

## Slide 30: Laser Protection System
Laser
Chiller
Sensor Box
Room Interlock
Temperature
Flow
Power/Fluorescence
Personal Protection
Machine Protection (Laser MPS)
Independent systems
Both have to agree to enable the Laser
Temperature
Interlock Box

## Slide 31: Laser Protection System Flow Chart
Monitored Parameters
Chiller Temp.
Chiller Flow
Module Temp.
Power
...
In Range to SetPoints?
YES
Look up Failure Mode
if Main_Amp_Chiller_Flow == 0 Interlock_Main_Amp;
Alert_Operators;
If Chiller_Pre_Amp_Flow == 0
Interlock_Entire_System;
Alert_Operators;
...
When Interlocked: Wait for RESET
Monitoring Loop
NO
If YES Enable System

## Slide 32: LEReC Laser Transport

## Slide 33: Relay Table Layout
f = 7m
Camera
+35“
+12“
To LEReC
To CEC
Camera

## Slide 34: Gun Table Layout
Camera
Power meter
Photodiode
Pinhole Wheel
Motorized
tilt-mirror
X-Motion
Y-Motion
f = 1m
Motorized
Zoom lens
Sampler
Flip-Mirror
Power Meter
From Relay Table
To DC-Gun
Filter Wheel