## Slide 1: How well do we understand NIR detectors? (focus on hybrid HgCdTe)
Roger Smith
Caltech
2018-12-03

## Slide 2: Photodiodes, with low bandgap
Bump bond low bandgap material to silicon readout IC.
Integrate charge on diode capacitance.
Material becomes transparent when
hc/λ < Eg

## Slide 3: Basic theory of PN junction
When majority carriers diffuse across PN junction they recombine leaving a charged but carrier free region. The charge on the donor atoms produces an E field which opposes further diffusion.
For given donor density profile, Q(x) at equilibrium, within space-charge region:
E  Q(x).dx
V  - E(x).dx
http://en.wikipedia.org/wiki/P-n_junction

## Slide 4: Built in voltage
Depletion width
Built in voltage
E  Q(x).dx
V  - E(x).dx
Equal areas

## Slide 5: Reverse bias
At “reset”, a reverse bias is applied : charge is removed, increasing depletion width.
Built in voltage
Depletion width increases
Reverse bias

## Slide 6: Photon makes electron hole pair
Q
E
V
Initial depletion width
h
e

## Slide 7: E field separates e-h  voltage drops
Q
E
V
Initial depletion width
Final depletion width
Narrowing of depletion region maps to a drop in the reverse bias.

## Slide 8: Charge accumulates during exposure
Saturation = forward bias occurs when depletion region collapses
Q
E
V
Initial depletion width
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
h
e
Signal

## Slide 9: PN junctions on common substrate
Silicon Multiplexor ~ 600µm thick
Undepleted Dsub (N)
Depleted N-type
Depleted P-type
CdZnTe substrate is fully removed
Undepleted implant (P)
HgCdTe
Indium bump bonds
~3m ?
18m
Epoxy wicked into voids between bumps prior to substrate removal
Silicon Carbide or Molybdenum package
Epoxy

## Slide 10: PN junctions on common substrate
Silicon Multiplexor ~ 600µm thick
Undepleted Dsub (N)
Depleted N-type
Depleted P-type
Undepleted implant (P)
HgCdTe
18m
Epoxy wicked into voids between bumps prior to substrate removal
Silicon Carbide or Molybdenum package
Epoxy
Dielectric
Areas with mobile carriers act as capacitor plates

## Slide 11: Inter-Pixel Capacitance
Silicon Multiplexor ~ 600µm thick
HgCdTe
Dielectric constant of epoxy underfill increases IPC
Silicon Carbide or Molybdenum package
Depletion region is dielectric:
More signal (or lower bias)
Higher capacitance
= non-linear
Conductive substrate shields pixel from its neighbors

## Slide 12: IPC non-linearity
Silicon Multiplexor ~ 600µm thick
HgCdTe
Silicon Carbide or Molybdenum package
At higher fluence, depletion region narrows
Carriers in implant with high signal get closer  increased IPC expected.
Greater IPC?
See Kevan Donlon’s talk

## Slide 13: IPC dependent on detector bias ??
Silicon Multiplexor ~ 600µm thick
HgCdTe
Silicon Carbide or Molybdenum package
Higher initial IPC
At high reverse bias depletion regions may merge and open up a new path for electric field lines between pixels
SPECULATION:

## Slide 14: Pixel boundaries after reset
Silicon Multiplexor ~ 600µm thick
HgCdTe
Silicon Carbide or Molybdenum package

## Slide 15: Signal contrast moves boundary
Silicon Multiplexor ~ 600µm thick
HgCdTe
Silicon Carbide or Molybdenum package
Pixel shrinks as charge accumulates so PSF seems to grow.
Brighter-fatter paper by Andres Plazas; see Eric Huff’s talk

## Slide 16: Image persistence
Depleted
Mobile holes
Mobile electrons
-
-
-
+
+
+
-
-
-
-
-
+
+
+
+
+
Signal (lower bias)
As signal accumulates the depletion width is reduced. Traps newly exposed to charge can capture some mobile carriers.
+
+
+
+
+
-
-
-
-
-
Idle / Reset (initial reverse bias)
Charge diffusing across the PN junction recombines leaving behind charge donor sites that produce an E field opposing the diffusion. When traps in depleted area release charge, it is swept towards depleted areas by this E field, and cannot be recaptured: eventually only undepleted area has populated traps.
P
N
-
-
-
-
-
+
+
+
+
+
Next Exp. (bias relapses)
The motion of the released charge reduces the E field: the reverse bias and depletion width tend to towards the previous values. The voltage change has the same polarity as signal and is“persistence”.
Trapped holes
Trapped electrons
-
-
-
+
+
+
+
+
+
+
+
-
-
-
-
-
Reset (increased bias)
At “reset” the reverse bias and wider the depletion region ware restored,but trapped charge stays behind. When this is later released, it is swept by the E field out of the depletion region to rejoin carriers of the same type.
-
+

## Slide 17: Reset
Reverse bias is reapplied at reset.
Ideally this would reestablish the initial depletion width, but..
some charge remains behind in traps.
Q
E
V
Before Reset
Ideal Reset
Actual reset

## Slide 18: Charge release  Signal in later frames
Q
E
V
persistence
Persistence

## Slide 19: MBE HgCdTe Dark current
~9
~5
~2.5
~1.7
Typical InSb
Dark Current
108
107
106
105
104
103
102
10
1
10-2
10-3
10-4
230
210
190
170
150
130
110
90
70
50
30
Temperature (K)
Dark
Current
Electrons
per pixel
per sec
18 micron
square
pixel
10-1
What causes floor ?

## Slide 20: Dark current floor
Dark current is high after turn-on due to de-trapping.
Floor continues to drop for weeks if dark and stable.
De-trapping is less noisy than equivalent photo-generated signal  our noise model may be wrong at low enough temperatures for de-trapping to dominate dark current.

## Slide 21: Circuit for one pixel
Photodiode is reverse biased by momentarily closing reset switch,
  DSUB =VRESET+0.5V
Low temperature makes leakage currents in silicon switch and buffer transistor negligible.
Photogenerated e-h pairs are collected on diode capacitance causing drop in reverse bias.
Voltage at output of buffer MOSFET is read as charge accumulates and dV/dt is calculated without effecting charge on diode.
Minimum of 2 reads are required to remove large electronic offsets, bias voltage drift and √kTC noise left behind by thermal noise in reset switch resistance.
….is higher than VRESET!

## Slide 22: Addressing multiple pixels
Single bit travelling in shift registers along two edges selects one row and one column at a time to implement raster addressing.
Voltage on selected pixel is connected to output buffer by analog switches.
Charge on pixel is unaffected.

## Slide 23: Pixel circuit
DSUB ~0.8V
VRESET ~0.3V
RG
0 – 3.3V
Vo
BiasGate 2.1V
RowSelect
ColSelect
Indium Bump
Power dissipated only when pixel is selected
Tempco = 600 to 800 e-/K
 Clock continuously whether idle, reset, exposure delay or read
BiasPower, 3.3V
10 kΩ

## Slide 24: Constant Cadence clocking
Temperature coefficient 600-800 e-/K
Both pixel and output amplifier (and load) contribute to temperature coefficient.
Depends on location on detector.
Pixel only dissipates power while selected.
Maintain constant clocking cadence across idle, reset, exposure_delay and reads
Pad timing to make reset and read timing identical, and to avoid changes in cadence when processing commands.

## Slide 25: Correlated Double Sampling
Exposure delay = p dummy reads for constant self heating
Subtract first frame from last frame
Equivalent to Fowler sampling with m = 1
Ignore p scans
At least one reset between frames
Reset while idling
Initial scan
Final scan
Exposure time
Frame time
Dummy reads to maintain constant power dissipation

## Slide 26: Fowler sampling, … same dissipation
Exposure delay is in units of full scan ties but need not be multiple of m.
Subtract mean of first group from mean of last group.
Coadd m
Ignore p scans
Coadd m
Exposure time
Frame time
“Fowler Sampling”

## Slide 27: Sample Up The Ramp … same dissipation
Store every scan (no real time coadd)
Use post facto least squares fit to measure slope with best S/N;
Effective exposure duty cycle due to weighting of shot noise by least squares ~ 90%; reduce this to include effect of the reset overhead.
Equivalent MultiAccumulate with m=1.

## Slide 28: Pixel circuit
DSUB ~0.8V
VRESET ~0.3V
RG
0 – 3.3V
Vo
BiasGate 2.1V
RowSelect
ColSelect
Indium Bump
BiasPower, 3.3V
10 kΩ

## Slide 29: Pixel circuit, showing caps
DSUB ~0.8V
VRESET ~0.3V
ResetGate
0 – 3.3V
Vo
BiasGate 2.1V
RowSelect
ColSelect
Indium Bump
All capacitors are non-linear
BiasPower, 3.3V
10 kΩ

## Slide 30: Pixel circuit, showing caps
DSUB ~0.8V
VRESET ~0.3V
ResetGate
0 – 3.3V
Vo
BiasGate 2.1V
RowSelect
ColSelect
Indium Bump
All capacitors are non-linear
Capacitive coupling of RG
changes diode voltage
 trapping charge, even in dark.
See poster by Tim Greffe
Explains Global Reset transient
10 kΩ
BiasPower, 3.3V

## Slide 31: Large contact Resistance
DSUB ~0.8V
VRESET ~0.3V
ResetGate
0 – 3.3V
BiasPower, 3.3V
Vo
BiasGate 2.1V
RowSelect
ColSelect
Indium Bump
Slows reset and signal
Poster by Tim Greffe:
This limits effectiveness of reset as electronic dark slide.
Deserves more attention
~1012 Ω
10 kΩ

## Slide 32: Initial signal deficit
SUTR data acquired at 100 Hz through a single channel. Several illuminated frames were averaged as were several matched darks which have been subtracted.
The fit was derived using the points marked in cyan to eliminate the area affected by the initial signal deficit.
The quadratic fit is extremely good except for the large deviation in the first 3s (300 frames).
Pixel contact resistance ? …. much larger effect and faster settling than charge trapping

## Slide 33: Signal loss due to charge trapping
Flux
Reset
Signal
Trapping
De-trapping
Image persistence
= reciprocity failure ?
Melissa Marshall
SURF 2013

## Slide 34: Initial signal deficit (AKA burn in)
1st frame / 4th frame -1
Image persistence
Engineering grade H2RG (2.3 µm cutoff from Euclid)
This suggests that charge trapping is dominant cause of initial signal deficit in some areas of this device.

## Slide 35: Read noise
The photodiodes generate more read noise than the ROIC and have much stronger 1/f component, as demonstrated by much lower noise and better reduction by multiple sampling for
  Permanent reset (shorting out detector noise)
  Reference pixels
Is this due to trapping/detrapping?
Contact resistance is about the correct magnitude to produce about this much Johnson noise.
  This is worth investigating since lower contact resistance could improve both dynamics (linearity) and noise.

## Slide 36: Linearity
First, subtract time-dependent offsets
  Self heating (minimize by Constant Cadence Clocking)
  Electronic drifts (reference pixels help)
  Reset induced (de)trapping
Correct for dependence of capacitance on voltage
Correct for trapping of charge (complicated)
Correct for photodiode time-constant (contact resistance)
Everything is spatially variable.
Rarely is this well characterized and corrected
See Bernie Rauscher’s talk

## Slide 37: Simple test
Applying all corrections to SUTR data for a wide range of fluxes, extrapolate back to zero exposure time.
Do all curves predict zero signal at zero exposure time?
Do they even intersect within some small error box?
Good luck !!

## Slide 38: To do list
Characterize contact resistance and its consequences
Study burn-in and reciprocity failure to see how much comes from contact resistance an how much from trapping.
Evaluate noise produced by dark current at floor. Hot pixels may be noisier.
Improve linearity calibration methods and/or fix these odd behaviors which make linearity correction so difficult.