## Slide 1: Uniformity and Stability of the LSST Focal Plane
P. O’Connor
Instrumentation Division, BNL

## Slide 2: Outline
Focal plane
Sensors and rafts
Electro-optic tests
Results
  Uniformity
  Stability
Cadence implications

## Slide 3: FOCAL PLANE, SENSORS, and RAFT TOWERS

## Slide 4: Focal Plane Layout
3.5° diameter
0.2” (10mm) pixels
189 4K  4K CCDs
21 “rafts”
Fully-butted mosaic (90% fill factor)
N. Kurita., et al. " Large Synoptic Survey Telescope camera design and construction." SPIE 2016, 9912-28.
guider
WFS

## Slide 5: Sensors
Uniform illumination
Test target image
CCD in test fixture
T-E2v CCD250
T-DALSA/ITL STA3800C
(overscan displayed)
(overscan suppressed)
4 parallel phases
2-stage output amplifiers
Entrance window: laser-annealed shallow-B impant
Antiblooming midline stop
3 parallel phases
single-stage output amplifier
JFET source follower buffer
Entrance window: chemisorption process
Both devices feature: Fully-depleted, high-r silicon 100mm thick, 10mm pixels, 4Kx4K format, 16 independent 512x2K segments, 4-side buttable packages with >90% fill factor.
Mechanical mount and electrical interface 100% interchangable
annular shims
plane parallel package body

## Slide 6: Requirements and verification matrix
Parameter | Req’t. | Meas. | unit
QE u | 41 | 65 | %
QE g | 78 | 83 | %
QE r | 88 | 90 | %
QE i | 81 | 94 | %
QE z | 75 | 88 | %
QE y | 21 | 25 | %
Diffusion | 5.0 | 3.6 | mm rms
Dark current (95%) | 0.2 | 0.02 | e-/pix/s
Defects | 0.5 | 0.002 | %
Frame read time (16Mpix) | 2 | 2.0 | sec
Read noise/pixel | 8 | 4.5 | e- rms
CTI (ser/par) at 1ke- signal | 5/3 | 3/0.7 | ppm
Nonlinearity | 2 | 0.6 | %
Video crosstalk | 2 | 0.08 | %
Pixel size | 10 | 10 | mm
Flatness (95%) | ± 5 | 1.9 | mm
Coplanarity (95%) | ± 9 | 6.3 | mm
Inactive area | 8 | 8 | %

## Slide 7: Focal plane is composed of 21 Science Raft Tower modules (RTM)
Photons in – bits out
Complete 12K x 12K imager
Support sensors mechanically to meet strict coplanarity
Thermal management of sensors and electronics
Provide bias, timing, and control signals for CCD operation
144 video channels with low noise analog signal processing, digitizing, multiplexing of pixel data
Diagnostics + monitoring
12.7 x 12.7 x 42cm3
20 kg
39W

## Slide 8: Raft tower (partially assembled condition)

## Slide 9: Raft tower in grid installation fixture

## Slide 10: Production flow
functional test
Data base
ASIC
REB
functional test
Data base
Burn-in,
Thermal cycle
Image archive
EO parameter catalog
Human-readable html report

## Slide 11: Example outputs:
144 channel 55Fe spectra
144 channel PTC
144 channel flatfield

## Slide 12: EO parameter summary

## Slide 13: Specialized diagnostics – video waveform reconstruction, noise correlations
144 x 144-channel noise correlation (overscan regions)
144-channel synthetic video waveforms (CCD-output and post-DSI)
Single-channel video chain waveforms
C. Juramy

## Slide 14: LSST Science Raft Tower Module Production Status (Dec. 2018)
Assembled and tested: 23 rafts (21 of 22 production deliverables + 2 Engineering/Test units): 207 CCDs, 3,312 video channels (3.3Gpix)
engineering and test units (2)
production units (21 + spare)
Key
RTM1
Feb ‘17
RTM2
May ‘17
RTM3
Sep ‘17
RTM4
Jul ‘17
RTM5
Oct ‘17
RTM6
Dec ‘17
RTM7
Dec ‘17
RTM8
Mar ‘18
RTM9
Feb ‘18
RTM10
Apr ‘18
RTM11
Jun ‘18
RTM12
Apr ‘18
RTM13
Jun ‘18
RTM14
Jul ‘18
RTM15
Aug ‘18
RTM16
Jul ‘18
RTM17
Aug ‘18
RTM18
Aug ‘18
RTM19
Aug ‘18
RTM20
Oct ‘18
RTM21
Nov ‘18
RTM22
Oct ‘18
ETU1
Jan ‘17
ETU2
Mar ‘17
e2v
ITL
vendor
assembled and tested
shipped
future production

## Slide 15: UNIFORMITY

## Slide 16: Parameter distributions – 20 rafts, 180 CCDs, 2880 video channels

## Slide 17: Parameter distributions – 20 rafts, 180 CCDs, 2880 video channels
0.0
0.2
0.1
0.3
E2V
ITL
E2V
ITL
Read noise, e-
Gain, e-/ADU
Full well, e-
Dark current (95th %ile @ -90C) e-/pix/s
PSF, s (mm)
Serial CTI, ppm

## Slide 18: Parameter distributions – 20 rafts, 180 CCDs, 2880 video channels
0.0
0.2
0.1
0.3
E2V
ITL
E2V
ITL
Read noise, e-
Gain, e-/ADU
Full well, e-
Dark current (95th %ile @ -90C) e-/pix/s
PSF, s (mm)
Serial CTI, ppm
3.1 ±5.7
4.8 ±28.3
143 ±8.5
173 ±33.4
0.69 ±0.04
0.93 ±0.08
0.015 ±0.034
0.054 ±0.046
4.09 ±0.15
4.53 ±0.26
4.68 ±0.24
6.84 ±1.8
<sITL/sE2V> = 3.6

## Slide 19: Parameter distributions – 20 rafts, 180 CCDs, 2880 video channels
 E2V
ITL

## Slide 20: Quantum efficiency – 20 rafts, 180 CCDs

## Slide 21: Temperature distributions – 20 rafts, 180 CCDs, 20 baseplates, 600 REB sensors
5.62 ± 0.80
1.81 ± 0.59
BASEPLATEs
ELECTRONICS
CCDs - BASEPLATEs
CCDs
setpoint = -90
coldplate setpoint = -40
CCD temperature set by makeup heaters on baseplate. For most rafts, collected full EO suite at -90C and -100C setpoints.
Raft electronics temperatures not actively stabilized and have 15-20C gradients across the long REB boards.

## Slide 22: Example: QE dependence on CCD temperature, RTM-4

## Slide 23: Example: gain and noise dependence on electronics temperature, RTM-9

## Slide 24: Parameter statistics and temperature coefficients
Pmt | <E2V> | <ITL> | <TOT> | sE2V | sITL | sTOT | dp/pdTE2V | dp/pdTITL | dp/pdTTOT | UNIT
QE-u | 69.8 | 62.2 | 67.9 | 6.3 | 7.8 | 7.8 | 0.0013 | 0.0050 | .00137 | %
QE-g | 89.9 | 88.4 | 89.5 | 2.0 | 3.4 | 2.8 | 0.00027 | -0.00023 | 5.3e-5 | %
QE-r | 95.6 | 94.3 | 95.6 | 2.4 | 3.4 | 3.1 | -0.00017 | -0.00011 | -.00014 | %
QE-I | 95.1 | 99.4 | 98.1 | 3.6 | 4.4 | 4.2 | -0.00016 | -0.00016 | -.00016 | %
QE-z | 84.2 | 92.7 | 86.6 | 2.4 | 4.6 | 5.3 | 0.00088 | 0.00063 | .00075 | %
QE-y | 25.9 | 31.5 | 28.9 | 3.5 | 4.2 | 4.8 | 0.00506 | 0.0080 | .0051 | %
read noise | 4.7 | 6.1 | 4.8 | 0.23 | 1.8 | 1.5 | -0.00035 | 0.0009 | .00017 | e-
gain | 0.69 | 0.94 | 0.71 | 0.034 | 0.11 | 0.14 | 0.00023 | 0.0002 | .00024 | e-/ADU
full well | 144 | 186 | 147 | 8.0 | 40 | 30 | nan | -0.00017 | -2.3e-5 | ke-
dark current | 0.013 | 0.038 | 0.017 | 0.03 | 0.048 | 0.041 | 0.0052 | -0.002 | 0.0029 | e-/pix/s
PSF | 4.12 | 4.48 | 4.2 | 0.14 | 0.28 | 0.28 | 0 | 0 | 0- | Um
CTI-serial | 1.52 | 2.07 | 1.7 | 6.6 | 45 | 27 | nan | nan | -.016 | ppm

## Slide 25: STABILITY

## Slide 26: RTM-21, 15-hour gain and offset stability
TEMPERATURES
GAINS
OFFSETS
1 ADU ~ 0.7e-

## Slide 27: RTM-21, 15-hour gain and offset stability
TEMPERATURES
GAINS
OFFSETS
1 ADU ~ 0.7e-

## Slide 28: CADENCE IMPLICATIONS

## Slide 29: LSST dither strategy
Baseline WFD
no dither
DDF
no field dither
sky rotation
Modified WFD
large field dither
sky rotation
Awan ApJ 2016

## Slide 30: FPA distribution of rafts
“rings”
“spokes”
“random”
Depending on the arrangement of rafts in the FPA, in DDF cadence objects will experience unequal fraction of visits to ITL and E2V rafts
(one of 203,490 possibilities)

## Slide 32: SUMMARY
All of the LSST science raft towers have been produced and EO performance measured
Key parameters meet requirements at >98% level
Distribution of EO parameters is strongly tied to CCD supplier
  Central value, spread, and temperature coefficient differences
Gain and offset remain stable at 1e-4 level over 15-hour runs in test cryostat
Choice of placement of rafts across the focal plane may impact the survey uniformity

## Slide 33: ACKNOWLEDGEMENTS
Justine Haupt
Ivan Kotov
Andrei Nomerotski
Guangyu Zhang
Peter Takacs
Sven Herrmann
Jim Chiang
Claire Juramy-Giles
Pierre Antilogus
Hye-Yun Park
Al DellaPenna
Matthew Rumore
Peter Doherty
John Kuczewski
Wei Lu
Homer Neal
Al DellaPenna
Bill Wahl
Sean Robinson
Brian Walsh
This presentation has been co-authored by employees of Brookhaven Science Associates, LLC. Portions of this work are supported by the Department of Energy under contract DE-AC02-98CH10886 with Brookhaven National Laboratory. LSST project activities are supported in part by the National Science Foundation through Governing Cooperative Agreement 0809409 managed by the Association of Universities for Research in Astronomy (AURA), and the Department of Energy under contract DE-AC02-76-SFO0515 with the SLAC National Accelerator Laboratory. Additional LSST funding comes from private donations, grants to universities, and in-kind support from LSSTC Institutional Members.

## Slide 34: Supplementary Material

## Slide 35: Raft electronics
Functions:
Video CDS processing, dual-slope integration
18-bit digitization and serialization
HV clock and bias generation
CCD temperature monitoring and control
Onboard diagnostic features including waveform recorder
Implementation:
Three 4” x 16”, 14-layer boards
Video processing in CMOS ASIC
144 video channels
Performance:
72 Mpixels/s readout rate, 2.3Gbits/s
7e- read noise
0.08% crosstalk
380mW/channel total power dissipation
Russo, S., et al. "The LSST Science Raft Electronics." RT2014-19th Real-Time Conference. 2014.

## Slide 36: Test stand cryocooler refrigeration capacities
Single PT-30 cryo head for RSA
DT
expected Q
Dual NF-55 cold head for electronics
2X
DT
expected Q

## Slide 37: Parameter distributions – 20 rafts, 180 CCDs, 2880 video channels
(if trap)
(if CTI)

## Slide 38: Example: gain and noise dependence on electronics temperature, RTM-

## Slide 39: Video sampling-scope mode
Arrange a repetitive CCD output waveform (i.e. dark, flat)
Acquire one sample per pixel period, incrementing the ADC sampling clock 10ns per pixel
Combined with “transparent mode” feature of DSI ASIC, creates virtual 100MHz, 18b oscilloscope for probing video channels
(a) ASPIC DSI mode
(b) ASPIC Transparent mode – amplified CCD output

## Slide 40: Virtual sampling scope
Time
Integration windows
Transparent mode
Reset feedthrough
Reset level
Video level
2msec

## Slide 41: Virtual sampling scope
Time
DSI
Transp.
Transparent and DSI modes
2msec

## Slide 42: Flatness
111-01 CCD250 #021
112-01 CCD250 #035
112-05 CCD250
99% range:
1.95μm
1.85μm
2.37μm
1μm
1μm
1μm
P.Z. Takacs, BNL

## Slide 43
Bias10%FS
max FWC 82%FS

## Slide 45: Specialized diagnostics – video waveform reconstruction