## Slide 1
LAPPDTM / MCP-PMT
Goal:
FY 19 Proposed tasks:
MCP-PMT after pulse study to better understand the ion feedback for LAPPD
Improvement of RMS timing by modifying the bias voltages
Produce a detector with 5μm pore size MCP and minimum spacing to further improve magnetic-field performance and fast timing (possibly <10 ps)
Test of MCP-PMT/LAPPD with different configurations (smaller pore size and reduced spacing) in lab and in a magnetic fields
Demonstration of capacitive-coupling pixelated readout through ALD coated glass
Adapt LAPPDTM to EIC requirements (magnetic fields, pixelated readout)
FY 18 Report:
6 cm MCP-PMT characterized in details in magnetic fields
20 cm LAPPD characterized in magnetic fields
Pixelated readout baseline tested in beamline

## Slide 2
LAPPDs – Characteristics (6 cm ones) in magnetic fields
θ = 0o
Independently biased design
MCP-PMT characteristics in dependence of: Magnetic field strength, HV, tilt angle, and gap voltages were all tested in magnetic fields
Baseline tests show for a device with Gain > 5x105 , Magnetic Field tolerance is 0.7 T
The MCP-PMT performance in magnetic field is clearly angle related, due to the 8o MCP bias angle, the highest gain is obtained around 8o.
Notice the two peaks around ±8o, indicating the effect from upper and lower MCP bias angles are different. Simulation is undergoing with collaborators to explain the different effect.
8o MCP bias angle

## Slide 3
LAPPDs – Characteristics (6 cm ones) in magnetic fields
HV(MCPs) was fixed, varies the HV of first gap ΔHV(photocathode–top MCP) by adjusting HV (Cathode).
Gain increases as ΔHV increases to a maximum then decreases, this can be explained by the MCP gain dependence of primary electron energy.
Ref: S. Jokela et al. Physics Procedia 37 (2012) 740
MCP-PMTs design/operation in magnetic fields resulted in a SBIR phase 1 award (Incomand ANL, $150k for 9 months) to develop magnetic field tolerant LAPPDTM
(> 1.5 Tesla)

## Slide 4
LAPPDs – Characteristics (20 cm) in magnetic fields
Commercial LAPPDTM delivered and installed at ANL B field facility
Pore size: 20 µm
Activation area: 195 mm x 195 mm

## Slide 5
Similar behavior as 6 cm MCP-PMT: gain decrease as the magnetic field increases
Two local gain maximum corresponding to the 13o bias angle of MCPs used in LAPPDTM
θ = 0o
B = 0.1 T
LAPPDs – Characteristics (20 cm) in magnetic fields
Due to the magnetic sensitive components (Kovar is used as shims in the current LAPPDTM), we can not go to high magnet field test, a new LAPPDTM with non-magnet components is scheduled to be fabricated and tested in Sep. 2018. The results here demonstrate the capability of the facility for 20 cm LAPPDTM.
13o MCP bias angle

## Slide 6
HV applied to all three gaps affects the gain of the LAPPD
HV between the photocathode and MCP1 gap has the greatest slope, indicating the strongest effect
LAPPD gain becomes a constant with the MCP2-Anode bias HV above a threshold
HV applied to MCPs seems to have NO preference, equally affects the LAPPD gain
Gap and MCP ΔHV dependence
LAPPDs – Characteristics (20 cm) in magnetic fields
Gap ΔHV
MCP ΔHV

## Slide 7
LAPPDs – Pixelated readout
Pad sizes:
2mm x 2mm
3mm x 3mm
4mm x 4mm
5mm x 5mm
Spacing between pads:
0.5 mm
Demountable chamber installed on the stage of Fermilab Test Beam Facility MT6.2C

## Slide 8
Expected position resolution σ (expected) = pixel size/√12
Beamline experiment preliminary results show that experimental position resolutions are close to the expected position resolutions
Pixel size | 2 mm x 2 mm | 3 mm x 3 mm | 4 mm x 4 mm
σ (x) | - | 1.01 mm | 1.11 mm
σ (y) | 0.73 mm | 0.93 mm | 1.43 mm
σ (expected) | 0.6 mm | 0.9 mm | 1.2 mm
LAPPDs – Pixelated readout
Example correlation between the y-axis of a 3 mm x 3 mm pad and the MWPC projection

## Slide 9
LAPPDs – Conclusion
LAPPD/MCP-PMTs were characterized in details in magnetic fields
Further development of magnetic field tolerant LAPPD progresses well
Pixelated readout development for LAPPD was started, will be the next focus
LAPPD R&D, DIRC R&D, electronics R&D are well aligned with each other for demonstration of DIRC prototype with LAPPDTM sensors