## Slide 1: Measurements and Calibration of a Deep Dielectric Charge Monitoring Sensor for the Europa Clipper Mission
Wousik Kim (JPL), John Goldsten & Zach Yokley (APL)
2019 ATF
User Meeting
2019 ATF Users Meeting: Experiment Completion Report

## Slide 2: Mission Challenges – Scientific Case
L/V: Launch Vehicle, VGA: Venus Gravity Assist, JOI: Jovian Orbit Insertion, TID: Total Ionizing Dose, RDF: Radiation Design Factor, iESD: internal ElectroStatic Discharge

## Slide 3: Experimental Setup
The pulsed, mono-energetic electron beam of 40-60 MeV was used.
Two motion stages allow charge stack and degraders to be moved in and out of electron beam
Degraders form low-energy tail as a way to exercise all three plates
Stepped degrader provides three positions: “none”, “thin” and “thick”
ATF Chamber Setup
The technical data in this document is controlled under the U.S. Export Regulations; release to foreign persons may require an export authorization. Pre-Decisional Information – For Planning and Discussion Purposes Only
Beam
Flight Configuration
Simulated Spectra
(by Geant4)
Detector plates
Aperture

## Slide 4: Experiment Results
100 pC pulses @ 1.5 Hz
Absolute level match prediction within factor of 2
Ratio of currents within 20%
50 MeV beam through thick degrader
Front plate
Middle plate
Back plate
Pulse Response
Beam Sync
The technical data in this document is controlled under the U.S. Export Regulations; release to foreign persons may require an export authorization. Pre-Decisional Information – For Planning and Discussion Purposes Only
Beam Degrader

## Slide 5: Conclusions & Further Work
Experiment at ATF was very successful in validating sensor response to high energy electrons
Considering adding 4th charge plate to better cover ultra-wide energy range
Would like to return and experiment again after development of the engineering model sensor
Simulations and beam experiment in agreement – good opportunity to publish results from final sensor