## Slide 1: LEReC Laser Review
Laser related MPS Patrick Inacker-Mix

## Slide 2: Agenda
LEReC Machine Protection requirements
Pulsed Mode
CW Mode
Laser System Protection

## Slide 3: LEReC MPS requirements
Parameter | Target Value | Achieved Pulse Mode | Achieved CW Mode
Reaction Time | 15µs | < 1µs | < 1µs
Extinction Ratio (Before Shutter closes) | 180 (30mA) 300 (50mA) | 10^7 | > 200 -
Required Extinction Ratios based on a Shutter closing time of 15ms.

## Slide 4: LEReC Machine Protection (Pulsed Mode)
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically Phase-Matched LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
PC
Power
~22W
Fast / Slow Shutter
Pockels Cell extinction ratio 800 – 1000 SHG: 5*10^5-10^6

## Slide 5: LEReC Machine Protection (CW Mode)
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically Phase-Matched LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
PC
Power
~22W
Fast / Slow Shutter
Pockels Cell extinction ratio 25-35 EOM extinction ratio 10-500
EOM extinction ratio is strongly correlated to Crystal temperature and drifts around
Bias Voltage 0 for Fail-Safe operation
CW Pockels Cell extinction ratio stable around 30 (Degraded within 1Week due to permanently applied HV

## Slide 6: EOM extinction ratio Measurements
EOM extinction ratio over 9 hours
EOM extinction ratio sensitive to laser power fluctuations
EOM alone is not capable of providing the needed extinction ratio for extended periods of time.
Added cw Pockels Cell to boost extinction by a factor of 20-30

## Slide 7: LEReC Machine Protection (CW Mode)
704 MHz High-Order-Modelocked Oscillator
EOM
PC
SHG
976nm Pump
60W
976nm Pump
Mach-Zehnder
1W
160-200W
EOM
85µm, 80cm fiber rod
SC
YDF
SC
YDF
DC
YDF
Non-Critically Phase-Matched LBO Crystal
ISO
Polarizer
976nm Pump
Intensity Control
Temporal Shaping
PC
Power
~22W
Fast / Slow Shutter
AOM will replace EOM + CW PC
AOM operates by deflecting the beam using a standing pressure wave. Refractive index changes due to temperature changes does not translate to power leakage but changes in deflection angle
AOM

## Slide 8: Laser System Protection

## Slide 9: Laser System Protection
Laser
Chiller
Sensor Box
Room Interlock
Temperature
Flow
Power/Fluorescence
Personal Protection
System Protection (Laser MPS)
Independent systems
Both have to Agree to enable the Laser
Temperature
Interlock Box

## Slide 10: Laser System Protection Sensor Box
Arduino driven, stand-alone system
Network access through RS232 for readout and remote control.
Full local control. Fully customizable sensor array (Analog&Digital)
Version 3:
5 RTD Temperature Sensor inputs
4 Flow Sensor inputs
2 Dry contact inputs
8 Individually controllable Interlock Channels
Programmable fault scenarios and system conditions. In continous, maintenance free 24/7 operations since 05/2018

## Slide 11: Laser MPS Flow Chart
Monitored Parameters
Chiller Temp.
Chiller Flow
Module Temp.
Power
...
Reset
Local / Remote
In Range to SetPoints?
YES
Look up Fault Scenario
if Main_Amp_Chiller_Flow == 0 Interlock_Main_Amp;
Alert_Operators;
If Chiller_Pre_Amp_Flow == 0
Interlock_Entire_System;
Alert_Operators;
...
Monitoring Loop
NO
If not Interlocked Enable System
Interlock according to Fault Scenario

## Slide 12: Summary
Experience from last year requires upgrade of CW MPS setup
  EOM+PC will be replaced by single AOM
Added Laser Protection enabled 24/7 operations
  Runs flawlessly since installation
  Can support design changes and new requirements

## Slide 13: Acknowledgements
Laser group: Michiko Minty, Brian Sheehy (retired), Zhi Zhao, Patrick Inacker & Linh Nguyen
Instrumentation, RF and Controls Hardware and Software Engineering
CAD Support groups