## Slide 1: Laser Beam Position Stabilization
Linh Nguyen

## Slide 2: Current Status
During Run18, the laser beam exhibited two types of position instabilities:
- fast jumping - slow drifting
Parameter | Target Value | Current Value
Pointing stability on cathode | <10 µm c.o.m rms | <75 µm c.o.m rms

## Slide 3: Fast Jumping
The fast jumping was easily observable while watching cameras and routinely logged by the gun-table camera (1 Hz refresh rate)
It is the fast instability that drives the values in the table (see next slide)
X position
Y position

## Slide 4: Slow Drifting
The slow drift was typically noticed in-between shifts and when orbits degraded:
Measurable effect on electron beam orbit
If left uncorrected, it will simply move the beam off the cameras (no bounded value)

## Slide 5: Sources of Fast Jumping
Previously collected data showed that the high-frequency component of the instability has a gaussian (non-deterministic) distribution, suggesting turbulence and random vibrations
Data over 30 minutes at 23 Hz
However, the air turbulence is well correlated with humidity in the trailer

## Slide 6: Sources of Slow Drifting
This is due to slow movement of tables and structures relative to one another with changing outside temperature
Abrupt transitions and periods of fluctuations correspond to changes in mode of operation (cw vs. pulsed vs. no beam)
Note that most drift occurs in x-direction, which translates into vertical motion in the trailer (change of orientation occurs in transport)

## Slide 7: Past Experience
Previous testing aimed at implementing a commercial fast stabilization system using quad diodes in transport showed limitations when used with LEReC
  Exception handling of shutter showed poor zeroing properties—beam would simply get lost upon resuming operation due to long transport
Recently collected data using CeC’s ~40 m transport arm between the relay table and the gun table. Each step is +10 mV on the steering mirror. ~Δ80 mV of change takes the beam through ~Δ1 mm of position. Corresponding LEReC arm is ~20 m.

## Slide 8: Other Obstacles
LEReC’s large and structured beam in the transport might complicate tracking (manufacturer specifies a 6mm beam on a 10mm x 10mm sensor)
  Beam is smaller and looks better in the trailer
Pulsed beam is outside the manufacturer’s specified bandwidth, which is 100 kHz
  Latency in electronics would not allow us to catch the same pulsed beam for differential measurement performed by quad diodes—might give spurious responses

## Slide 9: Current Approach
Transport
Separate fast- and slow-stabilization systems to maximize the impact of each system
  Fast, commercial unit in trailer before shutters
    Places fast correction where essentially all the air turbulence is
    Presence of ~50 µW of leakage light increases likelihood of effectiveness in pulsed mode
    Avoids problems with shutters opening and closing
  Slow, in-house, camera-based system corrects along the transport
    Leverages structure in beam to analyze and track changes
    Pulsed and cw beams appear identical with dynamic exposure and gain settings
    Built-in and expandable exception handling to handle shutters, etc.
Laser Trail
Tunnel
Slow Stabilization
Relay Table
Gun Table
Shutters
Laser
Fast Stabilization
Cathode
Moving air

## Slide 10: Fast-Feedback System Mock-Up
Initial test set-up in the temporary laser development area:
Actuator 1 (piezo mirror)
Camera
Laser
Detector 2 (quad diode)
Detector 1 (quad diode)
Actuator 2 (piezo mirror)
Distance between detectors ~1m

## Slide 11: Results
Initial 1m distance was too short. The following videos show the results when the distance was increased to ~1.4 m:
Environment was inherently noisy, with a lot of ambient light. Bandpass filters will be necessary in front of the quad diodes (the video above used none)
Without stabilization
With stabilization

## Slide 12: Slow-Feedback System
Uses camera and continuous image processing
Dedicated local loop—only interacts with Controls System via existing steering mirrors
Transport
Laser Trail
Tunnel
Relay Table
Gun Table
Shutters
Existing piezo mirror
Existing piezo mirror
New Camera
New Camera
Cathode

## Slide 13: Slow-Feedback Algorithm
Built-in exception handling and camera setting

## Slide 14: Edge-Based Image Processing
Highly structured beams are handled as easily (if not better) than purely gaussian beams
Optical artefacts, which are fixed features, essentially drop out
Image  gaussian 4-sigma linear filter  edge detect  difference  count (set threshold here)
Normal intensity fluctuations register less of a change than displacement. The opposite is true for a simple difference calculation.
LEReC beam at the start of shutdown
Example showing edge detection of 3 pixels (~16 μm)
Images showing a strict difference analysis for movement (left) and intensity change (right)

## Slide 15: Summary
Addressing the fast and slow instabilities separately will allow the final dual system not to have to make unnecessary compromises, which usually come with such huge dynamic ranges
With ~2m distance in the fast-stabilization system, we can expect to achieve the <50 µm rms fast stability requirement on the aperture, which becomes the <10 µm on the cathode
Based on data collected with CeC’s transport, we have enough grain to correct the slow drift to within 25 µm (corresponding to 5 mV on the piezos), which is also within the requirement
Timeline:
  Commercial fast-stabilization system is already on-site and verified as operational. Can be implemented in the trailer in about half a day once preceding stages (AOM and temporal shaping) are finalized
  Cables are expected to be pulled by mid-December, and a dedicated workstation is already available. Once hardware is in place, programming can continue into run if needed (non-invasive work)