## Slide 1
Magnet Mass Production and Field Measurements Results for TPS
C. S. Hwang, NSRRC, Hsinchu, Taiwan
IMMW 18, NSLSII (Long Island), 3 June 2013
Notes: Change name, proton, neutron, heavy iron & carbon therapy, I will try to give a overall introduction of magnet measurement system at NSRRC

## Slide 2: Outline
Field measurement system
- Hall Probe, Rotating coil, search coil
Magnet alignment method
Inspection of Storage ring magnets
Inspection of Booster ring magnets
Injection Magnets
Summary
Notes: Before the introduction, I will report the status of the new project of 3 GeV TPS.

## Slide 3
June 2005 | TPS Proposal to Government
Oct. 2007 | Lattice (circ. 518.4m) approved by BOT
Dec. 2007 | TPS final approval by Legislative Yuan
May 2008 | EPA approval; site plan completed
June 2008 | Accelerator Design Book (DRAFT) issued
Dec. 2008 | TPS budget (Civil + Accelerator + BL) lock in
Dec. 2009 | Contract out of the civil construction
Feb. 2010 | Ground Break
Dec. 2013 | Civil construction completed
Sep. 2014 | Installation completed & start to commission
Sep. 2015 | Open to users
Major Milestones of TPS project

## Slide 4
Status of the sub-components
Civil and utility construction will be completed before end of 2013.
The storage ring installation has been started and will be finished in the third quarter of 2014 and then start the commissioning.
150 MeV linac system has been completed and prepare to move into TPS location.
3 sets of superconducting RF system has been finished and the high power test is processing in house.
700 kW Cryogenic system is already and prepare to move to TLS location.
All magnets has been finished 80% by BSL company and will be done at the end of September.
All the vacuum chambers of the 24 sections will be finished before September.
All the power supply will be delivery to NSRRC before this year.
Control system were developed and will be installed and tested next year.

## Slide 5
Activity center &TPS outside view
Academic & Activity center
TPS building

## Slide 6
The inside view of TPS

## Slide 7: Overview of magnet Lab.
Measurement system | Measurement dimension | Purpose | Note
Hall probe (bench) | 6 m | ID and EPU | Pulley driven by DC motor
 | 5.4 m | Wiggle, Undulator & EPU | Linear motor with air bearing
 | 1.5 m | Superconducting undulator | In vertical dewar
 | 2 m | For Lattice magnet | Mechanical precision 20m
Rotating coil | 840 cm | QM and SM-storage ring | 
 | 570 cm | QM and SM-boost ring | 
NMR ESR | 0.043-13.7T | Hall sensor calibration | 
 | 0.55-3.2 mT |  | 
Search coil | Wire OD=0.12mm, 30turns Wire OD=0.05mm 30turns | Kicker and septum | 
Stretch wire | - | Insertion Devices | Measure 1st & 2nd field integral
Long loop coil | 5-axes | Insertion Devices | Measure 1st field integral
Helmholtz Coil | 3-axis rotation | Magnet block | 
Notes: We have many different length of Hall probe bench for different purpose of field measurement.

## Slide 8: Field measurement system for TPS magnets
Two Hall probe system (HPS):
Main field strength of HPS was calibrated by NMR & ESR (Nuclear magnetic resonance & electron spin resonance system) using dipole magnet.
This system used for all dipole magnet measurement and the inspection all of the prototype magnets.
Two Rotating coil system (RCS):
Main field strength of RCS was calibrated by HPS using quadrupole and sextupole magnets.
Field center of magnet was calibrated by CMM (3-D mechanical measurement system) using quadrupole and sextupole magnets.
Long coil & Search coil system
This system was used to measure Kicker and Septum. The pulse field and pulse current were recoded to be compared.

## Slide 9: HPS and measurement bench
Probe: 2m long
Granite bench
SR-dipole magnet adapter
BR-dipole magnet adapter

## Slide 10: SR-Dipole alignment on the HPS bench
SR-dipole adapter (3.75o): Rotate the SR-dipole magnet to parallel the Hall probe.
The magnet and adapter are located by the stop-pin and stop-surface.
The horizontal alignment is checked by the l theodolite.
The vertical height of magnet is checked by the eveling.
Stop-pin for z-axis
Stop-surface for x-axis
(touch surface)
Hall Probe
Target_1
Target_2
Target_3
3.75o trough
Adapter position check before SR-dipole installation. Hall probe, Tarage_1, Target_2 and Targer_3 in alignment.

+y
+x
+z
Target_4
Target_5
Target_2
Target_3
Line_1
Line_2
SR-dipole position check: Hall probe, Tarage_2, Target_3, Target_4 and Targer_5 in alignment.
Vertical height check: Hall probe, Line_1 (yoke seam) and Line_2 in alignment.

## Slide 11: HPS measurement and analysis
Hall probe
3D mapping: XY-Z
(circular & elliptical)
Hall probe trajectory

+x
+z
+y
Hall probe mapping trajectory: Curve and lamina
Hall probe
2D mapping: X-Z
(plane)
Normal (Skew) multipoles distribution along Z-axis.
Normalized multipoles.
Field homogeneity b0/b0 and b0L/b0L
Hall probe
1D mapping: Z
(on axis)
IB curve
Trim coil


## Slide 12: Rotating coil system (RCS)
SR-QM/SM magnet measured by SR-RCS
SR-FFC magnet measured by BR-RCS
Corrector magnet measured by BR-RCS
Magnet measured by BR-RCS

## Slide 13: Rotating coil - vertical offset calibration
The calibration value of RCS unit is -0.094 mm in the vertical offset.
The precision of magnet center measurement of this system in vertical axis is 5 m.
Average offset = 94 m
Magnet | Q1 | Q1 | Q9 | Q1 | Q5 | Q6 | R4Q8 | S1 | S2 | S3
SN | P01 | P02 | P01 | P03 | P01 | P01 | P01 | P01 | P01 | P01
CMM V-offset | 0.003 | 0.002 | -0.002 | -0.006 | -0.002 | 0.003 | -0.001 | -0.005 | 0.002 | 0.004
RCS V-offset | 0.106 | 0.098 | 0.084 | 0.085 | 0.087 | 0.100 | 0.085 | 0.089 | 0.099 | 0.106
Different | 0.103 | 0.096 | 0.086 | 0.091 | 0.089 | 0.097 | 0.086 | 0.094 | 0.097 | 0.102
STD = 5 m

## Slide 14
Rotating coil - horizontal offset
 | Q1 | Q10 | Q1 | Q9 | Q5 | Q6 | R4Q8 | S1 | S2 | S3
Note | P01 | P01 | P02 | P01 | P01 | P01 | P01 | P01 | P01 | P01
H nor. side | 0.005 (A) | 0.060 | 0.021 | 0.060 | 0.048 | 0.070 | 0.066 | 0.012 | 0.002 | 0.030
H opp. side | 0.018 (B) | 0.038 | 0.000 | 0.038 | 0.014 | 0.029 | 0.036 | 0.002 | 0.009 | 0.015
Decision | X1+X2 >> 2U | X1+X2 >> 2U | X1+X2 >> 2U | X1+X2 >> 2U | X1+X2 >> 2U | X1+X2 >> 2U | X1+X2 >> 2U | X1-X2 <<2U | X1-X2 << 2U | X1+X2 >> 2U
H offset | 0.012 (A+B)/2 | 0.049 (A+B)/2 | 0.011 (A+B)/2 | 0.049 (A+B)/2 | 0.031 (A+B)/2 | 0.050 (A+B)/2 | 0.051 (A+B)/2 | 0.005 (A-B)/2 | 0.004 (A-B)/2 | 0.023 (A+B)/2
Bench offset | 0.007 (A-B)/2 | 0.011 (A-B)/2 | 0.011 (A-B)/2 | 0.011 (A-B)/2 | 0.017 (A-B)/2 | 0.021 (A-B)/2 | 0.015 (A-B)/2 | 0.007 (A+B)/2 | 0.006 (A+B)/2 | 0.008 (A-B)/2
Test
差異<0.01mm
Average offset= 11 m
STD= 5 m
The calibration value of RCS unit is 11 m in the horizontal offset.
The precision of magnet center measurement of this system in horizontal axis is 10 m.

## Slide 15
Roll Angle Measurement-QM
All symbol is ABS
Normal side: N
Opposite side: O
PCB effective angle: P
Real angle: s
N = s + p ------(1)
N
P

N
N
S
S
(1)-(2)
s = (O+N)/2
O
P

N
N
S
S
Normal side of magnet
Opposite side of magnet
O = -s + p ------(2)
The resolution of rotary encoder is 9000.
The precision of the roll angle measurement is 0.01o.

## Slide 16
Roll Angle measurement by CMM
104.9
124.6
487.693
487.681
512.305
512.291
487.694 - 487.682
= 0.012 mm
(124.6+104.9)/2
= 114.75mm
=0.006o
487.682
487.694
512.314
512.285
512.314 - 512.285
= 0.029 mm
(124.6+104.9)/2
= 114.75mm
=0.014o
The roll angle of this magnet is 0.01o

## Slide 17: Magnet alignment & position check

## Slide 18
Positioning method of Quadrupole & Sextupole
Horizontal
reference
Vertical reference
Horizontal
reference
Vertical reference
The positioning precision of this method is within few m.
A PSD with laser is used to double check the system position.
Notes: Before the field measurement, the magnet should be align at the exact position. There are two touch reference plane to determine the horizontal transverse plane and four touch reference plane to determine the vertical plane.

## Slide 19
Position shimming method of Quadrupole &
Sextupole

## Slide 20
Type I
Type II
-z
+z
-z
+z
+y
+x
+z
Positioning method of SR-DM
Two reference plates are for defining the horizontal position and four reference plates are for defining vertical position.
45 degree clamper was used to push magnet to touch reference plane and fixed by the torque range tool.
The positioning precision of this method is also within few m.

## Slide 21: Magnet alignment method
45o bolt
Lift bolt
Tape-key block
Retractable block
RCS bench (similar as girder)
Both side measurement of RCS to indicate the field center of magnet.
Magnet on girder

## Slide 22: 1/24 section magnets

## Slide 23: TPS Magnets Alignment Inspection
A laser positioning system for TPS magnet alignment inspection is developed.
  System is composed of a Laser and 2 PSD with 2 granite blocks.
  PSD module is designed to install in quadrupole and sextupole magnet
  Position jig is designed to symbolize magnet
Position Jig
PSD module
  The optical axis of Laser is adjusted to parallel to girder through reference to position jig
PSD
Girder
Position Jig2
(granite block)
Position Jig1
(granite block)
Laser
PSD1
PSD2
Notes: t
23

## Slide 24: Magnets alignment on the girder
A PSD module with circular jig was developed to check the magnet center of all magnet in the same girder.
The measurement result was confirmed that all the magnet center is within 30 m.
PSD module
circular jig
Shielding
Laser

## Slide 26: Measurement of Storage ring magnets

## Slide 27
SR-magnet specification
(x10-4) | SR-DM
n | BnL/B0L
0 | 10000
1 | ±3
2 | ±3
3 | ±2
4 | ±3
Normalized at 25mm b0/b0<1E-4 b0L/b0L<0.5E-3 | 
(x10-4) | SR-QM | 
n | BnL/B1L | AnL/B1L
1 | 10000 | -
2 | ±2 | ±2.0
3 | ±2.0 | ±1.0
4 | ±0.5 | ±0.3
5 | ±0.8 | ±0.3
6 - 8 | ±0.3 | ±0.3
9 | ±0.3 | ±0.3
10 - 29 | ±0.3 | ±0.3
Normalized at 25mm |  | 
(x10-4) | SR-SM | 
n | BnL/B2L (x10-4) | AnL/B2L (x10-4)
0 | ±15 | ±10
2 | 10000 | -
3 | ±2 | ±2
4 | ±3 | ±1
5 - 7 | ±0.5 | ±0.5
8 | ±0.5 | ±0.3
9 | ±0.3 | ±0.3
10 - 13 | ±0.3 | ±0.3
14 | ±0.3 | ±0.3
15 - 29 | ±0.3 | ±0.3
Normalized at 25mm |  | 

## Slide 28
SR-DM-I and SR-DM-II
Field strength
21 SR dipole magnets were examined at NSRRC.
(50 magnets required)
The dispersion of b0L is better than 0.11%.

## Slide 29
Homogeneity b0/b0 and b0L/b0L
SR-DM-I and SR-DM-II
The b0/b0 and b0L/b0L are meet spec within good field region (GFR).

## Slide 30
SR-short QM
Field strength
159 SR short-QM magnets were examined at NSRRC until 2013/05/30. (195 magnets required)
The dediation of b1L will be compensated by independent power supply.

## Slide 31
Normalized multipoles
SR-short QM
The mostly normalized multipoles of SR-short QM are within spec.

## Slide 32
SR-short QM
Center offset and roll angle
The mean value of magnetic center offset & magnet roll are within 0.01mm in the vertical (horizontal) direction and 0.012o.
 | 159 short_QM magnets | 
 | Mean | Std. Dev.
Magnetic center, Vertical-offset (mm) | 0.002 | 0.009
Magnetic center, Horizontal-offset (mm) | 0.005 | 0.012
Magnetic Roll angle (degree) | -0.0063 | 0.012
Mechanical center V-offset (mm) | -0.002 | 0.004
Mechanical center H-offset (mm) | -0.002 | 0.006
Mechanical Roll angle (degree) | 0.003 | 0.005

## Slide 33
SR-long QM
Field strength
28 SR long-QM magnets were examined at NSRRC until 2013/05/30. (50 magnets required)
The prototype magnet have a large field dispersion.

## Slide 34
Normalized multipoles
SR-long QM
The mostly normalized multipoles of SR-long QM are within spec.

## Slide 35
SR-long QM
Center offset and roll angle
The mean value of magnetic center offset & magnet roll are within 0.01mm and 0.012o in the vertical and horizontal direction.
 | 28 long_QM magnets | 
 | Mean | Std. Dev.
Magnetic center, Vertical-offset (mm) | 0.003 | 0.008
Magnetic center, Horizontal-offset (mm) | -0.008 | 0.012
Magnetic Roll angle (degree) | -0.004 | 0.009
Mechanical center V-offset (mm) | -0.004 | 0.003
Mechanical center H-offset (mm) | -0.004 | 0.006
Mechanical Roll angle (degree) | 0.003 | 0.004

## Slide 36
Multipoles distribution of Q-magnet along Z-axis

## Slide 37
Multipoles distribution of Q-magnet along Z-axis
Exact Skew quadrupole term can be obtained after the angle calibration of Hall probe .

## Slide 38
SR-Sextupole Magnet (SM)
Field strength
132 SR-SM magnets were examined at NSRRC until 2013/05/30. (174 magnets required)
The large dispersion is dominated by prototype magnet.

## Slide 39
Normalized multipoles
SR-SM
The mostly normalized multipoles of SR-SM are within spec.

## Slide 40
SR-SM
Center offset and roll angle
The mean value of magnetic center offset & roll are within 0.011mm in the vertical (horizontal) direction and 0.017o.
 | 132 SM magnets | 
 | Mean | Std. Dev.
Magnetic center, Vertical-offset (mm) | 0.004 | 0.009
Magnetic center, Horizontal-offset (mm) | 0.000 | 0.011
Magnetic Roll angle (degree) | 0.008 | 0.017
Mechanical center V-offset (mm) | -0.001 | 0.003
Mechanical center H-offset (mm) | -0.002 | 0.005
Mechanical Roll angle (degree) | 0.006 | 0.006

## Slide 41: Measurement of Booster ring magnets

## Slide 42
BR-magnet specification
x10-4 | BR-DM
n | BnL/B0L
0 | 10000
1 | *
2 | **
3 | ±3
4 | ±2
Normalized at 15mm b0/b0<5E-4 b0L/b0L<0.5E-3 *b1L/b0L= -2.1043 **b2L/b0L= -7.5331 | 
x10-4 | BR-Pure QM | 
n | BnL/B1L | AnL/B1L
1 | 10000 | -
2 | ±4 | ±10
3 | ±4 | ±2
4 | ±1 | ±1.5
5 | ±3 | ±0.5
6-7 | ±1 | ±0.5
8 | ±0.5 | ±0.5
9 | ±4 | ±0.5
10-12 | ±0.5 | ±0.3
13 | ±1.5 | ±0.5
14-16 | ±0.5 | ±0.3
17 | ±1.7 | ±0.5
18-20 | ±0.3 | ±0.3
Normalized at 15mm |  | 
x10-4 | BR-SM | 
n | BnL/B2L | AnL/B2L
0 | 45 | 30
2 | 10000 | -
3 | 15 | 6
4 | 9 | 6
5-7 | 3 | 1.5
8 | 10 | 1.5
9-13 | 3 | 1.5
14 | 6 | 1.5
15-20 | 3 | 0.6
Normalized at 15mm |  | 
x10-4 | BR-combined QM | 
n | BnL/B1L | AnL/B1L
1 | 10000 | -
2 | - | -
3 | ±4 | ±2
4 | ±4 | ±1
5 | ±2 | ±0.5
6 | ±1 | ±1
7 | ±1 | ±0.5
8 | ±5 | ±0.5
9 | ±2 | ±0.5
10 | ±4 | ±0.5
11 | ±0.5 | ±0.5
12-13 | ±2 | ±0.3
14-17 | ±0.5 | ±0.3
18-19 | ±0.3 | ±0.3
20 | ±0.5 | ±0.3
Normalized at 15mm b2L/b1L=1.145 |  | 

## Slide 43
BR-combined function dipole magnet
Homogeneity of BR-BH-001 and BR-BD-002 magnets
Increase the slope of pole profile to increase the quadrupole component.
Cancel the end-chamfer of pole.
Yoke length of BH dipole magnet reduce 4 mm from 758 mm to 754 mm.
Magnet | Items | b1L/b0L |  | b2L/b0L | 
 |  | Design | Measure | Design | Measure
 | Spec. | -2.1043 | - | -7.5331 | -
BD (1.6m) | Original simulation | -2.0365 | -2.049 | -7.5524 | -8.480
 | Optimized simulation | -2.1041 | - | -7.6129 | -
BH (0.8m) | Original simulation | -1.9758 | - | -7.1151 | -
 | Optimized simulation | -2.1050 | - | -7.7164 | -

## Slide 44
BR-combined function dipole magnet
Homogeneity of BR-BH-001 and BR-BD-002 magnets
The b0/b0 of BR-BH-001 and BR-BD-002 is better than 4.510-4 and -3.010-4, respectively
The b0L/b0L of BR-BH-001 and BR-BD-002 is better than 1.910-4 and -5.310-4, respectively.

## Slide 45
Field strength
BR-combined function dipole magnet
The measurement of ByL/I is confirmed by the calculated value with ignore r and Br between 100 A to 900 A.
The ByL/I of BR-BH-001 magnet has clearly increasing is due to the remnant field effect in the low current.
The ByL/I of BR-BH-001 magnet has slowly decreasing is due to the yoke situation over 900A.

## Slide 46
Combined multipoles along electron-beam trajectoy (Z-axis)
BR-combined function dipole magnet
The multipole distributions of measurement are agreement with the simulation result.

## Slide 47
Combined function quadrupole magnet
Field strength: Pre-measurement at vendor site
<b1L> @ 82A = 3.3647 T , Std @ 82A = 0.0032T
<b1L> @ 5A = 0.2280 T , Std @ 5A = 0.0004T
<b2L> @ 82A = 3.7871 T, Std @ 82A = 0.0356T
<b2L> @ 5A = 0.2602 T, Std @ 5A = 0.0161T
<b2L/b1L> @ 82A = 1.1256 T, Std @ 82A = 0.0101T
<b2L/b1L> @ 5A = 1.1409 T, Std @ 5A = 0.0695T
The b2L/b1L deviation of 5 A is larger than 82 A, because the remnant field obviously influence in the low current.

## Slide 48
Combined function quadrupole magnet
Normalized multipoles: Pre-measurement at vendor site
Dominated by remnant field in the low current.
The b3 component is same in the 0A ~ 5A
All of multipoles distribution, exclude B3L/B1L and B5L/B1L, are quite similar in the 82 A and 5 A.

## Slide 49
Remnant field improvement of BR-QP magnet
 |  |  | Original |  | Improved | 
Magnet L=0.3 (m) |  | Spec. b1L (T) | Coil turns | Energy @ current | Coil turns | Energy @ current
QP | Q1 | 4.2985 | 18 | 3GeV @ 104A | 18 | 3GeV @ 104A
 |  | 0.2149 |  | 150MeV @ 5.1A |  | 150MeV @ 5.1A
 | Q2 | -2.7174 | 18 | 3GeV @ 66A | 12 | 3GeV @ 99A
 |  | -0.1359 |  | 150MeV @ 3.3A |  | 150MeV @ 5.0A
 | QM | -1.2623 | 18 | 3GeV @ 30A | 6 | 3GeV @ 90A
 |  | -0.0631 |  | 150MeV @ 1.5A |  | 150MeV @ 4.5A
The excitation current of BR-pure quadrupole magnet will be increased after coil no. reduced.
The remnant field effect will be solved by reducing length for the low-current (150MeV) increase.
Coil turns: 18
Coil turns: 6
Coil turns: 12

## Slide 50
Pulse magnet measurement

## Slide 51: Pulse Magnet Type
Half-sine Kickers
  SR injection kickers x 4
Half-sine Septum magnet
  Booster injection septum x 1
  Booster extraction septum x 1
  SR injection septum x 1
PFN Kickers
  Booster injection kicker x 1
  Booster extraction kicker x 1

## Slide 52: Long Coil and Search coil Measurement system for Storage Ring Kicker Magnet
Long Coil
Search Coil
SCOPE
Integrator
Function Generator
Long coil : One turn loop of length 1000 mm and width 1.7 mm
Search coil : The is 5 mm, from winding thirty turns of enamel wire (thickness 0.5 mm)

## Slide 53: Storage ring Kicker Magnet
Integral Bfield Roll-Off
B-I Curve

## Slide 54: Long Coil and Search coil Measurement Systems for SR Injection and BR Extraction Septum
Long Coil
Search Coil
SCOPE
Integrator
Function Generator
Long coil :
  Bending curve : Measures the on-axis magnetic field of the septum magnet
  Straight curve : Measures the leakage magnetic field of the bumped orbit and stored orbit
  One turn loop of length 1000 mm and width 1.7 mm.
Search coil : The is 5 mm, from winding thirty turns of enamel wire (thickness 0.5 mm)
Bending curve probe
Straight curve probe

## Slide 55: Storage ring Injection Septum Magnet
By Leakage field | 
Bumped orbit (G-cm) | 90 (0.15%)
Stored orbit (G-cm) | 18 (0.03%)
Bx Leakage field | 
Bumped orbit (G-cm) | 15 (0.03%)
Stored orbit (G-cm) | 4 (0.01%)
By leakage field
Bx leakage field

## Slide 56: Field measurement of Booster Injection and Extraction Kicker Magnet
Long Coil
Search Coil
SCOPE
Integrator
Function Generator
Long coil : One turn loop of length 1000 mm and width 1.7 mm
Search coil : The is 5 mm, from winding thirty turns of enamel wire (thickness 0.5 mm)

## Slide 57: Booster Injection Kicker Magnet
Parameters | Measurement
Pulse shape | Flat top
Fall time (ns) | 400
Flat top (us) | 1
Flatness (%) | ± 1
Current output
Integral Bfield Roll-Off

## Slide 58
Most of the magnet can be meet the specification. However, much effort to solve the technique problem is necessary.
The deviation of b0L of SR-dipole magnets is better than 0.1 %.
Deviation of magnetic center & roll angle are within 0.02 mm and 0.02o.
A PSD module with circular jig can be used to check all the magnet center on the same girder.
The ratio of quadrupole and sextupole of the combine dipole magnet is quite consistence between the TOSCA simulation and the construction results.
The deviation of b2L/b1L of QF magnet with current below 5 A is much larger than 82 A, because the remnant field exist in the hysteresis behavior.
The coil turns and length of booster magnets should be as small as possible to avoid the current ripple and the remnant field effect in the low energy injection region.
Summary

## Slide 59
Thanks for your attention